| APPROVALS | DATE | REV | DCN NO. | BY | CHECK | DCC | DATE |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AUTHOR: Betsy Bland, John Worden | 01/28/09 |  |  |  |  |  |  |
| CHECKED: |  |  |  |  |  |  |  |
| APPROVED: |  |  |  |  |  |  |  |
| DCC RELEASE |  |  |  |  |  |  |  |

Requirements

General

- Washer shall be a front loading machine.

- Maximum overall dimensions:  84x84” by 126” in height

- The exterior of the machine shall be insulated to limit thermal hazards and radiation loss.

- Washer will be installed on an existing concrete floor.

- Unit shall be powered with 480 volt, 3-phase.

- An 8” port shall be provided for possible exhaust provisions.

- Washer shall meet applicable safety and electrical codes.

- Washer shall fail safe in the event of loss of utilities.

- Washer shall be compatible with Liquinox detergent (pH 8.5).

Wash Chamber

- Minimum chamber load capacity:  36x36” by 50” in height and 500 lbs in weight.

- The wash chamber shall be constructed of stainless steel, as corrosion inhibiting chemicals will not be used.

- Other materials used in the assembly, such as gaskets, shall be selected to minimize particulate.

- No lubrication is permitted within the fluid path of the washer, or in the chamber.

- Machine shall include base model parts rack.

- The machine shall utilize rotating wash/spray arms and/or rotating turntable, or other mechanism to maximize spray coverage of parts.

Wash Controls

- Washer shall have controls for wash and rinse cycles, and display for wash and rinse temperatures.

- Unit shall have a heated wash tank for a water based solution.

- Wash tank and rinse tank heaters shall be adjustable from 150 deg F to 190 deg F.

- Cycle durations shall be adjustable from 1 minute to 30 minutes.

Pumps and Plumbing

- All plumbing shall be stainless steel to the extent possible.

- Wash water shall be recirculated.

- Unit shall have a separate heated rinse tank for non recirculating fresh water (possibly deionized water) cycle.

- Wash cycles shall be driven by a pump of 7.5 HP - 15 HP, with flow rates in the range of 100-250gpm.

- Rinse cycle shall be driven by a dedicated pump with flow rates in the range of 10-20gpm.

- All tank fill and drain operations may be manual.

- Wash tank shall have a minimum capacity of 65 gallons.

- Rinse tank shall have a minimum capacity of 20 gallons.

- All rinse and wash water shall be filtered and/or strained.

- Recycled wash water shall be filtered with a 10 micron filter.

- Rinse water shall be filtered with a 1 micron filter.

- Spray headers shall be configured to maximize coverage to all surfaces of parts.

Optional Features

- Forced hot air Drying System

- Spray gun for manual rinse

- Additional racks

- Multiple wash/rinse cycle control