LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T0900321-v1	Advanced LIGO UK       June 2009

Satellite Box Users’ Guide

R. M. Cutler, University of Birmingham

Distribution of this document:

Inform aligo_sus

This is an internal working note

of the Advanced LIGO Project, prepared by members of the UK team.

| Institute for Gravitational Research University of Glasgow Phone +44 (0) 141 330 5884 Fax +44 (0) 141 330 6833 E-mail k.strain@physics.gla.ac.uk | School of Physics and Astronomy University of Birmingham Phone +44 (0) 121 414 6447 Fax +44 (0) 121 414 3722 E-mail av@star.sr.bham.ac.uk |
| --- | --- |
| Engineering Department CCLRC Rutherford Appleton Laboratory Phone +44 (0) 1235 445 297 Fax +44 (0) 1235 445 843 E-mail J.Greenhalgh@rl.ac.uk | Department of Physics University of Strathclyde Phone +44 (0) 1411 548 3360 Fax +44 (0) 141 552 2891 E-mail N.Lockerbie@phys.strath.ac.uk |

http://www.ligo.caltech.edu/

http://www.physics.gla.ac.uk/igr/sus/

http://www.sr.bham.ac.uk/research/gravity/rh,d,2.html

http://www.eng-external.rl.ac.uk/advligo/papers_public/ALUK_Homepage.htm

Satellite Box Users Guide

Contents

1. Document Organization

2. LIGO

2.1 The Mirror Suspension System

2.2    The Control System

3    The Satellite Box

3.1 The Current Source

3.2 The Photodiode Amplifier

3.3 The Window Detector

3.4 Use of LIGO1 OSEMs

4     Using the Unit

4.1   Connectors

4.2   Connector Functions

4.2.1 J1 to the Vacuum Tank

4.2.2  J2 to the Coil Drive Unit in the Analogue Rack

4.2.3 J3 to the Coil Drive Unit in the Analogue Rack

4.2.4  J4 carrying local Monitor signals

5.   Connector Details

5.1   J1 Connections to the Vacuum Tank

5.2   J2 Connections to the Coil Driver in the Analogue rack

5.3   J3 Connections to the Coil Drive Unit in the Analogue Rack

5.4   J4 Connections local diagnostic Connector

6.   References

## 1.  DOCUMENT ORGANIZATION

Section 2 introduces the LIGO suspensions, and describes the control             system

Section 3 describes the functions of the Satellite Box

Section 4 describes how to connect the unit

Section 5 gives the pin functions of the connectors.

Section 6 gives references.

LIGO

The LIGO (Laser Interferometric Gravitational-wave Observatory) experiment is designed to detect gravity waves.

The laser interferometery method used involves mirrors which are held in a steady position to a very high degree of precision.

2.1 The Mirror Suspension System

FIG 1 The LIGO Quad suspension.

Mirror stability is achieved by means of a suspension consisting of 4 heavy masses which form a compound pendulum. The position and movement of the top three stages is controlled electromagnetically, while the bottom stage on which the mirror is mounted, is controlled electrostatically.

The purpose of position control is to maintain the output of the interferometer on a dark fringe. Movement control also damps the low frequency pendulum modes.

2.2 The Control System

The Satellite Boxes form part of the control system which controls the position and movement of the pendulum stages electromagnetically. Fig 2 is a simplified diagram of the control system, illustrating the role of the Satellite Box.

Fig 1 The Control System

The demand signals to the Drive Amplifiers are generated by the Digital Control System, converted to analogue signals, and then filtered in the AI Unit. The signals are then amplified and filtered by the Drive Amplifiers and pass through the Satellite Box to provide the demanded currents to the drive coils. The position of each stage is indicated by mechanical “flags” to the OSEM optical position detectors. The position detectors consist of LED and photodiode pairs. The photodiode outputs are amplified in the Satellite Box, and then fed back to the Digital Control System via the filters in the AA unit to the A/D convertors.

3. The Satellite Box

Fig 2 Satellite Box Block diagram

The Satellite Box contains three main functional blocks:

1. The Current Source

2. The Photodiode Amplifier

3. The Window Detector

3.1 The Current Source

The current source provides a stable, low noise current supply (35 mA nominal) to each of 4 OSEM LEDs.

3.2 The Photodiode Amplifier

The photodiode amplifiers amplify and filter the output current from the outputs of 4 OSEM photodiodes. The output is a differential pair. The nominal output is +/- 10 volts for 60μA input.

3.3 The Window Detector

The window detector monitors the operation of the current sources. If they are not operating within specification, an error light on the front panel becomes illuminated.

3.4 Use of LIGO1 OSEMs

The satellite box is designed to be used with either the new Birmingham OSEMs, or with the old LIGO 1 OSEMs as required.

The most important difference between these OSEMs is that the LIGO 1 OSEMs require a bias voltage on the photodiodes, whereas the Birmingham OSEMs do not. The satellite box has an internal link which may be connected when bias is required. To use the Satellite Box with Birmingham OSEMs, connect the links 105, 205, 305 and 405 in the ‘B’ position.  To use the Satellite Box with LIGO1 OSEMs, connect the links 105, 205, 305 and 405 in the ‘L’ position.

4 USING THE UNIT

To use the unit, first connect it to the system as described below. The power should be switched off at this stage. The input voltage should be checked before connection (Positive voltage = 17v nominal, 15.5v minimum, 18v maximum, negative voltage = -17v nominal, -15.5 volts minimum, -18 volts maximum.)

Having checked the power supply and connected all the cables correctly, the power may be switched on. The green Power LED indicator lights on the front panel should then become illuminated, indicating that the on board regulators are operating and that the 14v, -14v and 5v supplies are running.

During the power up sequence the red LED Fault indicators should illuminate momentarily.

4.1 Connectors

Figure 3 The Satellite Box interconnections

The connections to the Satellite Box were shown on Fig 3.

Details of the connectors follow.

4.2 Connector Functions

The unit should be connected to the system in the following way.

4.2.1 J1 to the Vacuum Tank

J1 is the 25way connector to the vacuum tank. It supplies the 35 mA nominal current to the OSEM LEDs, and collects the photodiode currents from the OSEMs. The coil drive currents, which have passed through the satellite box unprocessed, supply the OSEM coils via J1.

4.2.2 J2 to the Coil Drive Unit in the Analogue Rack

J2 is connected to the appropriate Coil Drive Unit in the Analogue Rack. It carries three types of connections: Power to the satellite box, LED current monitoring signals and the photodiode outputs.

4.2.3 J3 to the Coil Drive Unit in the Analogue Rack

J3 receives the Coil Drive signals from the appropriate Coil Drive Unit. These signals emerge from J1, subsequently driving the coils.

4.2.4  J4 carrying local Monitor signals

J4 is intended for local diagnostics, and makes most of the key signals available for checking and fault finding.

5. Connector Details

J1 Connections to the Vacuum Tank

J2 Connections to the Analogue rack

J3 Connections to the Analogue Rack

J4 Connections to the local diagnostic Connector

Details of these connections follow.

5.1    J1 Connections to the Vacuum Tank

25 way Female ‘D’ connector

| PIN | SIGNAL | DESCRIPTION |
| --- | --- | --- |
| 1 | Coil D ST | Coil D Start |
| 2 | LED D K | LED D Cathode |
| 3 | PD D A | Photodiode D Anode |
| 4 | Coil C ST | Coil C Start |
| 5 | LED C K | LED C Cathode |
| 6 | PD C A | Photodiode C Anode |
| 7 | Coil B ST | Coil B Start |
| 8 | LED B K | LED B Cathode |
| 9 | PD B A | Photodiode B Anode |
| 10 | Coil A ST | Coil A Start |
| 11 | LED A K | LED A Cathode |
| 12 | PD A A | Photodiode A Anode |
| 13 | Shield |  |
| 14 | Coil D FN | Coil D Winding Finish |
| 15 | LED D A | LED D Anode |
| 16 | PD D k | Photodiode D Cathode |
| 17 | Coil C FN | Coil C Winding Finish |
| 18 | LED C A | LED C Anode |
| 19 | PD C K | Photodiode C Cathode |
| 20 | Coil B FN | Coil B Winding Finish |
| 21 | LED B A | LED B Anode |
| 22 | PD B K | Photodiode B Cathode |
| 23 | Coil A FN | Coil A Winding Finish |
| 24 | LED A A | LED A Anode |
| 25 | PD A K | Photodiode A Cathode |

5.2    J2 Connections to the Coil Driver in the Analogue rack

25 way Male ‘D’ connector

Power and PD signals

| PIN | SIGNAL | DESCRIPTION |
| --- | --- | --- |
| 1 | DLD A+ | Differential Line Diver Ch A, ( + Sensor o/p) |
| 2 | DLD B+ | Differential Line Diver Ch B, ( + Sensor o/p) |
| 3 | DLD C+ | Differential Line Diver Ch C, ( + Sensor o/p) |
| 4 | DLD D+ | Differential Line Diver Ch D, ( + Sensor o/p) |
| 5 | I mon A | LED Current monitor |
| 6 | I mon B | LED Current monitor |
| 7 | I mon C | LED Current monitor |
| 8 | I mon D | LED Current monitor |
| 9 | +17v |  |
| 10 | +17v |  |
| 11 | -17v |  |
| 12 | -17v |  |
| 13 | Shield |  |
| 14 | DLD A- | Differential Line Diver Ch A, ( - Sensor o/p) |
| 15 | DLD B- | Differential Line Diver Ch B, ( - Sensor o/p) |
| 16 | DLD C- | Differential Line Diver Ch C, ( - Sensor o/p) |
| 17 | DLD D- | Differential Line Diver Ch D, ( - Sensor o/p) |
| 18 | Ref gnd A | LED Current monitor 0v ref |
| 19 | Ref gnd B | LED Current monitor 0v ref |
| 20 | Ref gnd C | LED Current monitor 0v ref |
| 21 | Ref gnd D | LED Current monitor 0v ref |
| 22 | 0v |  |
| 23 | 0v |  |
| 24 | 0v |  |
| 25 | 0v |  |

5.3   J3 Connections to the Coil Drive Unit in the Analogue Rack

15 way Male ‘D’ connector

Coil drive signals

| PIN | SIGNAL | DESCRIPTION |
| --- | --- | --- |
| 1 | ST A | Coil Winding A Start |
| 2 |  |  |
| 3 | ST B | Coil Winding B Start |
| 4 |  |  |
| 5 | ST C | Coil Winding C Start |
| 6 |  |  |
| 7 | ST D | Coil Winding D Start |
| 8 | Shield |  |
| 9 | FN A | Coil Winding A Finish |
| 10 |  |  |
| 11 | FN B | Coil Winding B Finish |
| 12 |  |  |
| 13 | FN C | Coil Winding C Finish |
| 14 |  |  |
| 15 | FN D | Coil Winding D Finish |

5.4    J4 Connections local diagnostic Connector

37 Way Female ‘D’ Connector

| PIN | SIGNAL | DESCRIPTION | Normal voltages |
| --- | --- | --- | --- |
| 1 | I mon A | LED Current Monitor A | 1.000 V |
| 2 | I mon B | LED Current Monitor B | 1.000 V |
| 3 | I mon C | LED Current Monitor C | 1.000 V |
| 4 | I mon D | LED Current Monitor D | 1.000 V |
| 5 | V mon A | LED Voltage Monitor A | 1.6v nominal |
| 6 | V mon B | LED Voltage Monitor B | 1.6v nominal |
| 7 | V mon C | LED Voltage Monitor C | 1.6v nominal |
| 8 | V mon D | LED Voltage Monitor D | 1.6v nominal |
| 9 | PD A+ | Photodiode Amp A + Mon | +10V to -10V |
| 10 | PD B+ | Photodiode Amp B + Mon | +10V to -10V |
| 11 | PD C+ | Photodiode Amp C + Mon | +10V to -10V |
| 12 | PD D+ | Photodiode Amp D + Mon | +10V to -10V |
| 13 | Coil A + Mon | Coil A + Voltage Monitor | +12V to -12V (max.) |
| 14 | Coil B + Mon | Coil B + Voltage Monitor | +12V to -12V (max.) |
| 15 | Coil C + Mon | Coil C + Voltage Monitor | +12V to -12V (max.) |
| 16 | Coil D + Mon | Coil D + Voltage Monitor | +12V to -12V (max.) |
| 17 | +14v | +14v monitor | +7v |
| 18 | +5V MON | + 5v monitor | +5v |
| 19 | Shield |  |  |
| 20 | I mon A return | Current mon A return | 0V |
| 21 | I mon B return | Current mon B return | 0V |
| 22 | I mon C return | Current mon C return | 0V |
| 23 | I mon D return | Current mon D return | 0V |
| 24 | V mon A return | Voltage mon A return | 0V |
| 25 | V mon B return | Voltage mon B return | 0V |
| 26 | V mon C return | Voltage mon C return | 0V |
| 27 | V mon D return | Voltage mon D return | 0V |
| 28 | PD A - | Photodiode Amp A - Mon | -10V to +10V |
| 29 | PD B - | Photodiode Amp B - Mon | -10V to +10V |
| 30 | PD C - | Photodiode Amp C - Mon | -10V to +10V |
| 31 | PD D - | Photodiode Amp D - Mon | -10V to +10V |
| 32 | Coil A - Mon | Coil A - Voltage Monitor | -12V to+12V (max.) |
| 33 | Coil B - Mon | Coil B - Voltage Monitor | -12V to+12V (max.) |
| 34 | Coil C - Mon | Coil C - Voltage Monitor | -12V to+12V (max.) |
| 35 | Coil D - Mon | Coil D - Voltage Monitor | -12V to+12V (max.). |
| 36 | -14v | -14v monitor | -7v |
| 37 | 0V | 0V Supply | 0V Supply |

6. REFERENCES

(1) Quadruple Suspension Design for Advanced LIGO 	P020001-A-R

N A Robertson, G Cagnoli, D R M Crooks, E Elliffe, J E Faller, P Fritschel, S Goßler, A Grant, A Heptonstall, J Hough, H Lück, R Mittleman, M Perreur-Lloyd, M V Plissi, S Rowan, D H Shoemaker, P H Sneddon, K A Strain, C I Torrie, H Ward, P Willems

# Text boxes

TOP

UIM

PUM

MIRROR

STAGE

Photodiode

Coils

LED

OSEM

Photodiode

Amplifiers

Current Sources

Satellite Box

Coil Drivers

Monitors

AA Unit

Coil Drive Unit

D/A

Convertors

Demand

Digital Control System

AI

Unit

D/A

Convertors

Local

Diagnostic

Monitoring

Signals

From Coil Drive Unit

Coil Drive

Signals

J3

J2

Photodiode

Outputs

I mon

Power

To

Analogue

Rack

J1

To

Vacuum

Tank

From Photodoides

To Coils & LEDs

SATELLITE BOX

Monitors

Output

Photodiode

Current in

Indicators

Monitors

LED Current

PHOTO DIODE

AMPLIFIER

WINDOW DETECTOR

CURRENT SOURCE

Flag                                    Magnet

Suspension Stage

J4