LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T0900325-v2	Advanced LIGO	12/5/11

H1 Squeezer Noise Model

Sheila Dwyer, Sheon Chua, Michael Stefszky, Daniel Sigg, Nergis Mavalvala

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note
of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu

Table of Contents

1	Sensing and Actuation	4

1.1	Overview	4

1.2	Main laser	4

1.3	Auxiliary laser	4

1.4	SHG	6

1.5	OPO	7

1.6	Squeezed Field	7

1.7	Coherent Locking	8

1.8	Local Oscillator	10

2	Feedback Compensation Networks	12

2.1	Laser Locking	12

2.2	SHG and OPO length	12

2.2.1	SHG	12

2.2.2	OPO	13

2.3	Coherent Locking	13

2.4	Local Oscillator	14

3	Noise Couplings and Requirements	16

3.1	Acoustic Couplings	16

3.1.1	Back Scattering	16

3.1.2	Beam Jitter	17

3.2	Squeezed Beam Requirements	18

3.2.1	Squeezer Level	18

3.2.2	Phase Noise	18

3.3	Losses	19

3.4	Modulation Sidebands	19

3.5	Noise Couplings	21

3.5.1	Laser Noise	21

3.5.2	Path Length Fluctuations	23

3.5.3	Displacement Noise	24

3.5.4	Shot Noise	26

4	Fabry-Perot Resonators	29

4.1	Cavity Parameters	29

4.1.1	SHG	29

4.1.2	OPO	29

4.1.3	Interferometer	30

4.2	Cavity Equations	30

4.2.1	Pound-Drever-Hall Signals	30

4.2.2	Electric Fields	31

4.2.3	Non-Linear Process	33

4.2.4	Offset Locking	34

5	Compensation Design	35

5.1	Laser Locking	35

5.2	SHG and OPO	36

5.3	Coherent Locking and Local Oscillator	37

# Sensing and Actuation

## Overview

A block diagram of the servo model and noise propagation model is shown in Figure 1: Servo model of the H1 squeezer.Figure 2.

## Main laser

The main laser, fMain, is locked to the fiber output which is taken from the interferometer laser, fPSL,. This will guarantee that the squeezer is operating on the correct frequency. In order to lock the phase a feedback path from the local oscillator (S3) adjusts the offset of the frequency stabilization servo (FSS). The error signal expressed as phase angles can be written as

|  | V0ϕs=ϕMain-ϕPSL+ϕ0 |  |
| --- | --- | --- |

where ϕ0 denotes the demodulation angle. One can always do the calculation in phase by simply multiplying the frequency by 1s. We absorb the scaling factor due to the light level on the photodetector into the photodetector transfer function

|  | S0(s)=2ηPSL,0PPSLηMain,0PMainηQE,0εRZ0 |  |
| --- | --- | --- |

where PPSL, PMain, ηQE, εR, Z0 denote the light available from the PSL, the light power of the main laser, the photodetector quantum efficiency, the photodetector responsitivity, and the photodetector transimpedance gain. The factor ηPSL,0 and ηMain,0 are the attenuation factors of the two interfering beams, respectively, and account for beamsplitter ratios and attenuators.

## Auxiliary laser

The auxiliary laser is used to generate a frequency shifted subcarrier which is used for coherent locking of the squeezer phase to main interferometer beam at the antisymmetric port. It deploys a second FFS. The offset point into the error signal of the servo can be used as a fast feedback for the phase angle between the green pump beam into the OPO and the coherent field. The error signal and photodetector transfer function are identical to the ones from the main laser:

|  | V1ϕs=ϕAux-ϕMain+ϕ1 |  |
| --- | --- | --- |

where ϕ1 denotes the demodulation angle. We absorb the scaling factor to the light levels into the photodetector transfer function

|  | S1(s)=2ηMain,1PMainηAux,1PAuxηQE,1εRZ1 |  |
| --- | --- | --- |

where Z1 denotes the photodetector transimpedance gain. The factors ηMain,1 and ηAux,1 again account for the beam attenuation in each interfering beam, respectively.

Figure 1: Servo model of the H1 squeezer.

## SHG

The SHG is locked using a Pound-Drever-Hall signal in transmission of the cavity. For the red light we get the following formulae:

|  | TTRANS,fMainSHG,ϕ(s)=gSR,TRANSs2πfS,FSR |  |
| --- | --- | --- |
|  | TTRANS,fMainSHG,δl(s)=gSR,TRANSkπ1+s2πfSR |  |

with

|  | k=2πλR |  |
| --- | --- | --- |
|  | gSR,TRANS=-4J0ΓS,MJ1ΓS,MFSRrSR,1rSR,21-δSR tSR,SB tSR,CR |  |
|  | tSR,SB=1-δSR1-rSR,121-rSR,221-e2πifS,MfS,FSRrSR,1rSR,21-δSR |  |
|  | tSR,CR=1-δSR1-rSR,121-rSR,221-rSR,1rSR,21-δSR |  |
|  | FSR=πrSR,1rSR,21-δSR1-rSR,1rSR,21-δSR |  |
|  | fSR=fS,FSR2FSR |  |

With the parameters from Table 2 we get

|  | tSR,CR2≈1.3×10-2 and tSR,SB2≈6.4×10-3 |  |
| --- | --- | --- |

For the green output light we get:

|  | TfGreenSHG,ϕ(s)=21+s2πfSR |  |
| --- | --- | --- |
|  | TfGreenSHG,δl(s)=-4FSRrSR,1rSR,21-δSRkπ1+s2πfSR |  |

with

|  | fSR=fSRrSR,1rSR,21-δSR |  |
| --- | --- | --- |

## OPO

The OPO is locked in reflection using the green light. The locking signals can be written as

|  | TREFL,fGreenOPO,ϕ(s)=gOG,REFLs2πfO,FSR |  |
| --- | --- | --- |
|  | TREFL,fGreenOPO,δl(s)=gOG,REFL2kπ1+s2πfOG |  |

with

|  | gOG,REFL=-4J0(ΓO,M)J1(ΓO,M)FOG1-rOG,121-rOG,1rOG,2rOG,1rOG,2 |  |
| --- | --- | --- |
|  | FOG=πrOG,1rOG,21-rOG,1rOG,2 |  |
|  | fOG=fO,FSR2FOG |  |

The wave vector in green is twice the one in red. This has been accounted for by an explicit factor of 2 in the equation for the length error signal.

## Squeezed Field

The squeezed field sees first the cavity pole of the green light. After it gets converted to red light it will also see the cavity pole of the red light. We write

|  | TfMainNL,ϕs=12 11+s2πfOG11+s2πfOR |  |
| --- | --- | --- |

where the factor one half is due to the conversion. The pole frequency is at

|  | fOR=fO,FSR2FOR |  |
| --- | --- | --- |

A length change in the OPO will change both the green light before it is converted and the resulting red light. The two effects add:

|  | TfMainNL,δls=-2FOGkπrOG,1rOG,211+s2πfOG11+s2πfOR -2FORkπrOR,1rOR,211+s2πfOR |  |
| --- | --- | --- |

In the first term there is a factor of one half from the conversion and a factor of two from the wave vector which cancel.

## Coherent Locking

We write the auxiliary fields reflected from the cavity as

|  | EREFL,aux=EfAuxeiωauxt+iϕaux+EfAux2e-iωauxt+iϕaux2 |  |
| --- | --- | --- |

The prompt reflected field is given by

|  | EfAux=-rOR,1+eiφtOR,12eiφ-rOR,1rOR,2 |  |
| --- | --- | --- |

whereas the auxiliary field at fAux2 is given by

|  | EfAux2=e-iφtOR,12e-iφ-rOR,1rOR,2eiφ-rOR,1rOR,2 |  |
| --- | --- | --- |

Measuring the power using a photodetector and demodulating by cos2ωauxt+ϕM-π2, we get

|  | Vϕs=2EfAuxEfAux2sinϕaux2-ϕaux+ϕM |  |
| --- | --- | --- |

Locking at the zero crossing by setting ϕM=ϕaux-ϕaux2, we see that only phase variations will remain to produce an error signal. Phase variations in the coherent locking can be due to phase variations on the green pump light, due to phase variations on the auxiliary laser, due to length variations in the OPO and due to the amplitude variations of the auxiliary laser or the green pump light:

|  | TREFL,∆fAuxOPO,NL,ϕs=TREFL,fAux2OPO,NL,ϕs=HOR,2s,-φHOG,2s,0 |  |
| --- | --- | --- |
|  | TREFL,∆fAuxOPO,ϕs=TREFL,fAux2OPO,ϕ(s)-HOR,2s,φ |  |
|  | TREFL,∆fAuxOPO,δls=TREFL,fAux2OPO,δls-γOR,δlφHOR,1s,φ |  |
|  | TREFL,∆fAuxOPO,NL,ΔAs=TREFL,fAux2OPO,NL,ΔAs=γOR,ΔA(-φ)sωO,FSRHOR,3(s,-φ)HOG,2s,0 |  |
|  | TREFL,∆fAuxOPO,ΔAs=TREFL,fAux2OPO,ΔAs-γOR,ΔAφHOR,3s,φ |  |

with

|  | TREFL,fAux2OPO,ϕs=-HOR,2s,-φHOR,2s,φ+γOR,ΔA(-φ)αOR,Γ(φ)sω0,FSR2HOR,3s,-φHOR,3(s,φ) |  |
| --- | --- | --- |
|  | TREFL,fAux2OPO,δls=+γOR,δl-φHOR,1s,-φ+2γOG,δl0HOR,2s,-φHOG,1s,0-γOR,δlφHOR,2s,-φHOR,1s,φ+γOR,ΔA(-φ)αOR,δl(φ)sωO,FSRHOR,3(s,-φ)HOR,3(s,φ) |  |
|  | TREFL,fAux2OPO,ΔAs=+γOR,ΔA-φsωFSRHOR,3s,-φHOR,2s,φ-γOR,ΔA(φ)sωO,FSRHOR,2s,-φHOR,3(s,φ) |  |

The transfer function TREFL,fAux2OPO,NL,ϕs has a single term arising from phase fluctuations on fGreen directly causing phase fluctuations on fAux2. The transfer function TREFL,fAux2OPO,ϕs has two terms arising from (i) a phase shift on fAux directly causing a phase shift on fAux2, and (ii) from a phase shift on fAux causing an amplitude modulation on fAux  and hence fAux2 which transfers back into a phase shift on fAux2. The transfer function TREFL,fAux2OPO,δls has four terms arising from (i) a length variation causing a phase shift on fAux2, (ii) from a length variation causing a phase shift on fGreen and hence fAux2, (iii) from a length variation causing a phase shift on fAux and hence fAux2, and (iv) from length variations on fAux causing am amplitude modulation on fAux  and hence fAux2 which transfers back into a phase shift on fAux2. The transfer function TREFL,fAux2OPO,NL,ΔAs only has a single term coming from amplitude variations on fGreen causing amplitude variations on ffAux2 which translate into phase variations. The transfer function TREFL,fAux2OPO,ΔAs has two terms arising from (i) an amplitude variations on fAux causing amplitude variations on fAux2 which translate into phase variations, and (ii) an amplitude variations on fAux translating into phase variations on fAux and hence fAux2. The prompt reflected field fAux gets contributions to its phase variations from input phase variations, from OPO length variations and from input amplitude variations.

For the auxiliary field, EfAux, entering the OPO we get

|  | γδl(φ)γδl(φ≈116 and γΔA(φ)γΔA(φ)≈122 |  |
| --- | --- | --- |

Or, in other words these terms make a much smaller contribution than the ones carried by EfAux2; and thus we have neglected them.

## Local Oscillator

We write the auxiliary fields transmitted through the OPO cavity as

|  | ETRANS,aux=EfAuxeiωauxt+iϕaux+EfAux2e-iωauxt+iϕaux2 |  |
| --- | --- | --- |

Propagating this field to the anti-symmetric port yields

|  | EAS=eiϕinjETRANS,aux+ELO |  |
| --- | --- | --- |

where ϕinj is due to the path length variations in the squeezer injection path, and the local oscillator field is given by

|  | ELO=ELOeiϕLO |  |
| --- | --- | --- |

Demodulating by cosωauxt+ϕM yields

|  | Wϕs=+ 2ELOEfAuxcosϕaux+ϕinj-ϕLO-ϕM+2ELOEfAux2cosϕaux2+ϕinj-ϕLO+ϕM |  |
| --- | --- | --- |

Setting ϕM,I=ϕaux-ϕaux22 and ϕM,Q=ϕaux-ϕaux2+π2, one gets

|  | WIϕs=2ELOEfAux+EfAux2cosϕaux+ϕaux22+ϕinj-ϕLO |  |
| --- | --- | --- |
|  | WQϕs=2ELOEfAux-EfAux2sinϕaux+ϕaux22+ϕinj-ϕLO |  |

respectively. The non-linear conversion in the OPO will set ϕaux2=ϕgreen-ϕaux, and hence the in-phase photodetector at the anti-symmetric port is sensitive to

|  | ϕAS=ϕgreen2+ϕinj-ϕLO+aϕaux-ϕgreen2 |  |
| --- | --- | --- |

with a the field amplitude asymmetry

|  | a=EfAux-EfAux2EfAux+EfAux2 |  |
| --- | --- | --- |

Again, because we are locked to a zero point, only phase fluctuations enter the measured error signal to first order. Phase variations in the local oscillator locking signal can be due to phase variations on the local oscillator, due to phase variations on the green pump light, due to phase variations on the auxiliary laser, due to length variations in the OPO and due to the amplitude variations of the auxiliary laser or the green pump light:

|  | TLOϕs=-1 |  |
| --- | --- | --- |
|  | TTRANS,LOOPO,NL,ϕs=1-a2 TREFL,fAuxOPO,NL,ϕs |  |
|  | TTRANS,LOOPO,ϕs=1-a2TREFL,fAux2OPO,ϕ+1+a2HOR,2(s,φ) |  |
|  | TTRANS,LOOPO,δls=1-a2TREFL,fAux2OPO,δls+1+a2γOR,δlφHOR,1s,φ |  |
|  | TTRANS,LOOPO,NL,ΔAs=1-a2TREFL,fAux2OPO,NL,ΔAs |  |
|  | TTRANS,LOOPO,ΔAs=1-a2TREFL,LOOPO,ΔAs+1+a2γOR,ΔA(φ)sωO,FSRHOR,3(s,φ) |  |

The two contributions from each of the sidebands, fAux and fAux2, have weighting factors 1+a2 and 1-a2, respectively. When the sidebands are equal in strength, a=0, any direct contribution from ϕaux vanishes. In order to be sensitive to the green light phase the amplitude of the fAux2 sideband cannot be zero, i.e. a<1.

The phase of the local oscillator field is related to the phase of the PSL laser by an interferometer transfer function which consists of the mode cleaner pole and the double cavity pole

|  | TfPSLIFO,ϕs= 11+s2πfMC11+s2πfIFO |  |
| --- | --- | --- |

# Feedback Compensation Networks

## Laser Locking

The laser frequency actuator is a composite between a modulator, a PZT and a heater. The modulator or Pockels cell makes fast feedback up to ~1 MHz possible. The PZT dominates in the low frequency band below a couple of kHz and has a large range. The heater is used at very low frequency below ~1 Hz. Its error signal is usually taken from the PZT control voltage. For this analysis we can neglect the temperature servo all together.

The two actuators can be expressed as

|  | A0,1=2πα0,1s |  |
| --- | --- | --- |
|  | A0,2=α0,2 |  |

where α0,2 is the response for the modulator in rad/V and where α0,1 is the response of the PZT in Hz/V. We write the feedback compensation as

|  | G0s=g0,0G0,1sA0,1+G0,2sA0,2 |  |
| --- | --- | --- |

with

|  | G0,1s=g0,111+s2πf0,P611+s2πf0,P711+s2πf0,P8 |  |
| --- | --- | --- |
|  | G0,2s=g0,2s+2πf0,Z11+s2πf0,P11+s2πf0,Z21+s2πf0,P21+s2πf0,Z31+s2πf0,P3×1+s2πf0,Z41+s2πf0,P41+s2πf0,Z51+s2πf0,P5 |  |

The crossover between the PZT path and the Pockels cell path will be around 3 kHz, whereas the unity gain frequency will be near 500 kHz. The actuators and transfer functions of the auxiliary laser are identical and can be deduced from the above equations by replacing each "0" with a "1".

## SHG and OPO length

### SHG

For the photodetector to lock the SHG we are using

|  | S4s=η4PMainηQE,4εRZ411+s2πf4 |  |
| --- | --- | --- |

where η4, PMain, Γ4, ηQE,4, εR, Z4, and f4 denote the fractional power on the photodetector, the power of the main laser, the modulation depth, the photodetector quantum efficiency, the photodetector responsitivity in A/W, the photodetector transimpedance gain in Ω, and the pole due to the tuned circuit, respectively. The actuator is a PZT with a response of

|  | A4(s)=α4 |  |
| --- | --- | --- |

where α4 is the displacement coefficient in m/V. We write the electronics transfer function G4(s) as

|  | G4s=g41+s2πf4,Z11+s2πf4,P111+s2πf4,P2 |  |
| --- | --- | --- |

with zeroes and poles at f4,Z1,  f4,P1 and f4,P2, respectively. The gain g4 is adjusted to yield a unity gain frequency of about 1 kHz.

### OPO

The OPO feedback is similar to the SHG servo. For the photodetector we write

|  | S5(s)=η5PMainηQE,5εR2Z511+s2πf5 |  |
| --- | --- | --- |

where η5, PMain, Γ5, ηQE,5, εG, Z5, and f5 denote the fractional power on the photodetector, the power of the main laser, the modulation depth, the photodetector quantum efficiency, the photodetector responsitivity in A/W, the photodetector transimpedance gain in Ω, and the pole due to the tuned circuit, respectively. The photodetector responsitivity in green is half the one in red. The actuator is a PZT with a response of

|  | A5(s)=α5 |  |
| --- | --- | --- |

where α5 is the displacement coefficient in m/V. We write the electronics transfer function G5(s) as

|  | G5s=g51+s2πf5,Z11+s2πf5,P111+s2πf5,P2 |  |
| --- | --- | --- |

with zeroes and poles at f5,Z1,  f5,P1 and f5,P2, respectively. The gain g5 is adjusted to yield a unity gain frequency of about 1 kHz.

## Coherent Locking

The photodetector for the coherent locking can be described by

|  | S2(s)=2η2ηREFLPauxηQE,2εRZ211+s2πf2 |  |
| --- | --- | --- |

with f2 the bandwidth of the detector, and with

|  | ηREFL=ηREFL,auxηREFL,aux2 |  |
| --- | --- | --- |
|  | ηREFL,aux=rOR,2-rOR,1e-2πifauxfO,FSRe-2πifauxfO,FSR-rOR,1rOR,2≈1 |  |
|  | ηREFLS,aux2=ηOPOPGreenrOR,2tOR,121-2rOR,1rOR,2cos(2πfauxfO,FSR)+rOR,12rOR,22× tOG,11-rOG,1rOG,2 |  |

The coherent locking signal is fed back to the additive offset of the auxiliary laser as well as a PZT in the path from the auxiliary laser to the OPO. The PZT and additive offset actuators can be expressed as

|  | A2,1=kα2,1 and A2,2=α2,2 |  |
| --- | --- | --- |

For a servo compensation filter we use

|  | G2s=G2,0sG2,1sA2,1+G2,2sA2,2 |  |
| --- | --- | --- |

with

|  | G2,0s=g2,01+s2πf2,Z11+s2πf2,P11+s2πf2,Z21+s2πf2,P211+s2πf2,P3 |  |
| --- | --- | --- |
|  | G2,1s=g2,111+s2πf2,P4 |  |
|  | G2,2s=g2,2s1+s2πf2,P5s1+s2πf2,P6 |  |

Adjusting the gain in the PZT path we can set the crossover frequency around a few kHz. The overall gain determines the bandwidth of the servo and can be set as high as about 100 kHz. This is limited by the bandwidth of the laser locking servos since we are running out of actuator gain near its unity gain frequency. The additive offset path is AC-coupled to limit the amount of low frequency feedback range needed in this path. The overall open loop transfer function has two boost stages to increase low frequency suppression.

## Local Oscillator

The photodetector for the local oscillator locking can be described by

|  | S3(s)=2ηAux,3ηTRANS,aux+ηTRANS,aux2PAuxηLO,3PLOηQE,3εRZ311+s2πf3 |  |
| --- | --- | --- |

with f3 the bandwidth of the detector, and with

|  | ηTRANS,aux=tOR,1tOR,2e2πifauxfO,FSR-r1r2 |  |
| --- | --- | --- |
|  | ηTRANS,aux2=ηOPOPGreentOR,1tOR,21-2rOR,1rOR,2cos(2πfauxfO,FSR)+rOR,12rOR,22× tOG,11-rOG,1rOG,2 |  |

The field amplitude asymmetry the becomes

|  | α=ηTRANS,aux-ηTRANS,aux2ηTRANS,aux+ηTRANS,aux2 |  |
| --- | --- | --- |

The local oscillator locking signal is fed back to the additive offset of the main laser as well as a PZT in the path of the green light. Alternatively, the PZT can be placed in the fiber path, or altogether eliminated by feeding back to a VCO which feeds an AOM in the fiber path.

The PZT and additive offset actuators can be expressed as

|  | A3,1=kα3,1 and A3,2=α3,2 |  |
| --- | --- | --- |

For a servo compensation filter is identical to the one used for the coherent locking

|  | G3s=G3,0sG3,1sA3,1+G3,2sA3,2 |  |
| --- | --- | --- |

with

|  | G3,0s=g3,01+s2πf3,Z11+s2πf3,P11+s2πf3,Z21+s2πf3,P211+s2πf3,P3 |  |
| --- | --- | --- |
|  | G3,1s=g3,111+s2πf3,P4 |  |
|  | G3,2s=g3,2s1+s2πf3,P5s1+s2πf3,P6 |  |

Adjusting the gain in the PZT path we can set the crossover frequency around a few kHz. The overall gain determines the bandwidth of the servo and can be set as high as about 100 kHz. This is limited by the bandwidth of the laser locking servos since we are running out of actuator gain near its unity gain frequency. The additive offset path is AC-coupled to limit the amount of low frequency feedback range needed in this path. The overall open loop transfer function has two boost stages to increase low frequency suppression.

Since the local oscillator servo effectively locks the main laser to the PSL through the DC offset field at the antisymmetric port, one might wonder if the laser locking of the main laser to the PSL through a fiber is still necessary. In principle, one could eliminate the laser locking loop, but in practice the main laser and the PSL have to be within a few 10s of kHz to make this work. Locking the main laser to the PSL through a fiber first will guarantee that the laser runs at the correct frequency. The local oscillator network then only needs to lock the phase. Feeding back using an additive offset to the main laser also has the advantage to greatly increase the bandwidth and the low frequency suppression. This almost certainly will make a fiber stabilization system unnecessary.

# Noise Couplings and Requirements

## Acoustic Couplings

### Back Scattering

Back scattering has been calculated for the output mode cleaner (OMC) in T060303-v1. We are following the same derivation. Using Eqn. (3) of the above paper, we require

|  | xSCRr<10-17 m/√Hz, f>10 Hz, |  |
| --- | --- | --- |

where xSC is the motion of the scattering surface along the direction of the beam propagation vector and Rr is the power reflectivity, into the interferometer TEM00 mode, of the scattering surface.

We assume that the fraction of light at the AS port deflected into the squeezer path by the thin-film polarizer is 10-3, and that the rejection of the in-vacuum Faraday isolator in the squeezer path is 10-3 as well. Any light reflected back into the squeezer path will then go straight back into the interferometer. Looking first at the output window on HAM4 and taking the reflectivity value from Table 1 in T060303, we get

|  | Rr≈3×10-12 and xSC<4×10-6. |  |
| --- | --- | --- |

Looking at Figure 4 in T060303 we can see that this is easily fulfilled.

Once on the squeezer breadboard the largest back scattering is due to the OPO itself. Since the OMC and the OPO are very similarly constructed cavities, we assume their reflectivities will be similar too. The OMC reflectivity is estimated to be 6×10-10, whereas high quality steering optics will be in the range from 2×10-13 (0.7 mm beam waist) to 6×10-14 (4.5 mm beam waist).

Figure 2 shows a typical spectrum of the motion of an optical table in the LVEA. We are approximating the spectra with

|  | xSC≈10-6f2 m∙Hz3/2 for an optics table straight on the floor,10-8f2 m∙Hz3/2 for an accoustically isolated table. |  |
| --- | --- | --- |

Substituting back into the requirement and including the factor 10-6 from the thin-film polarizer and the in-vacuum Faraday isolator, we get a requirement on the overall back-scattering reflectivity coefficient of the squeezer setup

|  | RSQZ≈10-4f100 Hz2 for an optics table straight on the floor,10-2f100 Hz2 for an accoustically isolated table. |  |
| --- | --- | --- |

For the H1 squeezer experiment we set f = 100 Hz and use the first requirement. This sets the requirement for RSQZ<10-8. The estimate for the OPO reflectivity is well below that and so are all the steering optics.

For advanced LIGO the lowest frequency of interest is 10 Hz. This will require the placement of the squeezer table into an acoustically isolated enclosure. Special care needs to be taken to avoid increasing the acoustically induced path fluctuations by introducing large resonances in the optical components on the squeezer table.

Figure 2: Motion of an optics table in the LVEA. IOT7 and ISCT7 are standard optics table with fixed legs resting on the ground. ISCT7 and ISCT10 are located inside an acoustic enclosure. Above 100 Hz the displacement spectra for ISCT7 and 10 are limited by the noise of the accelerometers.

### Beam Jitter

Beam jitter and static misalignment can reduce the power coupled through the output mode cleaner. This will look like a straight power loss and can be added to the power budget. To get a 10% loss in transmitted power the static misalignment needs to be as large as a third of the divergence angle. Static misalignment will be smaller than this—certainly, if an auto-alignment system is used.

Beam jitter can also couple with a static misalignment to generate a longitudinal phase shift. We write

|  | Δφ≈α0 Δα2 |  |
| --- | --- | --- |

These fluctuations will be suppressed by both the local oscillator feedback servo and a possible auto-alignment system. Hence, they are unlikely a significant contribution to the error budget.

## Squeezed Beam Requirements

### Squeezer Level

We require 3 dB of squeezing at the interferometer antisymmetric port. To allow for significant optical losses in the injection path and the interferometer we require the squeezer itself to demonstrate 6 dB of squeezing.

### Phase Noise

If the squeezer produces a squeezed state of 6 dB and 14 dB of anti-squeezing, the expected squeezed state after 25% of additional optical losses in the interferometer is 3.5 dB. This corresponds to a combined total optical loss of about 40%. Figure 3 shows that with 60% propagation efficiency 3.3 degrees of phase fluctuations will degrade the squeezed state to about 3 dB. We therefore set a requirement on the total phase noise of the squeezed vacuum state relative to the DC locking field at the antisymmetric port of 50 mrad rms.

Figure 3: Expected squeezing level from 0 dB to 10 dB as function of the propagation efficiency and the RMS fluctuation angle.

## Losses

The relevant optical losses from the squeezer output into the interferometer and out onto the photodetector at the antisymmetric port are shown in Table 1.

Table 1: Optical losses from the squeezer to the antisymmetric port photodetector.

| Component | Transmission | Uncertainty | Reference |
| --- | --- | --- | --- |
| Vacuum window | 0.998 | 0.002 |  |
| 1st Faraday isolator | 0.95 | 0.02 |  |
| 2nd Faraday isolator (side injection) | 0.95 | 0.02 |  |
| Reflection Michelson/arm cavities | 0.98 | 0.01 |  |
| 2nd Faraday isolator (through) | 0.94 | 0.02 |  |
| Septum window | 0. 998 | 0.2 |  |
| OMC mode matching | 0. 74 | 0.002 |  |
| OMC cavity losses | 0.967 | 0.005 |  |
| DC photodetector QE | 0.98 | 0.01 |  |
| Total | 0.745 | 0.043 |  |

## Modulation Sidebands

The RF modulation sidebands are large at the antisymmetric port. They will be rejected by the output mode cleaner which has a power transmission coefficient of TSB=0.9% at 25 MHz offset. The free-spectral-range is 287 MHz and the input and output coupler have a 8000 ppm transmission. With 8 W power into the H1 interferometer about 10 mW carrier light gets passed through the output mode cleaner. About 0.4 mW of this light is due to contrast defect whereas the remainder is due to the length offset. A scan of the OMC length revealed that the RF sideband power in the OMC mode is about the same for each sideband as for the carrier. The phase noise contribution has 2 terms: a term due to the sidebands beating against the carrier in the contrast defect and a term due to the imbalance of the sidebands.

If we write

|  | PCD=PCRη |  |
| --- | --- | --- |

with PCD the power due to the contrast defect and PCR the total carrier power after the OMC. For the sidebands we can write

|  | PSB=PSB++PSB- |  |
| --- | --- | --- |
|  | ΔPSB=PSB+-PSB-=2ε PSB |  |

The rms phase fluctuations then becomes

|  | Γrms=TSBPSBPCR1η+ε2η-1η |  |
| --- | --- | --- |

Figure 4 shows the phase noise due to the RF sidebands as function of η. The current value for η at the H1 interferometer is around 25 which yields a rms phase noise of about 1.5° which corresponds to about 25 mrad and is below the requirement.

Figure 4: Phase noise due to the RF sidebands at the AS port after the OMC. The curve is given as a function of the ratio of the total carrier power over the carrier power due to contrast defect. The red curve is for ε=0, whereas the blue curve is for a very large ϵ=0.2.

## Noise Couplings

Using the servo model from the previous section we are looking at the noise couplings into the phase angle of the squeezed field. Using worst case spectra we can then propagate the corresponding noise terms into phase fluctuations of the squeezed field. We have established a requirement on the phase fluctuations of the squeezed field of 50 mrad rms. We translate this into a requirement on each of the technical noise terms of no more than 5 mrad rms. All calculations in this section are made with an auxiliary laser field asymmetry of α=0.33.

### Laser Noise

We include frequency noise from the PSL, the main laser and the auxiliary laser, as well as intensity noise from the main and auxiliary laser. Figure 5 shows the magnitude transfer functions for these couplings. Due to our high bandwidth feedback compensation networks laser frequency noise is greatly suppressed. The coupling of laser intensity noise is intrinsically small, since its only direct coupling is due to the coherent field which is passed through the OPO off-resonance.

Figure 5: Transfer functions of laser noise into the squeezer phase noise.

We take the free running laser frequency noise from Figure 5 in T990025-00 and approximate a worst case spectrum with the following equation

|  | NΔffree(f)=200 HzHzfor f≤100 Hz200 HzHz×100 Hzf for f>100 Hz |  |
| --- | --- | --- |

The stabilized laser frequency noise is used for the PSL and is taken from Figure 3 in T000031-00. We use

|  | NΔfstab(f)=0.02 HzHzfor f≤1 MHz0.02 HzHz×1 MHzf for f>1 MHz |  |
| --- | --- | --- |

Figure 6: Squeezer phase noise due to laser noise.

The free running intensity noise is taken from Figure 6 in T990025-00. We are limiting the high frequency roll-off to the level of shot noise at 1 mW—again, a worst case assumption. We use the following equation for the relative intensity noise

|  | NΔAfreef=10-5 Hzfor f≤1 kHz10-5 Hz×1 kHzffor f>1 kHz and f≤100 kHz10-7 Hzfor f>100 kHz |  |
| --- | --- | --- |

Figure 6 shows these noise sources propagated to the phase angle of the squeezed field. The requirement on the rms phase noise is easily met.

### Path Length Fluctuations

We investigate path length fluctuations in the input path due to fiber noise. Path length fluctuations on the laser table itself due to flexing of the optical table or the non-rigidity of optical mounts are added to the main laser beam path, to path of the auxiliary laser, to the path of the green light and to the path towards the interferometer. Figure 7 shows the magnitude transfer functions for these couplings. Again, these fluctuations are greatly suppressed due to our servos.

Figure 7: Transfer functions of path length fluctuations into squeezer phase noise.

For the input to the path length fluctuations we are using Figure 2Figure 2: Motion of an optics table in the LVEA. IOT7 and ISCT7 are standard optics table with fixed legs resting on the ground. ISCT7 and ISCT10 are located inside an acoustic enclosure. Above 100 Hz the displacement spectra for ISCT7 and 10 are limited by the noise of the accelerometers., and approximate with

|  | NΔxf=10-10mHz×100 Hzf2 |  |
| --- | --- | --- |

Fiber noise was measured in Figure 2 of T0900376-v1. It is no worse than the free running frequency noise of a laser and we can just use its spectra.

Multiplying the noise spectra with the transfer functions we get the curves in Figure 8. Again, the requirement on the rms phase noise of the squeezed field is met easily.

Figure 8: Squeezer phase noise due to path length variations.

### Displacement Noise

Displacement noise couplings of both the SHG and the OPO cavity length are shown in Figure 9. There is a significant difference between the two: the displacement noise of the SHG is suppressed by the local oscillator feedback, whereas the one of the OPO is not. Displacement noise will affect the coherent fields at the antisymmetric port differently from the squeezed field. Therefore zeroing the effect on the local oscillator error signal will not suppress the effect on the squeezed field. On the other hand displacement noise of the SHG will add phase noise to the green light which then can be suppressed with the local oscillator servo.

We use the same length fluctuations as in the equation above. These are the length fluctuations as measured on the top of an optical table and therefore will overestimate the relative displacement noise of a rigid structure. Even with these assumptions we still meet the requirement. The curves are presented in Figure 10.

Figure 9: Transfer functions of cavity displacement noise into squeezer phase noise.

Figure 10: Squeezer phase noise due to cavity displacement noise.

### Shot Noise

Shot noise at the output of the nth photodetector can be written as

|  | Nnshotf=2q ηQE,nεRPn Zn1for n=0, 111+s2πfnfor n=2, 3, 412 11+s2πfnfor n=5 |  |
| --- | --- | --- |

where q=1.602176487(40)×10-19 is the elementary charge, Pn denotes the total light power incident on the photodetector,  ηQE,n denotes the quantum efficiency, εR denotes the photodetector responsitivity in A/W, Zn denotes the photodetector transimpedance gain in Ω, and fn denotes the pole due to the tuned circuit, respectively. Photodetector 5 measures green light and we include a factor of 12 for the reduction of the photodetector responsitivity.

For the incident power on each detector we estimate

| P0 | ηPSL,0PPSL+ηMain,0PMain | 1.5 mW |  |
| --- | --- | --- | --- |
| P1 | ηMain,1PMain+ηAux,1PAux | 1.5 mW |  |
| P2 | η2ηREFL,auxPAux | 0.5 mW |  |
| P3 | ηLO,3PLO+PCD | 1.0 mW |  |
| P4 | η4tSR,CR2PMain | 1.3 mW |  |
| P5 | η5PMain | 1.0 mW |  |

where we neglect the power in fAux2 on photodetector 2 and where we neglected the power in fAux and fAux2 on photodetector 3. PCD denotes the power in all other modes other an TEM00 of the carrier local oscillator field. It includes higher order modes as well as the RF sidebands. In reality, PCD may be ten times larger than PLO.

Figure 11: Transfer functions of shot noise contributions into squeezer phase angle.

Figure 11 shows the magnitude transfer functions for coupling shot noise into squeezer phase angle. The most critical coupling is due to shot noise in the coherent locking. This is mostly due to the fact the signal is rather small in reflection of the OPO and hence on photodetector S2. Figure 12 shows the resulting squeezer phase noise when we add the shot noise from the above power estimations. The resulting rms noise levels are well below our requirement.

Figure 12: Squeezer phase noise due to shot noise on the photodetectors.

# Fabry-Perot Resonators

## Cavity Parameters

### SHG

The SHG cavity is a linear configuration. Parameters are listed in Table 2.

Table 2: Optical parameters for the SHG cavity.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| LS | Physical length | 47.35 | mm |
| lS | Effective length (10 mm PPKTP) | 54.75 | mm |
| fS,FSR | Free-spectral-range | 2.7 | GHz |
| λR | Red wavelength | 1064 | nm |
| λG | Green wavelength | 531 | nm |
| rSG,12 | Green input mirror reflectivity | 0.01 |  |
| rSG,22 | Green output mirror reflectivity | 1.0 |  |
| rSR,12 | Red input mirror reflectivity | 0.9 |  |
| rSR,22 | Red output mirror reflectivity | 0.9992 |  |
| δSR | Red power loss (conversion coeff.) | 0.05 |  |
| FSG | Green Finesse | 1.1 |  |
| FSR | Red Finesse | 39.9 |  |
| fSR | Red cavity pole | 34.3 | MHz |
| fS,M | Modulation frequency | 35.5 | MHz |
| ΓS,M | Modulation depth | 0.01 |  |

### OPO

The OPO cavity is a bowtie configuration. Parameters are listed in Table 3.

Table 3: Optical parameters for the OPO cavity.

| Parameter | Parameter | Value | Unit |
| --- | --- | --- | --- |
| lO | Half length | 0.14 | m |
| fO,FSR | Free-spectral-range | 1.071 | GHz |
| rOG,12 | Green input mirror reflectivity | 0.7 |  |
| rOG,22 | Green output mirror reflectivity | 1.0 |  |
| rOR,12 | Red input mirror reflectivity | 0.99 |  |
| rOR,22 | Red output mirror reflectivity | 0.85 |  |
| FOG | Green Finesse | 17.6 |  |
| FOR | Red Finesse | 36.4 |  |
| fOG | Green cavity pole | 30.4 | MHz |
| fOR | Red cavity pole | 14.7 | MHz |
| fO,M | Modulation frequency | 71 | MHz |
| ΓO,M | Modulation depth | 0.01 |  |
| ηOPOPGreen | Conversion factor for coherent field | 0.0276 |  |

### Interferometer

Some relevant parameters of the LIGO interferometer are listed in Table 4.

Table 4: Optical parameters for the interferometer.

| Parameter | Parameter | Value | Unit |
| --- | --- | --- | --- |
| fMC | Mode cleaner pole | 4 | kHz |
| fIFO | Double cavity pole | 2 | Hz |

## Cavity Equations

### Pound-Drever-Hall Signals

Assuming the RF sidebands are well outside the cavity resonance and omitting the factor J0ΓJ1ΓP which describes the power and modulation depth of the incident beam, the cavity transfer function of a Pound-Drever-Hall reflection locking signal can be written as:

|  | TREFLδls=-4Fkπ1-r121-r1r2r1r211+sωcav |  |
| --- | --- | --- |

with F the cavity finesse, r1 and r2 the amplitude reflection coefficients of the input and output mirror, respectively, and ωcav the cavity pole angular frequency:

|  | F=πr1r21-r1r2 |  |
| --- | --- | --- |
|  | ωcav=πfFSRF |  |

and fFSR=c2L the free-spectral-range of the cavity. A similar equation can be written for phase fluctuations injected into the cavity:

|  | TREFLϕs=-4F1-r121-r1r2r1r2sωFSR |  |
| --- | --- | --- |

The equations for an error signal in transmission are practically the same, but have to include a factor from the sideband transmission amplitude.

|  | TTRANSδls=-4FkπtSB1-r121-r221-r1r2r1r211+sωcav |  |
| --- | --- | --- |
|  | tSB=1-r121-r221-e2πifMfFSRr1r2 |  |

with fM the modulation frequency. Again, we neglect contributions from the RF sidebands which are small as long as fM≫fcav. And for the laser phase noise:

|  | TTRANSϕs=-4FtSB1-r121-r221-r1r2r1r2sωFSR |  |
| --- | --- | --- |

For situations where the modulation frequency is close to the cavity pole one will experience an additional loss of signal. One might be tempted to estimate this factor with

|  | Tloss~ifMfcav+ifM |  |
| --- | --- | --- |

However, this can be quiet inaccurate for fM≲fcav, especially, in reflection when the rear mirror is a high reflector.

### Electric Fields

We are looking at an electric field entering a cavity. An off-resonant field can be characterized by its frequency shift away from resonance by a phase angle

|  | φ=2πffFSR |  |
| --- | --- | --- |

The field inside the cavity can always be decomposed into a phase-modulated and an amplitude-modulated component, ΓCAV and ΔACAV, respectively. If we make the ansatz that the input beam has a phase modulation Γ and a relative amplitude modulation ΔA, and if we denote a small length variation by δl, the components inside the cavity can be written as

|  | ΓCAV=γδl(φ)H1s,φδl+H2s,φΓ+ γΔA(φ)sfFSRH3(s,φ)ΔA |  |
| --- | --- | --- |
|  | ΔACAV=αδlφH3s,φδl+αΓφsfFSRH3s,φΓ+ H2(s,φ)ΔA |  |

For the fields propagating back to the input mirror the carrier field is

|  | ECAVcr=t1r2eiφ-r1r2EINcr |  |
| --- | --- | --- |

and the coefficients become

|  | γδl(φ)=-2k1-r1r2cosφ1-2r1r2cosφ+r12r22 |  |
| --- | --- | --- |
|  | γΔAφ=-r1r2sinφ1-2r1r2cosφ+r12r22 |  |
|  | αδl(φ)=2k γΔAφ |  |
|  | αΓφ=-γΔAφ |  |

The frequency dependency is contained in the transfer functions H1, H2 and H3. In the case of a detuned frequency, φ≠0, these transfer functions are more complicated than a single pole. However, we can still make a small frequency approximation using a single pole or a zero for frequency much below the cavity pole. In this case we write

|  | Hns,φ=1+sωnφif ωnφ>01if ωnφ=011+sωn(φ)if ωnφ<0 |  |
| --- | --- | --- |

where the pole/zero frequencies are given by

|  | ω1φ=-2fFSR 1-3r1r2cosφ+r12r222+cos2φ-r13r23cosφ1-r1r2cosφ-r12r222-cos2φ+r13r23cosφ |  |
| --- | --- | --- |
|  | ω2φ=-fFSR 1-2r1r2cosφ+r12r221-r1r2cosφ |  |
|  | ω3φ=-2fFSR 1-2r1r2cosφ+r12r223-2cosφr1r2-r12r22 |  |

When on resonance, φ=0,  we get the familiar

|  | γδl(0)=-2kr1r21-r1r2 |  |
| --- | --- | --- |
|  | γΔA(0)=αδl(0)=αΓ(0)=0 |  |
|  | ω10=-ωcav and ω20=-r1r2ωcav |  |

The transmitted field can be obtained from by replacing r2 with t2 in the numerator of the carrier field, i.e.,

|  | ETRANScr(φ)=t1t2eiφ-r1r2EINcr |  |
| --- | --- | --- |

and the above equations for amplitude and phase modulation transfer coefficients stay unchanged. In reflection one has to add the prompt reflected field and the transfer coefficients change significantly,

|  | EREFLcr(φ)=t11r2eiφ-r1r2-r1EINcr=r2-r1eiφeiφ-r1r2EINcr |  |
| --- | --- | --- |

and

|  | ΓREFL=γδl(φ)H1s,φδl+H2s,φΓ+ γΔA(φ)sfFSRH3(s,φ)ΔA |  |
| --- | --- | --- |
|  | ΔAREFL=αδlφH3s,φδl+αΓφsfFSRH3s,φΓ+ H2(s,φ)ΔA |  |

with

|  | γδl(φ)=-2kr2t12r11+r22cosφ-r21+r121-2r1r2cosφ+r12r22r12-2r1r2cosφ+r22 |  |
| --- | --- | --- |
|  | γΔAφ=r1r2t12t22sinφ1-2r1r2cosφ+r12r22r12-2r1r2cosφ+r22 |  |
|  | αδlφ=2k γΔAφ |  |
|  | αΓφ=-γΔAφ |  |

and

|  | ω1φ=-2fFSR 1-2r1r2cosφ+r12r22r21+r12-r11+r22cosφr2-r14r23-2-cos2φr12r2t22-r11+r221-r12r22cosφ |  |
| --- | --- | --- |
|  | ω2φ=fFSR 1-2r1r2cosφ+r12r22r12-2r1r2cosφ+r22r2t12r21+r12-r11+r22cosφ |  |
|  | ω3φ=2fFSR t221-2r1r2cosφ+r12r221-r223+3r12-r12r22+2r1r21+r22cosφ |  |

For a field on resonance we get

|  | EREFLcr(0)=r2-r11-r1r2EINcr |  |
| --- | --- | --- |
|  | γδl(0)=2kr2t121-r1r2r2-r1 |  |
|  | γΔA(0)=αδl(0)=αΓ(0)=0 |  |
|  | ω10=ωcavr1r2r2-r11-r1r22r1-r2+r12r2+r1r22-r13r22-r12r23+r14r23-r13r24 |  |
|  | ω20=ωcavr1r2r2-r1r2t12 |  |

### Non-Linear Process

The non-linear process in an OPO involving a seed field is given as Eout=ηOPOEseed*Epump, where η represents the conversion coefficient. Assuming that input fluctuations are small we can see that the relative amplitude modulation and the phase modulation of each input beam will just appear on the output field. However, the conversion factor is not independent of the field strength, and the phase modulation of the seed field acquires a minus sign at the output due to the fact that fout=fpump-fseed. We therefore write

|  | Γout=Γpump-Γseed and ΔAout=ΔApump+ΔAseed+… |  |
| --- | --- | --- |

Writing the seed and pump fields as

|  | Eseed=EauxeiΩt+2πifauxt+iϕaux and Epump=Egreene2iΩt+iϕgreen |  |
| --- | --- | --- |

the reflected output field can be written as

|  | Eout=Eaux2eiΩt-2πifauxt+i(ϕgreen-ϕaux) |  |
| --- | --- | --- |

with

|  | Eaux2=ηOPOrOR,2tOR,1e-2πifauxfO,FSR-rOR,1rOR,2 tOR,1e2πifauxfO,FSR-rOR,1rOR,2× tOG,11-rOG,1rOG,2EauxEgreen =ηOPOrOR,2tOR,121-2rOR,1rOR,2cos(2πfauxfO,FSR)+rOR,12rOR,22 ×tOG,11-rOG,1rOG,2EauxEgreen |  |
| --- | --- | --- |

where OR and OG represent the red and green OPO parameters, respectively.

In the case of the squeezed field generation we have no seed beam. The green pump beam then generates two correlated red photons for each green photon converted in the process. We write E1E2=ηEpump. Now setting E1≈E2 we see that the phase of the output field is half of the green input field.

For the SHG we can write Egreen=ηEred2 and we see that the phase and amplitude modulation of the red input field simply double at the output, Γgreen=2Γred and ΔAgreen=2ΔAred+…, respectively.

### Offset Locking

When two laser beams are interfered and locked with a frequency offset fS=f2-f1, the resulting error signal when demodulated by cos2πfSt+ϕM+π2 is given by

|  | Vϕs=ϕ2-ϕ1+ϕM |  |
| --- | --- | --- |

We are omitting the factor 2P1P2 which describes the phase-to-power conversion, with P1 and P2 the power in each beam, respectively.

# Compensation Design

The parameters of the feedback compensation network are given in the following tables.

## Laser Locking

Table 5: Parameters for the laser locking servos.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| α0,2 | Pockels modulation index | 15 | mrad/V |
| α1,2 | Pockels modulation index | 15 | mrad/V |
| α0,1 | Main laser PZT input | 300 | kHz/V |
| α1,1 | Aux. laser PZT input | 300 | kHz/V |
| ηQE,0 | Photodetector quantum efficiency | 0.95 |  |
| ηQE,1 | Photodetector quantum efficiency | 0.95 |  |
| εR | Red photodetector responsitivity | 0.858 | A/W |
| Z0 | Photodetector transimpedance | 40 | kΩ |
| Z1 | Photodetector transimpedance | 40 | kΩ |
| PPSL | Power from the PSL | 1 | mW |
| PMain | Main laser power | 500 | mW |
| PAux | Auxiliary laser power | 50 | mW |
| fAux | Auxiliary laser frequency offset | 29.507 | MHz |
| ηPSL,0 | Attenuation factor of PSL beam | 0.5 |  |
| ηMain,0 | Attenuation factor of main beam | 0.002 |  |
| ηMain,1 | Attenuation factor of main beam | 0.002 |  |
| ηAux,1 | Attenuation factor of aux. beam | 0.01 |  |
| g0 | Overall gain | 0.022 |  |
| g0,1 | Fast path gain | 32000 |  |
| g0,2 | Pockels cell path | 240 |  |
| f0,Z1 | Zero in electronics | 0 | kHz |
| f0,Z2 | Zero in electronics | 77.6 | kHz |
| f0,Z3 | Zero in electronics | 97.5+205i | kHz |
| f0,Z4 | Zero in electronics | 97.5‒205i | kHz |
| f0,Z5 | Zero in electronics | 10000 | kHz |
| f0,P1 | Pole in electronics | 1 | kHz |
| f0,P2 | Pole in electronics | 2.13 | kHz |
| f0,P3 | Pole in electronics | 3.1 | kHz |
| f0,P4 | Pole in electronics | 150 | kHz |
| f0,P5 | Pole in electronics | 100 | kHz |
| f0,P6 | Pole in electronics | 10 | kHz |
| f0,P7 | Pole in electronics | 23 | kHz |
| f0,P8 | Pole in electronics | 0.01 | kHz |

## SHG and OPO

Table 6: Parameters for the SHG length servo.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| η4 | Fraction of detected power | 0.2 |  |
| ηQE,4 | Photodetector quantum efficiency | 0.95 |  |
| Z4 | Photodetector transimpedance | 1000 | Ω |
| f4 | Photodetector pole | 5 | MHz |
| α4 | PZT displacement coefficient | 5 | nm/V |
| g4 | Servo gain | 31,600,000 |  |
| f4,Z1 | Zero in electronics | 500 | Hz |
| f4,P1 | Pole in electronics | 1 | Hz |
| f4,P2 | Pole in electronics | 1 | Hz |

Table 7: Parameters for the OPO length servo.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| η5 | Fraction of detected power | 0.002 |  |
| ηQE,5 | Photodetector quantum efficiency | 0.95 |  |
| Z5 | Photodetector transimpedance | 200 | Ω |
| f5 | Photodetector pole | 10 | MHz |
| α5 | PZT displacement coefficient | 5 | nm/V |
| g5 | Servo gain | 10,000,000 |  |
| f5,Z1 | Zero in electronics | 500 | Hz |
| f5,P1 | Pole in electronics | 1 | Hz |
| f5,P2 | Pole in electronics | 1 | Hz |

## Coherent Locking and Local Oscillator

Table 8: Parameters for the coherent locking servo.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| ηQE,2 | Photodetector quantum efficiency | 0.95 |  |
| Z2 | Photodetector transimpedance | 1000 | Ω |
| η2 | Attenuation factor in path to S2 | 0.01 |  |
| f2 | Photodetector pole | 10 | MHz |
| α2,1 | PZT displacement coefficient | 5 | nm/V |
| α2,2 | Auxiliary laser additive offset | 1 | V/V |
| g2,0 | Servo common gain | 320,000 |  |
| g2,1 | Servo PZT gain | 1,000,000 |  |
| g2,2 | Servo additive offset gain | 7.08 |  |
| f2,Z1 | Zero in common path electronics | 20 | kHz |
| f2,Z2 | Zero in common path electronics | 20 | kHz |
| f2,P1 | Pole in common path electronics | 1 | kHz |
| f2,P2 | Pole in common path electronics | 1 | kHz |
| f2,P3 | Pole in common path electronics | 40 | Hz |
| f2,P4 | Pole in PZT path electronics | 10 | Hz |
| f2,Z3 | Zero in additive offset path | 0 | Hz |
| f2,Z4 | Zero in additive offset path | 0 | Hz |
| f2,P5 | Pole in additive path electronics | 5 | Hz |
| f2,P6 | Pole in additive path electronics | 5 | Hz |

Table 9: Parameters for the local oscillator locking servo.

| Parameter | Description | Value | Unit |
| --- | --- | --- | --- |
| ηQE,3 | Photodetector quantum efficiency | 0.95 |  |
| Z3 | Photodetector transimpedance | 1000 | Ω |
| ηAux,3 | Attenuation factor in aux. path to S3 | 0.01 |  |
| ηLO,3 | Attenuation factor in LO path to S3 | 0.01 |  |
| PLO | Power in the LO | 10 | mW |
| f3 | Photodetector pole | 5 | MHz |
| α3,1 | PZT displacement coefficient | 5 | nm/V |
| α3,2 | Main laser additive offset | 1 | V/V |
| g3,0 | Servo common gain | 100,000 |  |
| g3,1 | Servo PZT gain | 1,000,000 |  |
| g3,2 | Servo additive offset gain | 7.08 |  |
| f3,Z1 | Zero in common path electronics | 20 | kHz |
| f3,Z2 | Zero in common path electronics | 20 | kHz |
| f3,P1 | Pole in common path electronics | 1 | kHz |
| f3,P2 | Pole in common path electronics | 1 | kHz |
| f3,P3 | Pole in common path electronics | 40 | Hz |
| f3,P4 | Pole in PZT path electronics | 10 | Hz |
| f3,Z3 | Zero in additive offset path | 0 | Hz |
| f3,Z4 | Zero in additive offset path | 0 | Hz |
| f3,P5 | Pole in additive path electronics | 5 | Hz |
| f3,P6 | Pole in additive path electronics | 5 | Hz |

The combined reflectivity of the OMC and the Faraday isolator has been measured to be Rr=3×10-8 (see LLO ilog from 1/28/2009, and corrected for the 50/50 beam splitter which was installed during the measurement).

LHO ilog from 8/4/2008

LHO ilog from 8/12/2008

LHO ilog from 8/4/2008

LHO ilog from 8/1/2008

LHO ilog from 7/8/2009