# LASER INTERFEROMETER GRAVITATIONAL WAVE OBSERVATORY

# -LIGO-

# CALIFORNIA INSTITUTE OF TECHNOLOGY

# MASSACHUSETTS INSTITUTE OF TECHNOLOGY

| Document Type Test Procedure and Results | DCC Number LIGO-T0900346-v2 | Date 16 July, 2009 |
| --- | --- | --- |
| STS-2 Seismometer Chassis Test Procedure |
| Ben Abbott |

Distribution of this draft: NSF reviewers, LIGO scientists

This is an internal working note of the LIGO Laboratory

California Institute of Technology	Massachusetts Institute of Technology

LIGO Project – MS 18-33	LIGO Project – MS 20B-145

Pasadena, CA 91125	Cambridge, MA 01239

Phone (626) 395-2129	Phone (617) 253-4824

Fax (626) 304-9834	Fax (617) 253-7014

E-mail: info@ligo.caltech.edu	E-mail: info@ligo.mit.edu

www:  http://www.ligo.caltech.edu/

LIGO-T0900346-v1

Performed by: ______________

Date: ______________

Board Serial Number: ___________

- Overview

The HEPI STS-2 Seismometer Interface Chassis provides power and control channels for a single STS-2 seismometer.  The chassis receives 6 differential signals from the seismometer that are sent as outputs to the AdL Anti-Alias Chassis for ultimate transmission to the ADCs.  A summary of functions for the STS-2 Seismometer Interface are:

  - DC power to the remotely located STS-2 seismometer

  - Receives 6 channels of differential signals from the STS-2 seismometer

  - Interfaces analog signals to the AdL Anti-Alias Chassis by differential interface

  - Provides local front panel switches and remote PCIX based control of STS-2 functions

The function of this procedure is to check each channel from its input to the respective output, test binary controls and to verify proper DC power consumption.

- Test Equipment

  - Power Supply capable of +/- 18 volts

  - Function generator (Stanford Research DS360 or the like)

  - Oscilloscope

- Preliminaries

  - Perform visual inspection on board to check for missing components or solder deficiencies

  - Before connecting the power to the chassis, set power supplies to +/- 18 Volts, and then turn them off. Connect the power supplies to the chassis under test at the back panel 3-pin power connector.

- DC Tests

  - Turn on the power supplies to the system under test and record the total current.  The specification assumes all inputs are not driven and the front panel switches are clicked down.

| Total Current | Specification | Observation |
| --- | --- | --- |
| +18V Supply | 110mA +/- 10mA |  |
| -18V Supply | 100mA +/- 10mA |  |
| Power LEDs | Lit with equal brightness |  |

LIGO-T0900346-v1

- Dynamic Tests

  - The following tests verify the proper gain, for each signal channel.  The test consists of applying a 100 Hz signal to test the channel gain

  - Using a function generator and an oscilloscope, enter 0.1V p-p (HI-Z) sine wave on the function generator and apply to the prescribed input and observe the amplitude at the designated output (either the Anti Alias signals on J7, the on-board testpoints, or the front panel BNC).  For differential outputs, use two scope probes in “differential mode” that is probe 1 minus probe 2.

STS-2 Response Data

| Input (+, -) | Output POINTs | NOMINAL mAG (100Hz) | MEAS. MAG. (100Hz) |
| --- | --- | --- | --- |
| 50 pin D-sub pin 3/36 or TP15/TP14 | STS2-X, J7 pins 4&12 or TP4&TP12 & X-MON BNC | Pins 4&12: 4V p-p +/- 0.5V BNC: 2v p-p +/-0.5V | Pins 4&12: BNC: |
| 50 pin D-sub pin 20/4 or TP15/TP14 | STS2-Y, J7 pins 5&13 or TP4&TP12 & Y-MON BNC | Pins 5&13: 4V p-p +/- 0.5V BNC: 2v p-p +/-0.5V | Pins 5&13: BNC: |
| 50 pin D-sub pin 37/21 or TP15/TP14 | STS2-Z, J7 pins 6&14 or TP4&TP12 & Z-MON BNC | Pins 6&14: 4V p-p +/- 0.5V BNC: 2v p-p +/-0.5V | Pins 6&14: BNC: |

  - Enter 10V p-p (HI-Z) on the function generator and use only the positive output relative to GND.  This requires referencing the common of the function generator to STS-2 GND.

| Input (+, -) | Output POINTs | NOMINAL mAG (100Hz) | MEAS. MAG. (100Hz) |
| --- | --- | --- | --- |
| 50 pin D-sub pin 41/GND | MASSPOS-U J7 pins 1&9 TP2&TP3 BNC U/X Pos | Pins 1&9: 20v p-p +/- 0.5v BNC: 10V +/- 0.5V | Pins 1&9: BNC: |
| 50 pin D-sub pin 8/GND | MASSPOS-V J7 pins 2&10 TP2&TP3 BNC V/Y Pos | Pins 2&10: 20v p-p +/- 0.5V BNC: 10V +/- 0.5V | Pins 2&10: BNC: |
| 50 pin D-sub pin 24/GND | MASSPOS-W J7 pins 3&11 TP2&TP3 BNC W/Z Pos | Pins 3&11 20v p-p +/- 0.5V BNC: 10V +/- 0.5V | Pins 3&11: BNC: |

LIGO-T0900346-v1

  - Using a piece of cable to short the indicated pins together, or by actuating the front panel switches, verify the operation of the following binary functions (FP indicates Front Panel, RP indicates Rear Panel). For the rear panel (Binary I/O) functions to operate normally, it is necessary to have all the front panel toggle switches toggled down:

| input | Output | expected response | actual response |
| --- | --- | --- | --- |
| CAL switch from NORM to CAL FP | CAL LED, J6 | LED is lit, J6 pin 38 = 15V |  |
| Basis switch from XYZ to UVW FP | J6 pin 22 | Transition from 0 to 15V |  |
| Period switch from 120SEC to 1SEC FP | J6 pin 6 | Transition from 0 to 15V |  |
| Push AZ button FP | AZ LED, J6 | LED is lit, J6 pin 5 = 15V |  |
| Binary I/O Interface pin 1 to 6 RP | CAL LED, J6 | LED is lit, J6 pin 38 = 15V |  |
| Binary I/O Interface Pin 2 to 7 RP | J6 pin 22 UVW | Transition from 0 to 15V |  |
| BinaryI/O Interface pin 3 to8 RP | J6 pin 6 PERIOD | Transition from 0 to 15V |  |
| Bin I/O Interface pin 4 to 9 RP | AZ LED, J6 | LED is lit, J6 pin 5 = 15V |  |