## outgas calc
Pressure Contribution from Flag Hydrocarbons
40M Lab RGA Scan Results
Job #C71509 |  | Description: |  | First Contact OTF samples |  | Date: | 2009-07-15
|  |  | Oven Used: | C
from ASCI data:
AMU 41 | 2.57e-17 | amps | from RGA scan listing |  | 4.59e-17
AMU 43 | 8.18e-17 | amps | from RGA scan listing |  | 4.53e-17
AMU 53 | 5.18e-17 | amps | from RGA scan listing |  | 6.95e-18
AMU 55 | 5.47e-17 | amps | from RGA scan listing |  | 1.72e-16
AMU 57 | 6.25e-17 | amps | from RGA scan listing |  | 8.61e-18
Sum Flag H/C AMUs | 2.765e-16 | amps |  |  | 2.7876e-16
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
AMU 40 (w/leak open) | 1.2e-14 | amps
AMU 40 (background) | 1.5e-15 | amps
Calib leak contributes | 1.05e-14 | amps | = (w/leak open) - (background)
Flag H/C Outgassing | 6.2147e-12 | torr l/s | = (Sum Flag H/C AMUs) x (Calib leak rate)/(Calib leak contrib.)
Test item surf area | 56.25 | cm2
Normalized outgassing | 1.1048e-13 | torr l/s-cm2 |  | = Flag H/C Outgassing/Test item surf area
Full description: | 3 Microscope slides with approx. 2 mm thickcoating on one surface.
Pre-scan bake: | Light water bake off at 40c for 12 hrs.
total for all AMUs | 7.7552e-12 | amps
approx. total outgas | 1.7431e-07 | torr-l/s
approx. total outgas | 3.0988e-09 | torr-l/s/cm^2
pumping speed | 60 | l/s
estimated total pressure | 2.9051e-09 | torr | (usually 3E-08 torr as measured by ion gauge or FC mode)

## RGA Scan data
2009-07-15 | 2009 11:41:45 AM
Residual Gas Analyzer Software
Software Version | 3.005
Analog Scan Setup:
Data Points in Scan | 991
Units | Amps
Noise Floor | 1
CEM Status | OFF
Points Per AMU | 10
Scan Start Mass | 1 | amu
Scan Stop Mass | 100 | amu
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.1
Filament Current | 1 | mAmps
|  | 1e-18 | HC AMUs
1 | 4.11e-17 | 4.11e-17
1.1 | 8.21e-17 | 8.21e-17
1.2 | 1.71e-16 | 1.71e-16
1.3 | 5.17e-15 | 5.17e-15
1.4 | 1.05e-12 | 1.05e-12
1.5 | 1.72e-12 | 1.72e-12
1.6 | 1.46e-12 | 1.46e-12
1.7 | 9.93e-13 | 9.93e-13
1.8 | 2.18e-13 | 2.18e-13
1.9 | 8.36e-15 | 8.36e-15
2 | 2.03e-15 | 2.03e-15
2.1 | 5.25e-16 | 5.25e-16
2.2 | 1.45e-16 | 1.45e-16
2.3 | 4.35e-17 | 4.35e-17
2.4 | 1.33e-17 | 1.33e-17
2.5 | -2.32e-17 | 1e-18
2.6 | -6.29e-17 | 1e-18
2.7 | -1.05e-16 | 1e-18
2.8 | -8.36e-17 | 1e-18
2.9 | -7.03e-17 | 1e-18
3 | -3.83e-17 | 1e-18
3.1 | 4.96e-16 | 4.96e-16
3.2 | 1.04e-15 | 1.04e-15
3.3 | 2.33e-15 | 2.33e-15
3.4 | 2.87e-15 | 2.87e-15
3.5 | 3.11e-15 | 3.11e-15
3.6 | 3.18e-15 | 3.18e-15
3.7 | 2.39e-15 | 2.39e-15
3.8 | 1.31e-15 | 1.31e-15
3.9 | 3.89e-16 | 3.89e-16
4 | 6.9e-17 | 6.9e-17
4.1 | -3.15e-17 | 1e-18
4.2 | 4.33e-17 | 4.33e-17
4.3 | -4.1e-17 | 1e-18
4.4 | -1.15e-16 | 1e-18
4.5 | -5.08e-17 | 1e-18
4.6 | -8.47e-17 | 1e-18
4.7 | -3.99e-17 | 1e-18
4.8 | -1.22e-17 | 1e-18
4.9 | -3.08e-17 | 1e-18
5 | 1.52e-18 | 1.52e-18
5.1 | -7.62e-19 | 1e-18
5.2 | -8.09e-18 | 1e-18
5.3 | -4.18e-17 | 1e-18
5.4 | -2.46e-17 | 1e-18
5.5 | -2.95e-17 | 1e-18
5.6 | -1.23e-16 | 1e-18
5.7 | -1.09e-16 | 1e-18
5.8 | -5.3e-17 | 1e-18
5.9 | -9.44e-17 | 1e-18
6 | 1e-17 | 1e-17
6.1 | -5.54e-17 | 1e-18
6.2 | -8.44e-17 | 1e-18
6.3 | -2.08e-16 | 1e-18
6.4 | -4.65e-17 | 1e-18
6.5 | -7.96e-17 | 1e-18
6.6 | -1.32e-16 | 1e-18
6.7 | -1e-16 | 1e-18
6.8 | 1.71e-18 | 1.71e-18
6.9 | -9.41e-17 | 1e-18
7 | -1.27e-16 | 1e-18
7.1 | -1.27e-17 | 1e-18
7.2 | -8.4e-17 | 1e-18
7.3 | -5.23e-17 | 1e-18
7.4 | -2.9e-18 | 1e-18
7.5 | 1.59e-17 | 1.59e-17
7.6 | 1.27e-17 | 1.27e-17
7.7 | -2.33e-17 | 1e-18
7.8 | -3.23e-17 | 1e-18
7.9 | 7.81e-18 | 7.81e-18
8 | -8.02e-17 | 1e-18
8.1 | -1.66e-17 | 1e-18
8.2 | -5.72e-17 | 1e-18
8.3 | -8.17e-17 | 1e-18
8.4 | -1.68e-17 | 1e-18
8.5 | -1.26e-16 | 1e-18
8.6 | -3.86e-17 | 1e-18
8.7 | -3.72e-17 | 1e-18
8.8 | -7.98e-17 | 1e-18
8.9 | -1.19e-16 | 1e-18
9 | -9.62e-17 | 1e-18
9.1 | -1.54e-17 | 1e-18
9.2 | 5.81e-18 | 5.81e-18
9.3 | -3.4e-17 | 1e-18
9.4 | -6.63e-17 | 1e-18
9.5 | -6.78e-17 | 1e-18
9.6 | 3.55e-17 | 3.55e-17
9.7 | -6.09e-18 | 1e-18
9.8 | -9.26e-17 | 1e-18
9.9 | -1.34e-17 | 1e-18
10 | -5.91e-17 | 1e-18
10.1 | -2.59e-17 | 1e-18
10.2 | -4.17e-17 | 1e-18
10.3 | -3e-18 | 1e-18
10.4 | -2.44e-17 | 1e-18
10.5 | -4.95e-17 | 1e-18
10.6 | -1.17e-16 | 1e-18
10.7 | -8.74e-17 | 1e-18
10.8 | -5.94e-17 | 1e-18
10.9 | 1.42e-17 | 1.42e-17
11 | -1.4e-17 | 1e-18
11.1 | 2.35e-17 | 2.35e-17
11.2 | 3.04e-16 | 3.04e-16
11.3 | 1.04e-15 | 1.04e-15
11.4 | 1.52e-15 | 1.52e-15
11.5 | 1.55e-15 | 1.55e-15
11.6 | 1.84e-15 | 1.84e-15
11.7 | 1.69e-15 | 1.69e-15
11.8 | 1.29e-15 | 1.29e-15
11.9 | 4.99e-16 | 4.99e-16
12 | 9.9e-17 | 9.9e-17
12.1 | -1.44e-17 | 1e-18
12.2 | 3.53e-17 | 3.53e-17
12.3 | 6.93e-17 | 6.93e-17
12.4 | -2.2e-17 | 1e-18
12.5 | -3.6e-17 | 1e-18
12.6 | 3.24e-17 | 3.24e-17
12.7 | 5.18e-17 | 5.18e-17
12.8 | 5.32e-17 | 5.32e-17
12.9 | 7.9e-18 | 7.9e-18
13 | 2.07e-17 | 2.07e-17
13.1 | 2.12e-16 | 2.12e-16
13.2 | 7.99e-16 | 7.99e-16
13.3 | 3.61e-15 | 3.61e-15
13.4 | 8.85e-15 | 8.85e-15
13.5 | 1.3e-14 | 1.3e-14
13.6 | 1.48e-14 | 1.48e-14
13.7 | 1.35e-14 | 1.35e-14
13.8 | 1e-14 | 1e-14
13.9 | 4.68e-15 | 4.68e-15
14 | 1.69e-15 | 1.69e-15
14.1 | 5.65e-16 | 5.65e-16
14.2 | 1.67e-16 | 1.67e-16
14.3 | 5.3e-17 | 5.3e-17
14.4 | 1.38e-16 | 1.38e-16
14.5 | 2.01e-16 | 2.01e-16
14.6 | 2.34e-16 | 2.34e-16
14.7 | 2.17e-16 | 2.17e-16
14.8 | 2.93e-16 | 2.93e-16
14.9 | 4.07e-16 | 4.07e-16
15 | 6.19e-16 | 6.19e-16
15.1 | 1.13e-15 | 1.13e-15
15.2 | 2.8e-15 | 2.8e-15
15.3 | 1.28e-14 | 1.28e-14
15.4 | 2.48e-14 | 2.48e-14
15.5 | 2.84e-14 | 2.84e-14
15.6 | 2.97e-14 | 2.97e-14
15.7 | 2.73e-14 | 2.73e-14
15.8 | 2.08e-14 | 2.08e-14
15.9 | 8.98e-15 | 8.98e-15
16 | 2.8e-15 | 2.8e-15
16.1 | 1.28e-15 | 1.28e-15
16.2 | 1.38e-15 | 1.38e-15
16.3 | 7.76e-15 | 7.76e-15
16.4 | 2.37e-14 | 2.37e-14
16.5 | 3.4e-14 | 3.4e-14
16.6 | 3.45e-14 | 3.45e-14
16.7 | 3.64e-14 | 3.64e-14
16.8 | 2.68e-14 | 2.68e-14
16.9 | 1.32e-14 | 1.32e-14
17 | 3.45e-15 | 3.45e-15
17.1 | 1.96e-15 | 1.96e-15
17.2 | 3.42e-15 | 3.42e-15
17.3 | 2.97e-14 | 2.97e-14
17.4 | 1.04e-13 | 1.04e-13
17.5 | 1.3e-13 | 1.3e-13
17.6 | 1.42e-13 | 1.42e-13
17.7 | 1.4e-13 | 1.4e-13
17.8 | 1.15e-13 | 1.15e-13
17.9 | 4.9e-14 | 4.9e-14
18 | 6.07e-15 | 6.07e-15
18.1 | 2.3e-15 | 2.3e-15
18.2 | 2.21e-15 | 2.21e-15
18.3 | 5.02e-15 | 5.02e-15
18.4 | 1.25e-14 | 1.25e-14
18.5 | 1.92e-14 | 1.92e-14
18.6 | 2.14e-14 | 2.14e-14
18.7 | 2.09e-14 | 2.09e-14
18.8 | 1.71e-14 | 1.71e-14
18.9 | 9.83e-15 | 9.83e-15
19 | 3.11e-15 | 3.11e-15
19.1 | 1.1e-15 | 1.1e-15
19.2 | 2.96e-16 | 2.96e-16
19.3 | 7.07e-17 | 7.07e-17
19.4 | 1.86e-16 | 1.86e-16
19.5 | 4.84e-16 | 4.84e-16
19.6 | 6.82e-16 | 6.82e-16
19.7 | 8.63e-16 | 8.63e-16
19.8 | 1.01e-15 | 1.01e-15
19.9 | 9.44e-16 | 9.44e-16
20 | 3.65e-16 | 3.65e-16
20.1 | 5.27e-17 | 5.27e-17
20.2 | -2.23e-17 | 1e-18
20.3 | 4.93e-17 | 4.93e-17
20.4 | -1.9e-17 | 1e-18
20.5 | -4.74e-17 | 1e-18
20.6 | -1e-16 | 1e-18
20.7 | 5.1e-17 | 5.1e-17
20.8 | 1.76e-17 | 1.76e-17
20.9 | -2.33e-17 | 1e-18
21 | -6.44e-17 | 1e-18
21.1 | -1.12e-16 | 1e-18
21.2 | -1.86e-17 | 1e-18
21.3 | -3.04e-17 | 1e-18
21.4 | -5.69e-17 | 1e-18
21.5 | -3.06e-17 | 1e-18
21.6 | 1.36e-17 | 1.36e-17
21.7 | 3.36e-17 | 3.36e-17
21.8 | -4.25e-17 | 1e-18
21.9 | 2.44e-17 | 2.44e-17
22 | 2.61e-17 | 2.61e-17
22.1 | -3.38e-17 | 1e-18
22.2 | -4.69e-17 | 1e-18
22.3 | -8.78e-17 | 1e-18
22.4 | -1.77e-17 | 1e-18
22.5 | -1.03e-16 | 1e-18
22.6 | 1.29e-17 | 1.29e-17
22.7 | -5e-17 | 1e-18
22.8 | -5.54e-17 | 1e-18
22.9 | -1.94e-17 | 1e-18
23 | -4.39e-17 | 1e-18
23.1 | -4.3e-17 | 1e-18
23.2 | 1.56e-17 | 1.56e-17
23.3 | -8.98e-17 | 1e-18
23.4 | -4.14e-17 | 1e-18
23.5 | -3.24e-17 | 1e-18
23.6 | -5.66e-17 | 1e-18
23.7 | -3.62e-17 | 1e-18
23.8 | -1.09e-17 | 1e-18
23.9 | -1.28e-16 | 1e-18
24 | -5.97e-17 | 1e-18
24.1 | -7.45e-17 | 1e-18
24.2 | -5.52e-18 | 1e-18
24.3 | -4.36e-17 | 1e-18
24.4 | -4.2e-17 | 1e-18
24.5 | 1.37e-17 | 1.37e-17
24.6 | -2.4e-17 | 1e-18
24.7 | -1.28e-17 | 1e-18
24.8 | -4.21e-17 | 1e-18
24.9 | 2.86e-18 | 2.86e-18
25 | 6.76e-18 | 6.76e-18
25.1 | 2.97e-17 | 2.97e-17
25.2 | 7.15e-17 | 7.15e-17
25.3 | -2.51e-17 | 1e-18
25.4 | 1.88e-17 | 1.88e-17
25.5 | -2.2e-17 | 1e-18
25.6 | 2.11e-16 | 2.11e-16
25.7 | 1.39e-16 | 1.39e-16
25.8 | 3.05e-16 | 3.05e-16
25.9 | 3.01e-16 | 3.01e-16
26 | 1.96e-16 | 1.96e-16
26.1 | -4.33e-18 | 1e-18
26.2 | 1.67e-17 | 1.67e-17
26.3 | 1.15e-16 | 1.15e-16
26.4 | -1.21e-17 | 1e-18
26.5 | -4.22e-17 | 1e-18
26.6 | 1.11e-16 | 1.11e-16
26.7 | 1.63e-16 | 1.63e-16
26.8 | 1.86e-16 | 1.86e-16
26.9 | 3.06e-16 | 3.06e-16
27 | 2.91e-16 | 2.91e-16
27.1 | 2.9e-16 | 2.9e-16
27.2 | 7.61e-16 | 7.61e-16
27.3 | 1.57e-15 | 1.57e-15
27.4 | 4.37e-15 | 4.37e-15
27.5 | 3.94e-14 | 3.94e-14
27.6 | 1.31e-13 | 1.31e-13
27.7 | 1.83e-13 | 1.83e-13
27.8 | 2.07e-13 | 2.07e-13
27.9 | 1.89e-13 | 1.89e-13
28 | 9.12e-14 | 9.12e-14
28.1 | 9.79e-15 | 9.79e-15
28.2 | 2.34e-15 | 2.34e-15
28.3 | 7.56e-16 | 7.56e-16
28.4 | 1.67e-16 | 1.67e-16
28.5 | 1.09e-16 | 1.09e-16
28.6 | 4.95e-16 | 4.95e-16
28.7 | 8.97e-16 | 8.97e-16
28.8 | 1.48e-15 | 1.48e-15
28.9 | 1.47e-15 | 1.47e-15
29 | 9.89e-16 | 9.89e-16
29.1 | 3.57e-16 | 3.57e-16
29.2 | 2.52e-17 | 2.52e-17
29.3 | -6.41e-17 | 1e-18
29.4 | -7.26e-17 | 1e-18
29.5 | -4.87e-17 | 1e-18
29.6 | 1.84e-17 | 1.84e-17
29.7 | 2.27e-17 | 2.27e-17
29.8 | 1.39e-16 | 1.39e-16
29.9 | 1.17e-16 | 1.17e-16
30 | 1.36e-16 | 1.36e-16
30.1 | 4.44e-17 | 4.44e-17
30.2 | -3.82e-17 | 1e-18
30.3 | -2.2e-17 | 1e-18
30.4 | -1.79e-17 | 1e-18
30.5 | -1.68e-16 | 1e-18
30.6 | -9.23e-17 | 1e-18
30.7 | 2.47e-17 | 2.47e-17
30.8 | -2.92e-17 | 1e-18
30.9 | -6.25e-17 | 1e-18
31 | -9.15e-17 | 1e-18
31.1 | 2.74e-17 | 2.74e-17
31.2 | 5.54e-17 | 5.54e-17
31.3 | -1.16e-17 | 1e-18
31.4 | 2.18e-16 | 2.18e-16
31.5 | 8.31e-16 | 8.31e-16
31.6 | 4.38e-15 | 4.38e-15
31.7 | 1.44e-14 | 1.44e-14
31.8 | 2.19e-14 | 2.19e-14
31.9 | 2.61e-14 | 2.61e-14
32 | 2.12e-14 | 2.12e-14
32.1 | 7.41e-15 | 7.41e-15
32.2 | 2.23e-15 | 2.23e-15
32.3 | 7.79e-16 | 7.79e-16
32.4 | 1.96e-16 | 1.96e-16
32.5 | 1.2e-16 | 1.2e-16
32.6 | 8.71e-18 | 8.71e-18
32.7 | -1.43e-18 | 1e-18
32.8 | -1.07e-16 | 1e-18
32.9 | -8.18e-17 | 1e-18
33 | -1.48e-18 | 1e-18
33.1 | -2.18e-17 | 1e-18
33.2 | 6.62e-17 | 6.62e-17
33.3 | 1.02e-17 | 1.02e-17
33.4 | -2.22e-17 | 1e-18
33.5 | -3.9e-17 | 1e-18
33.6 | -5.24e-18 | 1e-18
33.7 | 7.3e-17 | 7.3e-17
33.8 | 2.29e-17 | 2.29e-17
33.9 | 8.22e-17 | 8.22e-17
34 | 7.6e-17 | 7.6e-17
34.1 | -1.21e-17 | 1e-18
34.2 | -5.97e-17 | 1e-18
34.3 | -9.85e-17 | 1e-18
34.4 | -1e-16 | 1e-18
34.5 | -7.23e-18 | 1e-18
34.6 | -7.81e-17 | 1e-18
34.7 | 2.54e-17 | 2.54e-17
34.8 | 4.94e-17 | 4.94e-17
34.9 | 2.48e-17 | 2.48e-17
35 | 1.86e-17 | 1.86e-17
35.1 | -1.24e-16 | 1e-18
35.2 | -1.82e-17 | 1e-18
35.3 | -4.39e-17 | 1e-18
35.4 | 2.86e-19 | 1e-18
35.5 | -3.47e-17 | 1e-18
35.6 | 1.49e-17 | 1.49e-17
35.7 | 3.45e-17 | 3.45e-17
35.8 | 5.68e-17 | 5.68e-17
35.9 | 1.26e-16 | 1.26e-16
36 | 8.14e-17 | 8.14e-17
36.1 | 4.68e-17 | 4.68e-17
36.2 | -1.33e-17 | 1e-18
36.3 | -1.58e-17 | 1e-18
36.4 | -8.27e-17 | 1e-18
36.5 | -3.33e-17 | 1e-18
36.6 | 9e-18 | 9e-18
36.7 | -5.39e-17 | 1e-18
36.8 | 3.61e-17 | 3.61e-17
36.9 | 5.97e-17 | 5.97e-17
37 | 3.57e-18 | 3.57e-18
37.1 | -1.78e-17 | 1e-18
37.2 | -1.7e-17 | 1e-18
37.3 | -6.3e-17 | 1e-18
37.4 | -1.46e-17 | 1e-18
37.5 | 1.37e-17 | 1.37e-17
37.6 | 3.38e-17 | 3.38e-17
37.7 | 7e-18 | 7e-18
37.8 | -6.07e-17 | 1e-18
37.9 | -6.33e-17 | 1e-18
38 | 2.09e-17 | 2.09e-17
38.1 | 1.46e-17 | 1.46e-17
38.2 | -2.14e-17 | 1e-18
38.3 | 2.95e-18 | 2.95e-18
38.4 | 3.16e-17 | 3.16e-17
38.5 | -3.51e-17 | 1e-18
38.6 | -7.87e-17 | 1e-18
38.7 | 3.86e-18 | 3.86e-18
38.8 | -1.95e-18 | 1e-18
38.9 | -3.75e-17 | 1e-18
39 | 2.83e-17 | 2.83e-17
39.1 | 4.13e-17 | 4.13e-17
39.2 | -6.56e-17 | 1e-18
39.3 | -4.86e-17 | 1e-18
39.4 | -7.89e-17 | 1e-18
39.5 | -5.77e-17 | 1e-18
39.6 | -5.09e-17 | 1e-18
39.7 | 3.66e-18 | 3.66e-18
39.8 | 7.55e-17 | 7.55e-17
39.9 | 4.04e-16 | 4.04e-16
40 | 6.48e-16 | 6.48e-16
40.1 | 6.24e-16 | 6.24e-16
40.2 | 2.43e-16 | 2.43e-16
40.3 | 3.47e-17 | 3.47e-17
40.4 | -6.23e-17 | 1e-18
40.5 | -5.41e-17 | 1e-18
40.6 | -9.21e-17 | 1e-18
40.7 | 1.89e-17 | 1.89e-17
40.8 | 6.81e-18 | 6.81e-18
40.9 | 6.71e-18 | 6.71e-18
41 | -2.57e-17 | 1e-18
41.1 | 4.59e-17 | 4.59e-17 | 4.59e-17
41.2 | 4.31e-17 | 4.31e-17
41.3 | -3.38e-17 | 1e-18
41.4 | -5.27e-17 | 1e-18
41.5 | -5.6e-17 | 1e-18
41.6 | -5.37e-17 | 1e-18
41.7 | -2.1e-17 | 1e-18
41.8 | -6.04e-17 | 1e-18
41.9 | 6.71e-18 | 6.71e-18
42 | -1.87e-17 | 1e-18
42.1 | 9.56e-17 | 9.56e-17
42.2 | 7.18e-17 | 7.18e-17
42.3 | -1.52e-17 | 1e-18
42.4 | -4.34e-17 | 1e-18
42.5 | -3.4e-17 | 1e-18
42.6 | -1.98e-17 | 1e-18
42.7 | -2.01e-17 | 1e-18
42.8 | -3.53e-17 | 1e-18
42.9 | -3.38e-17 | 1e-18
43 | -8.18e-17 | 1e-18
43.1 | -7.62e-19 | 1e-18
43.2 | -3.05e-18 | 1e-18
43.3 | 4.53e-17 | 4.53e-17 | 4.53e-17
43.4 | 1.71e-17 | 1.71e-17
43.5 | 1.66e-17 | 1.66e-17
43.6 | 9.27e-17 | 9.27e-17
43.7 | 1.72e-16 | 1.72e-16
43.8 | 1.01e-15 | 1.01e-15
43.9 | 3.38e-15 | 3.38e-15
44 | 7.64e-15 | 7.64e-15
44.1 | 1.15e-14 | 1.15e-14
44.2 | 8.79e-15 | 8.79e-15
44.3 | 3.14e-15 | 3.14e-15
44.4 | 1.16e-15 | 1.16e-15
44.5 | 3.55e-16 | 3.55e-16
44.6 | 2.31e-17 | 2.31e-17
44.7 | 4.71e-18 | 4.71e-18
44.8 | 3.14e-18 | 3.14e-18
44.9 | -4.3e-17 | 1e-18
45 | 1.75e-17 | 1.75e-17
45.1 | 9.97e-17 | 9.97e-17
45.2 | 3.75e-17 | 3.75e-17
45.3 | 3.33e-19 | 1e-18
45.4 | -7.51e-17 | 1e-18
45.5 | -8.84e-17 | 1e-18
45.6 | -6.39e-17 | 1e-18
45.7 | -4.64e-17 | 1e-18
45.8 | -4.2e-17 | 1e-18
45.9 | -6.78e-17 | 1e-18
46 | -8.64e-17 | 1e-18
46.1 | 2.28e-18 | 2.28e-18
46.2 | -1.98e-17 | 1e-18
46.3 | -7.27e-17 | 1e-18
46.4 | -3.32e-17 | 1e-18
46.5 | -7.66e-18 | 1e-18
46.6 | 8.81e-18 | 8.81e-18
46.7 | -7.76e-17 | 1e-18
46.8 | -1.11e-16 | 1e-18
46.9 | -4.69e-17 | 1e-18
47 | -5.85e-17 | 1e-18
47.1 | -3.61e-17 | 1e-18
47.2 | 5.64e-17 | 5.64e-17
47.3 | 8.9e-18 | 8.9e-18
47.4 | -5.91e-17 | 1e-18
47.5 | 1.05e-17 | 1.05e-17
47.6 | -7.27e-17 | 1e-18
47.7 | 1.31e-17 | 1.31e-17
47.8 | -1.3e-16 | 1e-18
47.9 | 6.09e-17 | 6.09e-17
48 | -3.81e-18 | 1e-18
48.1 | -2.3e-17 | 1e-18
48.2 | -1.81e-18 | 1e-18
48.3 | -4.83e-17 | 1e-18
48.4 | -8.51e-17 | 1e-18
48.5 | -6.67e-17 | 1e-18
48.6 | -4.99e-17 | 1e-18
48.7 | -7.44e-17 | 1e-18
48.8 | -6.65e-17 | 1e-18
48.9 | -6.79e-17 | 1e-18
49 | -4.47e-18 | 1e-18
49.1 | 8.81e-18 | 8.81e-18
49.2 | -1.11e-16 | 1e-18
49.3 | -3.15e-17 | 1e-18
49.4 | -4.81e-17 | 1e-18
49.5 | -6.58e-17 | 1e-18
49.6 | -9.22e-17 | 1e-18
49.7 | -7.26e-17 | 1e-18
49.8 | -1.29e-17 | 1e-18
49.9 | -2.91e-17 | 1e-18
50 | 4.76e-18 | 4.76e-18
50.1 | 4.26e-17 | 4.26e-17
50.2 | -3.97e-17 | 1e-18
50.3 | -3.5e-17 | 1e-18
50.4 | -8.8e-17 | 1e-18
50.5 | -4.73e-17 | 1e-18
50.6 | -5.73e-17 | 1e-18
50.7 | 5.92e-17 | 5.92e-17
50.8 | -1.18e-17 | 1e-18
50.9 | 4.76e-18 | 4.76e-18
51 | -8.03e-17 | 1e-18
51.1 | -8.95e-18 | 1e-18
51.2 | -1.22e-16 | 1e-18
51.3 | -9.74e-17 | 1e-18
51.4 | 3.33e-19 | 1e-18
51.5 | -7.88e-17 | 1e-18
51.6 | -3.59e-17 | 1e-18
51.7 | -2.1e-17 | 1e-18
51.8 | -6.17e-17 | 1e-18
51.9 | -8.6e-17 | 1e-18
52 | -1.57e-16 | 1e-18
52.1 | -8e-17 | 1e-18
52.2 | -5.82e-17 | 1e-18
52.3 | -5.78e-17 | 1e-18
52.4 | -5.46e-17 | 1e-18
52.5 | -2.94e-17 | 1e-18
52.6 | 1.69e-17 | 1.69e-17
52.7 | -7.03e-17 | 1e-18
52.8 | 1.29e-18 | 1.29e-18
52.9 | -4.13e-17 | 1e-18
53 | -5.18e-17 | 1e-18
53.1 | -3.33e-18 | 1e-18
53.2 | -8.71e-17 | 1e-18
53.3 | -5.71e-18 | 1e-18
53.4 | 6.95e-18 | 6.95e-18 | 6.95e-18
53.5 | 2.57e-17 | 2.57e-17
53.6 | -7.53e-17 | 1e-18
53.7 | -6.82e-17 | 1e-18
53.8 | -7.92e-17 | 1e-18
53.9 | -3.28e-17 | 1e-18
54 | -6.81e-17 | 1e-18
54.1 | -5.07e-17 | 1e-18
54.2 | -1.61e-16 | 1e-18
54.3 | -8.35e-17 | 1e-18
54.4 | -1.39e-16 | 1e-18
54.5 | -1.41e-17 | 1e-18
54.6 | 2.29e-17 | 2.29e-17
54.7 | 6.15e-17 | 6.15e-17
54.8 | -2.55e-17 | 1e-18
54.9 | -4.16e-17 | 1e-18
55 | -5.47e-17 | 1e-18
55.1 | -1.55e-16 | 1e-18
55.2 | -5.2e-17 | 1e-18
55.3 | 1.72e-16 | 1.72e-16 | 1.72e-16
55.4 | -6.37e-17 | 1e-18
55.5 | -6.1e-17 | 1e-18
55.6 | -4.9e-17 | 1e-18
55.7 | -3.31e-17 | 1e-18
55.8 | -3.62e-17 | 1e-18
55.9 | -9.66e-17 | 1e-18
56 | -5.63e-17 | 1e-18
56.1 | -1.09e-18 | 1e-18
56.2 | 1.55e-17 | 1.55e-17
56.3 | -1.19e-16 | 1e-18
56.4 | -8.98e-17 | 1e-18
56.5 | -2.57e-17 | 1e-18
56.6 | 2.71e-17 | 2.71e-17
56.7 | -1.24e-17 | 1e-18
56.8 | 8.61e-18 | 8.61e-18 | 8.61e-18
56.9 | -6.07e-17 | 1e-18
57 | -6.25e-17 | 1e-18
57.1 | -5.23e-17 | 1e-18
57.2 | -9.32e-17 | 1e-18
57.3 | -4.18e-17 | 1e-18
57.4 | -1.14e-16 | 1e-18
57.5 | -5.3e-17 | 1e-18
57.6 | -8.92e-17 | 1e-18
57.7 | -3.87e-17 | 1e-18
57.8 | -5.67e-17 | 1e-18
57.9 | -2.96e-17 | 1e-18
58 | -2.1e-17 | 1e-18
58.1 | -2.93e-17 | 1e-18
58.2 | -7.03e-17 | 1e-18
58.3 | -1.56e-17 | 1e-18
58.4 | -6.57e-17 | 1e-18
58.5 | -6.69e-17 | 1e-18
58.6 | -7.4e-17 | 1e-18
58.7 | 2.68e-17 | 2.68e-17
58.8 | 8.83e-17 | 8.83e-17
58.9 | 8.72e-17 | 8.72e-17
59 | 6.8e-17 | 6.8e-17
59.1 | 2.24e-17 | 2.24e-17
59.2 | -8.02e-17 | 1e-18
59.3 | -2.69e-17 | 1e-18
59.4 | -7.22e-17 | 1e-18
59.5 | -9.1e-17 | 1e-18
59.6 | -1.47e-16 | 1e-18
59.7 | -5.98e-17 | 1e-18
59.8 | -4.5e-17 | 1e-18
59.9 | -5.93e-17 | 1e-18
60 | -8.76e-17 | 1e-18
60.1 | -9.21e-17 | 1e-18
60.2 | -4.34e-17 | 1e-18
60.3 | -8.39e-17 | 1e-18
60.4 | -6.58e-17 | 1e-18
60.5 | -5.64e-17 | 1e-18
60.6 | -7.66e-17 | 1e-18
60.7 | -1.02e-16 | 1e-18
60.8 | -9.9e-17 | 1e-18
60.9 | -1.38e-16 | 1e-18
61 | -4.79e-17 | 1e-18
61.1 | -3.81e-17 | 1e-18
61.2 | -1.01e-17 | 1e-18
61.3 | -4.71e-18 | 1e-18
61.4 | -4.78e-17 | 1e-18
61.5 | -6.43e-17 | 1e-18
61.6 | -4.28e-17 | 1e-18
61.7 | -5.55e-17 | 1e-18
61.8 | -1.12e-16 | 1e-18
61.9 | -5.48e-17 | 1e-18
62 | -1.42e-17 | 1e-18
62.1 | 1.16e-16 | 1.16e-16
62.2 | -4.71e-18 | 1e-18
62.3 | -1.32e-16 | 1e-18
62.4 | -8.84e-17 | 1e-18
62.5 | -1.53e-17 | 1e-18
62.6 | -5.95e-17 | 1e-18
62.7 | 6.24e-18 | 6.24e-18
62.8 | -9.19e-17 | 1e-18
62.9 | 1.02e-17 | 1.02e-17
63 | -5.74e-17 | 1e-18
63.1 | -7.94e-17 | 1e-18
63.2 | -1.06e-16 | 1e-18
63.3 | -8.77e-17 | 1e-18
63.4 | 5.81e-18 | 5.81e-18
63.5 | -1.1e-16 | 1e-18
63.6 | -4.83e-17 | 1e-18
63.7 | -8.7e-17 | 1e-18
63.8 | -3.95e-17 | 1e-18
63.9 | -4.35e-17 | 1e-18
64 | -1.06e-16 | 1e-18
64.1 | -3.81e-17 | 1e-18
64.2 | -2.27e-17 | 1e-18
64.3 | -3.54e-17 | 1e-18
64.4 | -5.11e-17 | 1e-18
64.5 | -7.38e-17 | 1e-18
64.6 | -4.14e-18 | 1e-18
64.7 | -3.86e-18 | 1e-18
64.8 | -1.78e-17 | 1e-18
64.9 | -1.48e-17 | 1e-18
65 | -4.62e-17 | 1e-18
65.1 | -2.31e-17 | 1e-18
65.2 | -5.76e-18 | 1e-18
65.3 | -8.39e-17 | 1e-18
65.4 | -1.43e-16 | 1e-18
65.5 | -6.09e-17 | 1e-18
65.6 | -4.08e-17 | 1e-18
65.7 | 2.7e-17 | 2.7e-17
65.8 | -1.58e-17 | 1e-18
65.9 | -6.38e-17 | 1e-18
66 | -1.11e-16 | 1e-18
66.1 | -1.26e-16 | 1e-18
66.2 | -7e-17 | 1e-18
66.3 | -4.64e-17 | 1e-18
66.4 | -9e-17 | 1e-18
66.5 | -5.18e-17 | 1e-18
66.6 | -5e-18 | 1e-18
66.7 | -1.03e-16 | 1e-18
66.8 | -4.69e-17 | 1e-18
66.9 | -1.21e-16 | 1e-18
67 | -2.05e-17 | 1e-18
67.1 | -4.32e-17 | 1e-18
67.2 | -5.74e-17 | 1e-18
67.3 | -6.78e-17 | 1e-18
67.4 | 1.32e-17 | 1.32e-17
67.5 | -5.95e-18 | 1e-18
67.6 | -1.14e-16 | 1e-18
67.7 | -9.36e-17 | 1e-18
67.8 | -2.67e-17 | 1e-18
67.9 | -2.22e-17 | 1e-18
68 | -3.06e-17 | 1e-18
68.1 | 9.33e-18 | 9.33e-18
68.2 | -5.42e-17 | 1e-18
68.3 | -8.52e-17 | 1e-18
68.4 | 1.47e-17 | 1.47e-17
68.5 | -4.4e-17 | 1e-18
68.6 | -1.4e-16 | 1e-18
68.7 | -6.77e-17 | 1e-18
68.8 | -1.9e-16 | 1e-18
68.9 | -3.2e-17 | 1e-18
69 | -9.2e-17 | 1e-18
69.1 | -9.35e-17 | 1e-18
69.2 | -2.49e-17 | 1e-18
69.3 | 1.15e-17 | 1.15e-17
69.4 | -4.01e-17 | 1e-18
69.5 | -3.96e-17 | 1e-18
69.6 | -3.63e-17 | 1e-18
69.7 | -6.16e-17 | 1e-18
69.8 | -8.14e-17 | 1e-18
69.9 | -8.48e-17 | 1e-18
70 | 2.92e-17 | 2.92e-17
70.1 | -1.09e-16 | 1e-18
70.2 | -5.04e-17 | 1e-18
70.3 | -2.66e-17 | 1e-18
70.4 | -1.33e-16 | 1e-18
70.5 | -1.09e-16 | 1e-18
70.6 | -2.14e-17 | 1e-18
70.7 | -3.14e-18 | 1e-18
70.8 | -5.06e-17 | 1e-18
70.9 | -6.94e-17 | 1e-18
71 | -1.11e-16 | 1e-18
71.1 | -5.03e-17 | 1e-18
71.2 | -2.27e-17 | 1e-18
71.3 | -7.16e-17 | 1e-18
71.4 | -5.37e-17 | 1e-18
71.5 | -3.48e-17 | 1e-18
71.6 | -9.21e-17 | 1e-18
71.7 | -7.41e-17 | 1e-18
71.8 | -9.53e-17 | 1e-18
71.9 | -2.31e-17 | 1e-18
72 | -2.22e-17 | 1e-18
72.1 | -3e-17 | 1e-18
72.2 | -8.27e-17 | 1e-18
72.3 | -7.06e-17 | 1e-18
72.4 | -6.37e-17 | 1e-18
72.5 | -5.3e-17 | 1e-18
72.6 | 5.08e-17 | 5.08e-17
72.7 | -2.07e-17 | 1e-18
72.8 | -8.5e-17 | 1e-18
72.9 | -4.49e-17 | 1e-18
73 | -5.49e-17 | 1e-18
73.1 | 1.9e-18 | 1.9e-18
73.2 | -1.29e-16 | 1e-18
73.3 | -1.33e-16 | 1e-18
73.4 | 1.98e-17 | 1.98e-17
73.5 | -1.72e-16 | 1e-18
73.6 | 3.76e-18 | 3.76e-18
73.7 | -5.79e-17 | 1e-18
73.8 | 3.19e-17 | 3.19e-17
73.9 | -1.78e-17 | 1e-18
74 | -6.88e-17 | 1e-18
74.1 | -4.78e-17 | 1e-18
74.2 | 2.14e-17 | 2.14e-17
74.3 | -8.23e-18 | 1e-18
74.4 | -5.85e-17 | 1e-18
74.5 | -8.48e-17 | 1e-18
74.6 | -7.47e-17 | 1e-18
74.7 | -6.61e-17 | 1e-18
74.8 | -1e-17 | 1e-18
74.9 | -2.1e-17 | 1e-18
75 | -2.14e-17 | 1e-18
75.1 | -1.14e-17 | 1e-18
75.2 | -7.43e-17 | 1e-18
75.3 | -8.62e-17 | 1e-18
75.4 | -6.58e-17 | 1e-18
75.5 | -1.17e-16 | 1e-18
75.6 | -4.56e-17 | 1e-18
75.7 | -7.64e-17 | 1e-18
75.8 | -2.9e-17 | 1e-18
75.9 | -7.97e-17 | 1e-18
76 | 4.42e-17 | 4.42e-17
76.1 | -1.67e-17 | 1e-18
76.2 | -5.98e-17 | 1e-18
76.3 | -4.53e-17 | 1e-18
76.4 | 1.52e-17 | 1.52e-17
76.5 | -7.08e-17 | 1e-18
76.6 | -4.47e-18 | 1e-18
76.7 | -5.36e-17 | 1e-18
76.8 | 5.81e-17 | 5.81e-17
76.9 | -1.03e-16 | 1e-18
77 | -6.94e-17 | 1e-18
77.1 | -6.94e-17 | 1e-18
77.2 | -3.56e-17 | 1e-18
77.3 | -1.11e-16 | 1e-18
77.4 | 5.51e-17 | 5.51e-17
77.5 | -5.78e-17 | 1e-18
77.6 | -4.09e-18 | 1e-18
77.7 | -8.6e-17 | 1e-18
77.8 | -3.66e-18 | 1e-18
77.9 | 5.32e-17 | 5.32e-17
78 | 1.66e-17 | 1.66e-17
78.1 | -6.28e-17 | 1e-18
78.2 | -9.5e-17 | 1e-18
78.3 | -8.56e-17 | 1e-18
78.4 | -2.07e-17 | 1e-18
78.5 | -3.01e-17 | 1e-18
78.6 | -4.19e-17 | 1e-18
78.7 | -3.1e-17 | 1e-18
78.8 | -8.71e-17 | 1e-18
78.9 | -6.82e-17 | 1e-18
79 | -1.88e-17 | 1e-18
79.1 | -7.68e-17 | 1e-18
79.2 | -7.95e-18 | 1e-18
79.3 | -5.93e-17 | 1e-18
79.4 | -3.88e-17 | 1e-18
79.5 | -1.03e-16 | 1e-18
79.6 | -1.94e-17 | 1e-18
79.7 | -1e-17 | 1e-18
79.8 | -7.04e-17 | 1e-18
79.9 | -8.38e-17 | 1e-18
80 | -1.29e-18 | 1e-18
80.1 | -5.58e-17 | 1e-18
80.2 | -9.45e-17 | 1e-18
80.3 | -8.17e-17 | 1e-18
80.4 | -1.15e-16 | 1e-18
80.5 | -9.05e-17 | 1e-18
80.6 | -6.87e-17 | 1e-18
80.7 | -7.31e-17 | 1e-18
80.8 | -7.76e-18 | 1e-18
80.9 | -5.38e-17 | 1e-18
81 | -8.05e-17 | 1e-18
81.1 | 1.23e-17 | 1.23e-17
81.2 | -6.33e-17 | 1e-18
81.3 | -1.32e-16 | 1e-18
81.4 | -9.99e-17 | 1e-18
81.5 | -2.62e-17 | 1e-18
81.6 | -7.11e-17 | 1e-18
81.7 | -1.09e-17 | 1e-18
81.8 | -1.07e-16 | 1e-18
81.9 | -1.02e-16 | 1e-18
82 | -1.01e-16 | 1e-18
82.1 | -8.11e-17 | 1e-18
82.2 | -4.84e-17 | 1e-18
82.3 | -1.23e-16 | 1e-18
82.4 | -2.39e-17 | 1e-18
82.5 | -1.74e-16 | 1e-18
82.6 | -1.64e-17 | 1e-18
82.7 | -7.05e-17 | 1e-18
82.8 | -1.09e-16 | 1e-18
82.9 | -2.13e-17 | 1e-18
83 | -7.23e-17 | 1e-18
83.1 | -5.45e-17 | 1e-18
83.2 | -4.54e-17 | 1e-18
83.3 | -2.69e-17 | 1e-18
83.4 | -3.39e-17 | 1e-18
83.5 | -1.12e-16 | 1e-18
83.6 | -5.18e-17 | 1e-18
83.7 | -7.6e-17 | 1e-18
83.8 | -1.25e-16 | 1e-18
83.9 | -4.62e-17 | 1e-18
84 | -2.52e-18 | 1e-18
84.1 | -4.84e-17 | 1e-18
84.2 | -1.31e-17 | 1e-18
84.3 | -2.1e-17 | 1e-18
84.4 | -2.92e-17 | 1e-18
84.5 | -2.37e-17 | 1e-18
84.6 | -2.78e-17 | 1e-18
84.7 | -3.32e-17 | 1e-18
84.8 | -8.27e-17 | 1e-18
84.9 | -3.77e-17 | 1e-18
85 | -1.69e-17 | 1e-18
85.1 | 2.8e-17 | 2.8e-17
85.2 | -2.71e-18 | 1e-18
85.3 | -6.04e-17 | 1e-18
85.4 | -1.28e-17 | 1e-18
85.5 | 1.11e-17 | 1.11e-17
85.6 | -8.52e-17 | 1e-18
85.7 | -9.14e-17 | 1e-18
85.8 | -3.68e-17 | 1e-18
85.9 | -8.8e-17 | 1e-18
86 | -4.99e-17 | 1e-18
86.1 | -3.86e-17 | 1e-18
86.2 | -6.31e-17 | 1e-18
86.3 | -9.61e-17 | 1e-18
86.4 | -8.33e-17 | 1e-18
86.5 | -3.97e-17 | 1e-18
86.6 | 1.99e-17 | 1.99e-17
86.7 | -2.04e-17 | 1e-18
86.8 | -6.63e-17 | 1e-18
86.9 | -3.37e-17 | 1e-18
87 | -4.86e-17 | 1e-18
87.1 | -3e-17 | 1e-18
87.2 | -8.19e-17 | 1e-18
87.3 | -3.09e-17 | 1e-18
87.4 | -3.12e-17 | 1e-18
87.5 | -3.08e-17 | 1e-18
87.6 | 6.81e-17 | 6.81e-17
87.7 | 2.68e-17 | 2.68e-17
87.8 | 1.52e-17 | 1.52e-17
87.9 | -7.03e-17 | 1e-18
88 | -2.97e-17 | 1e-18
88.1 | -3.79e-17 | 1e-18
88.2 | -4.59e-17 | 1e-18
88.3 | -6.73e-17 | 1e-18
88.4 | -3.77e-17 | 1e-18
88.5 | -3.92e-17 | 1e-18
88.6 | -4.79e-17 | 1e-18
88.7 | 1.1e-17 | 1.1e-17
88.8 | 1.19e-18 | 1.19e-18
88.9 | -1.19e-16 | 1e-18
89 | -3.17e-17 | 1e-18
89.1 | -6.95e-18 | 1e-18
89.2 | -5.34e-17 | 1e-18
89.3 | -9.38e-18 | 1e-18
89.4 | -1.1e-16 | 1e-18
89.5 | -7.14e-19 | 1e-18
89.6 | -2.37e-17 | 1e-18
89.7 | -3.38e-17 | 1e-18
89.8 | -4.82e-17 | 1e-18
89.9 | -2.57e-18 | 1e-18
90 | -4.57e-17 | 1e-18
90.1 | -3.77e-17 | 1e-18
90.2 | -5.27e-17 | 1e-18
90.3 | -5.76e-17 | 1e-18
90.4 | 1.08e-17 | 1.08e-17
90.5 | 8.92e-17 | 8.92e-17
90.6 | -9.8e-18 | 1e-18
90.7 | -4.79e-17 | 1e-18
90.8 | -2.02e-17 | 1e-18
90.9 | -2.17e-17 | 1e-18
91 | -8.89e-17 | 1e-18
91.1 | -1.03e-16 | 1e-18
91.2 | -4.01e-17 | 1e-18
91.3 | -1.19e-17 | 1e-18
91.4 | 6.2e-17 | 6.2e-17
91.5 | -1.07e-16 | 1e-18
91.6 | -8.69e-17 | 1e-18
91.7 | -1.21e-16 | 1e-18
91.8 | -1.19e-17 | 1e-18
91.9 | 2.91e-17 | 2.91e-17
92 | -3.03e-17 | 1e-18
92.1 | -3.24e-17 | 1e-18
92.2 | -1.68e-17 | 1e-18
92.3 | -1.94e-17 | 1e-18
92.4 | -4.28e-18 | 1e-18
92.5 | -5.15e-17 | 1e-18
92.6 | -8.42e-17 | 1e-18
92.7 | 3.01e-17 | 3.01e-17
92.8 | -8.04e-17 | 1e-18
92.9 | -6.05e-17 | 1e-18
93 | -8.04e-17 | 1e-18
93.1 | -3.96e-17 | 1e-18
93.2 | -4.64e-17 | 1e-18
93.3 | -3.64e-17 | 1e-18
93.4 | -6.41e-17 | 1e-18
93.5 | 5.76e-18 | 5.76e-18
93.6 | 4.93e-17 | 4.93e-17
93.7 | -1.98e-17 | 1e-18
93.8 | -1.21e-16 | 1e-18
93.9 | -7.16e-17 | 1e-18
94 | -2.76e-18 | 1e-18
94.1 | -1.43e-17 | 1e-18
94.2 | -9.67e-17 | 1e-18
94.3 | -7.94e-17 | 1e-18
94.4 | -3.88e-17 | 1e-18
94.5 | -3.34e-17 | 1e-18
94.6 | 2.99e-17 | 2.99e-17
94.7 | -3.96e-17 | 1e-18
94.8 | -6.09e-17 | 1e-18
94.9 | -1.94e-17 | 1e-18
95 | -4.62e-17 | 1e-18
95.1 | -2.55e-17 | 1e-18
95.2 | -8.37e-17 | 1e-18
95.3 | -7.95e-17 | 1e-18
95.4 | -8.83e-17 | 1e-18
95.5 | -7.03e-17 | 1e-18
95.6 | -4.29e-17 | 1e-18
95.7 | -2.78e-17 | 1e-18
95.8 | -7.73e-17 | 1e-18
95.9 | 6.47e-18 | 6.47e-18
96 | -4.49e-17 | 1e-18
96.1 | -8e-17 | 1e-18
96.2 | -1.33e-17 | 1e-18
96.3 | 3.15e-17 | 3.15e-17
96.4 | -9.77e-17 | 1e-18
96.5 | -1.15e-16 | 1e-18
96.6 | -9.14e-17 | 1e-18
96.7 | -2.25e-17 | 1e-18
96.8 | 2.19e-17 | 2.19e-17
96.9 | -7.1e-17 | 1e-18
97 | -7.75e-17 | 1e-18
97.1 | -7.61e-17 | 1e-18
97.2 | -3.6e-17 | 1e-18
97.3 | -6.17e-17 | 1e-18
97.4 | -5.63e-17 | 1e-18
97.5 | -8.19e-18 | 1e-18
97.6 | 5.18e-18 | 5.18e-18
97.7 | -1.79e-17 | 1e-18
97.8 | -6.94e-17 | 1e-18
97.9 | 5.58e-17 | 5.58e-17
98 | -5.07e-17 | 1e-18
98.1 | -1.83e-17 | 1e-18
98.2 | -6.4e-17 | 1e-18
98.3 | -1.15e-16 | 1e-18
98.4 | -7.45e-17 | 1e-18
98.5 | -6.15e-17 | 1e-18
98.6 | -1.02e-16 | 1e-18
98.7 | -7.18e-17 | 1e-18
98.8 | -6.48e-17 | 1e-18
98.9 | -5.6e-17 | 1e-18
99 | 1.54e-17 | 1.54e-17
99.1 | -2.97e-17 | 1e-18
99.2 | -1.38e-17 | 1e-18
99.3 | -2.36e-17 | 1e-18
99.4 | -7.65e-17 | 1e-18
99.5 | -2.5e-17 | 1e-18
99.6 | -6.57e-17 | 1e-18
99.7 | -1.14e-16 | 1e-18
99.8 | -8.8e-17 | 1e-18
99.9 | 5.9e-17 | 5.9e-17
100 | 1.42e-17 | 1.42e-17
|  | 7.7552e-12 | 2.7876e-16

## Optical Contam
Sample Test:
| absorption | 0 | ± | 0.2 | ppm/yr
| scatter | 5.4 | ± | 1.2 | ppm/yr
| area |  |  | 56.25 | cm^2
| max. normalized absorption |  |  | 0.0035556 | ppm/yr/cm^2
| max. normalized scatter |  |  | 0.11733 | ppm/yr/cm^2
Scaled to LIGO:
| pumping speed ratio |  |  | 0.02 | (sample test/LIGO)
| area |  |  | 150 | cm^2
| max. absorption |  |  | 0.010667 | ppm/yr | *Limit is < 0.1
| max. scatter |  |  | 0.352 | ppm/yr | *Limit is < 0.4
* The overall limit for AdL is < 1 ppm/yr absorption and < 4 ppm/yr scatter from all sources. It is assumed that ~10 sources could contribute.