LASER INTERFEROMETER GRAVITATIONAL WAVE OBSERVATORY

-LIGO-

CALIFORNIA INSTITUTE OF TECHNOLOGY

MASSACHUSETTS INSTITUTE OF TECHNOLOGY

| Document Type | DCC Number T080321-00-C | November 4, 2008 |
| --- | --- | --- |
| BSC Suspension Test Stand Electronics and Controls Quick Start Guide |
| J. Heefner |

Distribution of this draft:

This is an internal working note of the LIGO Laboratory

California Institute of Technology	Massachusetts Institute of Technology

LIGO Project – MS 18-33	LIGO Project – MS 20B-145

Pasadena, CA 91125	Cambridge, MA 01239

Phone (626) 395-2129	Phone (617) 253-4824

Fax (626) 304-9834	Fax (617) 253-7014

E-mail: info@ligo.caltech.edu	E-mail: info@ligo.mit.edu

www:  http://www.ligo.caltech.edu/

# Introduction

This quick start guide is meant to as a quick reference and guide to the BSC Suspension Test Stand electronics hardware and software. The test stand will be used at each of the AdL sites, LLO and LHO, to aid in the installation and checkout of each of the quadruple suspension systems. The test stand is built to run as a standalone, self-contained unit and does not require any connection to the outside world to operate. All equipment is located in one rack that can be moved, either by rolling or craning, to the location where the suspension to be tested is located.  Additionally, the test stand only includes those electronics that are needed for local damping and suspension parameter measurements. For example, the low noise output monitors that are essential for proper operation of the suspensions in AdL are not included in the test stand electronics since low noise operation of the quad suspension can not be achieved and is not required during initial assembly and checkout of a quadruple suspension. The software has allotted for future inclusion of the Electro-static driver hardware which may be installed if the test stands are used to checkout and test full quadruple suspensions installed in the AdL vacuum system. Detailed descriptions and documentation for each of the individual modules can be found in the schematics and test procedures for those components.

The figure below is a basic block and signal flow diagram for the installed system.

Each of the components is described in the sections below.

# Electronics Hardware

## System Wiring Diagram and Rack layout

The system wiring diagram can be found in the LIGO DCC or at:

http://www.ligo.caltech.edu/~jay/drawings/D080273-00-C.pdf

The figure below shows the rack layout for the system.

## UK Satellite Amplifier

The satellite amplifier provides two main functions in the system:

- Constant 35 mA bias current for the OSEM LEDs.

- Gain stage for the OSEM photodiode. The gain of the stage is 309 Kohms.

The coil driver signals from the coil driver to the OSEM coils only pass through the satellite amplifier. As can be seen from the figure above, the satellite amplifiers are located in the field near the quadruple suspension. Note that the design requires that the vacuum flange, or similar connector be included in the cabling from the satellite amplifier to the osems. This is because the vacuum flange actually mirrors the connector pins as they are passed through the flange. If the flange is not installed in the test setup, then an adapter cables that mirrors the pins must be used.

## UK TOP Coil Driver

The figure below shows the schematic for a single channel of the UK Top Coil Driver module.

The nominal response of the circuit is TBD the final design of the TOP coil driver.

## UK UIM Coil Driver

The figure below shows the schematic for a single channel of the UK UIM Coil Driver module.

The nominal response of the circuit is TBD the final design of the UIM coil driver.

## UK PUM Coil Driver

The figure below shows the schematic for a single channel of the UK PUM Coil Driver module.

The nominal response of the circuit is TBD the final design of the PUM coil driver.

## Anti-Alias and Anti-Image Chassis

The anti-alias and anti-image filters used for the BSC suspension test stand are the AdL standard 3rd order, 10KHz Butterworth low pass filters with a notch at 65536 Hz. The fundamental sample frequency for the system is 65536 SPS and the data is software decimated to 16384 SPS for the controls. The figure below shows the schematic for a single channel of the anti-alias/anti-image filter.

The schematic for the individual filter boards used in the AA and AI chassis can be found in the DCC and at:

http://www.ligo.caltech.edu/~jay/drawings/D070081-01-C.pdf

Photos of the fully assembled AA and AI chassis can be found in section 4.2 below. Note that the fully assembled AA and AI chassis contain a number of filter boards (4 and 2 respectively), a chassis power board (D060279) and an interface board that is used to match the signals to/from the ADCs and DACs into the appropriate connectors used inside the chassis.

The schematic for the chassis power board can be found in the DCC and at:

http://www.ligo.caltech.edu/~jay/drawings/D060279-00-C.pdf

The schematic for the ADC interface board can be found in the DCC and at:

http://www.ligo.caltech.edu/~jay/drawings/D070100-01-C.pdf

The schematic for the DAC interface board can be found in the DCC and at:

http://www.ligo.caltech.edu/~jay/drawings/D070101-01-C.pdf

Note that the jumper used to connect the ADC and DAC signal grounds to the AA/AI chassis signal ground on both interface boards are nominally installed. If these jumpers are not installed, excess noise can be induced into the signal chain.

## ADC and DAC Modules

The ADCs and DACs used in the controls for the BSC SUS Test Stand are the same ADC and DAC modules used in a variety of control systems at LASTI and elsewhere in LIGO. They are both manufactured by General Standards Corporation. The user manuals can be found at:

- ADC- http://www.generalstandards.com/user-manuals/pmc66_16ai64ssa_c_man_082006.pdf

- DAC- http://www.generalstandards.com/user-manuals/pmc66_16ao16_man_090406.pdf

## Rack DC Power

DC power for the chassis installed is supplied using commercial Kepco power supplies (model JQE 25-10M) mounted in the rack on a Kepco model RA24 shelf. Additionally, a DC power distribution bus chassis (D070124-A) is mounted in the top rear of the rack. AdL standard three conductor cable and connectors are used to distribute power from the bus chassis to the individual electronics chassis in the rack.

# Control Software

## Simulink Diagram

The control software for the BSC SUS Test Stand was generated using Simulink and the real-time and Linux compilers used for all CDS control systems. The top level Simulink diagram for the system is shown in the figure below.

The figure below shows the inside of the Q1 subsystem block.

As can be seen from the figure, the OSEM channels for each stage feed subsystem blocks for the corresponding stages. Note that each subsystem stage also provides a watchdog output that is summed and used as an interlock for the BSC seismic controls.

The figure below shows the control diagram for the M0 stage. The control diagrams for the other stages are similar.

## Code Generation and Operating the Front End

### Code Generation

All code for the front end is generated by compiling user generated simulink diagrams like those shown above. Use of the code generator is described in LIGO document T080135-00-C, “AdvLigo CDS Realtime Code Generator (RCG) Application Developer’s Guide”.

### Operating the Front End

The Test Stand has been designed and built to be a standalone self-contained unit that does not require the user to be an expert in the operation of real-time linux control and data acquisition systems. Normal starting and stopping of the front end and configuration of the DAQ system is controlled via scripts. These scripts are in the /cvs/cds/llo/scripts directory.

“startqts” is used to start or restart the front end as necessary. This script also starts the frameBuilder.

“killqts” is used to stop the front end

“daqconfig” is used to start the GUI used to set and remove DAQ signals in the data stored by the system.

## Operator Screens

MEDM is used to create and display data on operator screens. The operator screens for the test stand can be found in the directory /cvs/cds/llo/medm/l1/qts on the control computer. In order to start medm login into the test stand computer as user controls and type “medm” at the command prompt. The main control screen for the test stand is named L1QTS_QUAD1.adl and is shown in the figure below. From this screen the user can see an overview of the controls and reach all other operator screens.

In addition to all filter module and matrix screens needed for controls, there are several other very useful screens that are shown and described below. The first screen, “L1QTS_GDS_TP.adl” can be reached by clicking on the FE Diagnostics button in the bottom left corner of the main control screen. A picture of this screen is shown in the figure below. On this screen, the operator can check the status of the frameBuilder, the CPU computation time, monitor ADC and DAC overflows and monitor which test points have be set. Another important feature of this screen is the BURT Restore field in the bottom left corner of this screen. In order for the front end to begin execution a “1” must be entered into this field by the operator. If a “1” is not entered into this field the front end code will not begin execution. Under normal operating conditions, the test stand settings would be restored using the backup and restore tool (BURT) prior to starting the front end real-time code. This would prevent the damping servos from beginning execution with values that may cause the servo to go unstable. Once the proper or a set of safe settings for the system have been restored a “1” is entered into this field by the operator and the front end real-time code will begin execution.

The next very useful screen is the watchdog screen shown in the figure below. This screen can be accessed by clicking on the Watchdog button in the bottom right portion of the main control screen.

The setpoints on this screen are the operator adjustable trip levels for the watchdog logic. The reset button is used to reset any trips that have occurred and the on/off buttons are used to enable the outputs to the suspension coil drivers. The on/off buttons will not enable the outputs if a fault has not been reset, or one the osem monitors in the fields above exceeds the setpoint. In the event of a fault condition, the background color around the on/off buttons will turn from green to red.

The screen shown below is an input signal monitor screen that has been created to monitor the 12 osem levels in the TOP level of the main and reaction chains of the quad. This screen is very useful during assembly and osem adjust and can be reached by clicking the Input Signals button in the middle of the main control screen.

## Backup and Restore Tool

The test stand includes the standard EPICS backup and restore tool BURT. This tool is started by typing the command “burtgooey” at the command prompt. When this command is typed, the following window appears.

Using this window the operator can backup current settings and restore previous settings in addition to other operations. When the backup button is pushed, the following window appears.

Using this window the operator can select the request file or files to be used and specify the name and location of the settings snapshot file that will be created.

A basic request file for the test stand has been created and stored in the directory “/cvs/cds/llo/burt/qts”. This request file stores the settings for all epics channels used by the test stand. The name of the file is “qts.req”.

## Data Display and Analysis

All of the standard data display and analysis tools used in LIGO and AdL can be used on the test stand. These tools include:

- Dataviewer—started by typing “dataviewer” at the command prompt

- Dtt—started by typing “diaggui” at the command prompt

- Foton—started by typing “foton” at the command prompt

- Awg—started by typing “awggui” at the command prompt

## Desktop Shortcuts and Icons

Shortcuts and desktop icons have been created for all the tools described in the sections above.

# Documentation List

## Schematics

BSC SUS Test Stand Wiring—D080273-00-C

AdL AA and AI Filter (3rd order 10KHz, 65536 notch for 65536 sample rate)—D070081-01-C

UK TOP Coil Driver—drawing number TBD

UK UIM Coil Driver—drawing number TBD

UK PUM Coil Driver—drawing number TBD

UK Satellite Amplifier—drawing number TBD

## System and Chassis Photos

The picture below shows the prototype UK Satellite Amplifier removed from the box. It is assumed that the pre-production prototype is similar and this picture will be updated when the units are received.

The picture below shows the prototype TOP Coil Driver chassis with the top cover removed. It is assumed that the pre-production prototype is similar and this picture will be updated when the units are received.

The picture below shows the prototype Penultimate Coil Driver Chassis with the top cover removed. It is assumed that the pre-production prototype is similar and this picture will be updated when the units are received.

The picture below shows the Anti-Alias Chassis with the top cover removed.

The picture below shows the Anti-Image Chassis with the top cover removed.

The picture below shows the IO chassis (Blue-Box) with the cover removed and two PCIX boards installed.