1	Safety	2

2	Objective and Scope	2

3	Documents	3

4	Documenting the Assembly Process	4

5	Vacuum Compatibility	5

6	Fasteners	6

7	Overview of Assembly Process	7

8	Handling Suspension Wire	8

9	Assembling Upper Wires	9

10	Assembling Intermediate Wires	12

11	Assembling Lower Loop Wire	15

12	Assembling Upper Blade Guards	18

13	Assembling the Rotational Adjusters	19

14	Assembling Magnets – Upper Mass	23

15	Assembling Upper Mass and Coil Holder (M1)	25

16	Assembling Intermediate Mass (M2)	32

17	Assembling Bottom Mass and Optic (M3)	34

18	Assembling EQ Stops For Intermediate and Bottom Masses	36

19	Preparing The Weldment	38

20	Installing the Mounting Pads	39

21	Installing EQ Stop Mounts and Brackets	40

22	Installing Rotational Adjusters and Top Blade Guards	42

23	Installing Upper Mass and Coil Holder	45

24	Installing the Intermediate Mass	47

25	Installing Wire Assemblies	48

26	Assembling AOSEM Brackets	51

27	Installing Bottom Mass	54

28	Suspension and Alignment of Masses	55

29	Assembling Magnets To Intermediate Mass	59

30	Assembling Magnets To Bottom Mass	62

31	Installing the BOSEMs	65

32	Installing AOSEM Brackets	65

33	Aligning OSEMs	67

| Fig 1: HLTS | Fig 2: Front | Fig 3: Rear |
| --- | --- | --- |
|  | Note: Front vs. Rear of Weldment is identified by 3 hole patterns. |  |

# Safety

Please review E0900332 for concerns related to safe assembly and installation.

# Objective and Scope

Subassembly and Final Assembly of the aLIGO HAM Large Triple Suspension, including:

- General considerations for assembly

- Use of the assembly fixtures shown in D080718

# Documents

D070308	HLTS Top Blade Guard Assembly

D070319	HLTS Bridge EQ Stops

D070326	HLTS Rotational Adjuster

D070334	HLTS Intermediate Mass Assembly

D070335	HLTS Upper Mass Assembly

D070337	HLTS Bottom Mass Assembly

D070340	HLTS Upper Wire Assembly

D070393	HLTS Intermediate Wire Assembly

D070436	HLTS Lower Loop Wire Assembly

D070442	HLTS Overall Assembly

D070447	HLTS Assembly Drawings

D080677	HLTS Coil Holder Assembly

D080718	HLTS Assembly Fixtures

D080726	HLTS EQ Stop, Long Mount

D080727	HLTS EQ Stop, Long Bracket

D0900626	HLTS Mounting Pad Assembly

D1001695	Magnet and Flag Assembly, BOSEM

D1002821	HLTS EQ Stop Assembly, Bottom Mass, Lower

D1102071	HLTS EQ Stop, Bridge, Upper

E080208	HLTS Assembly Instructions (this document)

E0900047	Contamination Control Plan

E1000043	HLTS Assembly and Installation Hazard Analysis

E1000045	HLTS Installation Procedure

E1100109	Suspensions Controls Arrangement

E960022	LIGO Vacuum Compatibility, Cleaning Methods and Qualification Procedures

E990196	Magnet/Standoff Assembly Preparation

E1100472	HLTS Assembly and Installation Documentation

G1100147	HLTS Introduction

M0900034	RODA - Use of Magnets in Suspensions

T000053	Advanced LIGO Universal Suspension Subsystem Design Requirements

T010007	Cavity Optics Suspension Subsystem Design Requirements

T010103	aLIGO Suspension System Conceptual Design

T1000012	HLTS Final Design Document

T1000089	HLTS Test Plan

T1100066	General Torque Recommendations for Socket Head Cap Screws

T1100003	Building Suspensions Subassemblies in ICS

# Documenting the Assembly Process

## Documents

T1100003	Building Suspensions Subassemblies in ICS.

## Procedure

See the above document.

Data for each Final Assembly will be stored in ICS; using a Process Traveler is optional:

| Item | Assembly 1 Part Name | Assembly 1 Part Number | Serial Number | Position | Variant | Weight |
| --- | --- | --- | --- | --- | --- | --- |
| Each Mass | X | X | X | X |  | X |
| Each Blade | X | X | X | X |  |  |
| Each Blade Clamp | X | X | X | X | X |  |
| Each OSEM | X | X | X | X |  |  |
| Each Optic | X | X | X |  |  | X |

Note regarding Subassembly weights:  Each Subassembly must have 3 distinct weights recorded:

- Estimated Weight 	Calculated by SolidWorks;

- Actual Weight	Measured by a lab scale after built to the nominal mass;

- Balanced Weight	Totaled after Suspension is balanced (i.e. Actual Weight + Add-On Weights).

When Addable weights are used, note their location on the Mass.

# Vacuum Compatibility

## General Handling

All procedures must be performed in a clean room environment while suited up in:

- Coverall with Hood

- Boot style shoe covers

- LIGO-approved latex gloves

- Glove Liners and Safety Glasses when working with Wire

- All Tables surfaces used for Class A components should be wiped down daily with Isopropanol.

- Review E0900047 Contamination Control Plan for details. All HSTS parts are Class A hardware and once cleaned and baked should not come into contact with anything but Class A or B hardware.

## Cleaning Components

Clean items per E960022.

## Inspection

After baking, sample check the cleanliness of blind-tapped and through-tapped holes with a clean swab dampened with alcohol for a minimum of 10% of the holes in case any material has leached out during baking. If any discoloration of the swab is evident, then the part must go through at least one more wash cycle before repeating the bake.  After inspection, wrap items per E960022.

# Fasteners

## Documents

T1100066	General Torque Recommendations for Socket Head Cap Screws

## Helicoils

Helicoils are specified for:

- Certain SSTL applications to avoid using AgPlated fasteners;

- Certain applications where assembly / disassembly recurs.

Helicoils are cleaned, baked and installed with Class B clean tools in a Class 100 clean room.

## Silver Plated Stainless Steel

All Silver Plated fasteners are also SSTL, and so are labeled simply “AgPlated”, not “AgPlated SSTL”.

## Jam Nuts

All ¼-20 Jam (thin) Nuts have been replaced with ¼-20 AgPlated SSTL Hex Nuts (standard thickness).

## Screw Applications

| Screw Type | Screw Description | Receiver Application |
| --- | --- | --- |
| AgPlated SSTL | Silver-Plated Stainless Steel | - Stainless Steel threads |
| SSTL | Stainless Steel | - Aluminum Threads - Helicoil Threads |
| Vented | SSTL Screw with holes | - Rare Vacuum Compatibility situations |

## Torque Values

- Except where noted, Socket Head Cap Screws are to be tightened per the following table.

- “Generic” applies to Screws that are non-plated, non-vented, and not marked as Holokrome.

- Holokrome Screws are marked as such on the Screw.

- UC (UC Components, Supplier) Screws are AgPlated.

- All Screws are UNC (coarse threaded).

|  | Torque (in-lb) |  |  |
| --- | --- | --- | --- |
| Supplier | Generic (unmarked) | Holokrome UC | UC |
| Type | Unplated | Unplated AgPlated | Vented |
| Size |  |  |  |
| 2 | 3 | 4 | 3 |
| 4 | 5 | 6 | 7 |
| 8 | 20 | 30 | 25 |
| ¼” | 75 | 100 | 86 |
| 5/16” | 132 | 184 |  |

# Overview of Assembly Process

## General sequence:

- Prepare HLTS Structure (D070537)

- Assemble subassemblies, in any order:

  - Top Blade Guard Assemblies (2X D070308)

  - Earthquake Stop Assemblies (8X D080726 and 8X D080727 and 2X D080728 and 4X D1102071 and 2X D1002821)

  - Coil Holder Mounting Brackets (4X D1002133)

  - AOSEM Alignment Assemblies (2X D0902024 and 2X D0901551 and 2X D0901552 and 2X D0901553)

  - Rotational Adjusters (2X D070326)

  - Upper Mass Assembly (1X D070335)

  - Intermediate Mass Assembly (1X D070334)

  - Metal Bottom Mass Assembly (1X D070337)

- Attach Top Blade Guard Assemblies to Structure

- Attach Rotational Adjusters to Structure and flatten Upper Blades

- Attach Earthquake Stop Assemblies to Structural Weldment

- Install Intermediate and Metal Lower Mass Assemblies

- Assemble Intermediate Wire Assemblies (4X D070393) and attach to Upper Mass Assembly

- Assemble Upper Wire Assemblies (2X D070340) and attach to Upper Mass Assembly

- Place Coil Holder/Tablecloth (D070449) over Upper Mass Assembly and lock the two together

- Attach Upper Mass/Coil Holder Assembly (D1001754) to Structure

- Connect Upper Wire Assemblies to Upper Blades

- Connect Intermediate Wire Assemblies to Intermediate Mass Assembly

- Assemble Lower Loop Wire Assembly (D070436)

- Attach Lower Loop Wire Assembly to Intermediate Mass Assembly

- Suspend all masses

- Initial balancing

- Remove all masses and Rotational Adjusters

- In parallel:

  - Creep baking of Upper Blades (in Rotational Adjusters) and Lower Blades (in Upper Mass Assembly)

  - Magnet gluing for Intermediate Mass and Metal Bottom Mass Assemblies

- Reinstall Rotational Adjusters, Wires and Masses.

- Rebalancing.

- Install AOSEM Alignment Brackets.

- Install BOSEMs on Coil Holder/Tablecloth.

- Metal-Build Testing

- Transport HLTS to chamberside using a storage container.

- Metal-Build Testing, Part 2

- Replace Metal Bottom Mass with Glass Optic

- Rebalancing

- Glass-Build Testing

- Install HSTS into chamber

- In-Chamber Testing

# Handling Suspension Wire

The Wire used for all Suspensions is a hard temper carbon steel, delivered on large spools.  When unwound for cleaning, cutting and preparation for clamp-wire-clamp assembly, care must be taken such that the wire’s strong potential energy (making it act like a coiled spring) does not cause injury.

- Safety Glasses, provided in all Clean Room garbing areas, must be worn during all wire work.

Glove Liners must be worn under the latex clean-room gloves as a protective layer and extra barrier.  The E0900047 Contamination Control Plan, p. 13, provides further information on Glove Liners.

For easier holding, bend a small section (~3”) of the end of the Wire. The bent section can be hooked around your thumb and held by your index finger.  Un-spool the proper length of Wire – including extra for handling – and control the area of the Wire that will be cut.  Add a 2nd bend at the newly cut end for easier handling.

Change your gloves and wipe each Wire at least 3 times each, and until no residue appears on each Wipe, using:

  - A Cleaning Wipe with Methanol;

  - A Cleaning Wipe with Acetone;

  - A Cleaning Wipe with Isopropanol;

changing Wipes until the wire is completely clean.  Clean the Wire while it is coiled; do not stretch the wire until it is taut for cleaning. It can be laid down on a clean surface during this process. Clean one section at a time.

Transfer the Wire to the Assembly Jig. Use the Jig clamps to hold the Wire in place, and then cut off and discard the bent Wire ends.

# Preparing The Weldment

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 8 | Ea | 1185-2EN492 | Helicoil 8-32 x 3.0D |
| 8 | Ea | D980184 | LOS Clamps |
| 8 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.5” AgPlated |

## Procedure

| - Identify the Front vs Rear of the Weldment. The front has 4 pairs of adjacent, large-diameter through-holes. Install at the base of the Weldment: 8 1185-2EN492 Helicoils - 8-32 x 3.0D - Install these BEFORE securing the Weldment to the Optical Table! | Fig 70: Helicoils in Weldment |
| --- | --- |

| Secure the D070442 Structural Weldment to the Optical Table with: 8 D980184 LOS Clamps, 2 per corner 8 Socket Head Cap Screws - ¼-20 x 1.5” AgPlated | Fig 71: Securing the Weldment |
| --- | --- |

# Installing the Mounting Pads

## Documents

D0900626	HLTS Mounting Pad Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D070442 | HLTS Structural Weldment |
| 8 | Ea | D980184 | LOS Clamps |
| 2 | Ea | D070374 | HLTS Mounting Pad |
| 8 | Ea | NA | Helicoil 8-32 x 0.246” |
| 4 | Ea | 1185-2EN246 | Helicoil 8-32 2B x 2.0D |
| 2 | Ea | D0900628 | HLTS Mounting Pad Side Bracket |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” SSTL |
| 14 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” AgPlated |
| 18 | Ea | NA | Flat Washer #8 SSTL |

## Procedure

| - Assemble to each of the 2 D070374 Mounting Pads: Top of the Pad 4 Helicoil 8-32 x 0.246 or longer Side of the Pad 2 1185-2EN246 Helicoils 8-32 2B x 2.0D | Fig 72: Helicoils in 1 of 2 Mounting Pads |
| --- | --- |
| Create 2 D0900626 Mounting Pad Assemblies: 2 D070374 Mounting Pad Body 2 D0900628 Mounting Pad Side Bracket 4 Socket Head Cap Screw - 8-32 x 0.5” SSTL 4 Flat Washer #8 SSTL - Torque to 20 in-lb | Fig 73: Mounting Pad Assembly |
| Assemble the Mounting Pad Assembly to the Structural Weldment using: 14 Socket Head Cap Screw - 8-32 x 0.5” AgPlated 14 Flat Washer #8 SSTL - Torque to 30 in-lb - The Serial Number for each Pad is stamped on a specific Weldment corner. Ensure each Pad is mounted on the correct Weldment location. | Fig 74: Mounting Pad on Weldment |

# Assembling Top Blade Guards

## Documents

D070308	HLTS Top Blade Guard Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 4 | Ea | D070309 | Top Blade Guard Riser |
| 2 | Ea | D070310 | Top Blade Guard Bar |
| 8 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.0” SSTL |
| 8 | Ea | NA | Flat Washer, ¼” Vented SSTL |
| 4 | Ea | D0900999 | Socket Head Cap Screw ¼-20 x 2.0” SSTL Round-Tip |
| 4 | Ea | NA | Hex Nut ¼-20 AgPlated |

## Procedure

| - Assemble 2 D070308 Top Blade Guard Assemblies, each with: 2 D070309 Top Blade Guard Riser 1 D070310 Top Blade Guard Bar 4 Socket Head Cap Screws - ¼-20 x 1.0” SSTL 4 Flat Washers, ¼-20 Vented SSTL 2 D0900999 Socket Head Cap Screw - 2 ¼-20 x 2.0” SSTL Round-Tip 2 Hex Nut ¼-20 AgPlated | Fig 28: Top Blade Guard Assembly |
| --- | --- |

# Assembling EQ Stops For Intermediate and Bottom Masses

## Documents

D070447	HLTS Overall Assembly

D080726	HLTS EQ Stop, Long Mount

D080727	HLTS EQ Stop, Long Bracket

D1002821	HLTS EQ Stop Assembly, Bottom Mass, Lower

D1102071	HLTS EQ Stop, Bridge, Upper

D070319	HLTS Bridge EQ Stops

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 10 | Ea | D070460 | Long EQ Stop Mount |
| 8 | Ea | D070322 | Long EQ Stop Bracket |
| 4 | Ea | D1002823 | EQ Stop Bracket Back, Bottom Mass, Lower |
| 4 | Ea | NA | Socket Head Cap Screws 8-32 x 1.0” SSTL |
| 4 | Ea | NA | Flat Washer #8 Vented SSTL |
| 4 | Ea | D070321 | EQ Stop Bridge Crossbar |
| 8 | Ea | D1102072 | Bridge EQ Stop Mount, Upper |
| 28 | Ea | D080725 | EQ Stop for Metal or Glass |
| 28 | Ea | NA | Hex Nut ¼-20 AgPlated |
| 8 | Ea | NA | Socket Head Cap Screw ¼-20 x 0.75” SSTL |
| 8 | Ea | D1100785 | Flat Washer, ¼”, Nitronic 60 |

## Procedure

| - Assemble 10 D080726 Long Mount EQ Stops, each using: 1 D070460 Long EQ Stop Mount 1 D080725 EQ Stop for Metal or Glass 1 Hex Nut ¼-20 AgPlated | Fig 62: D080726 Long Mount EQ Stop |  |  |
| --- | --- | --- | --- |
| - Assemble 8 D080727 Long Bracket EQ Stops, each using: 1 D070322 Long EQ Stop Bracket 1 D080725 EQ Stop for Metal or Glass 1 Hex Nut ¼-20 AgPlated | Fig 63: D080727 Long Bracket EQ Stop |  |  |
| - Assemble 2 D1002822 EQ Stop Bracket Assemblies, 1 LH and 1 RH version, each using: 1 D1002823 EQ Stop Bracket, Back 1 D1002824 EQ Stop Bracket, Side 2 Socket Head Cap Screw - 8-32 x 1.00” SSTL 2 Flat Washer #8 SSTL Vented - Assemble 2 D1002821 EQ Stop Assemblies, 1 LH and 1 RH version, each using: 1 D1002822 EQ Stop Bracket (see above) 1 D080725 EQ Stop for Metal or Glass 1 Hex Nut ¼-20 AgPlated | Fig 64: D1002822 Left |  | Fig 65: D1002822 Right |
|  | Fig 66: D1002821 Left |  | Fig 67: D1002821 Right |
| - Assemble 4 D1102071 Bridge EQ Stops in 2 configurations. - Initial Build: Upper RH diagram: 2 above Metal; Lower RH diagram: 2 beneath Metal. - Optic Installation: Reconfigure existing assembly (lower RH diagram) as upper RH diagram. - Assemble all 4 with: 1 D070321 Crossbar 2 D1102072 Mount 2 Socket Head Cap Screw - ¼-20 x 0.75” SSTL 2 D1100785 Flat Washer, ¼” Nitronic 60 2 D080725 EQ Stop for Metal or Glass - The tips (shown red & black) are not used for the Metal Mass 2 Hex Nut ¼-20 AgPlated - Hand-tighten all fasteners at this point |  | Tips not used for Metal Mass Fig 68: Above Metal / Optic & beneath Optic Fig 69: Use beneath the Metal Mass |  |

# Assembling Coil Holder Brackets

# Assembling AOSEM Brackets

1 of each configuration is required.  The difference between LH / RH is simply the orientation of the Mounting and / or Alignment Brackets.

| D0902024 AOSEM Alignment Assembly Upper Side of Intermediate Mass | Fig 103: D0902024 | Alignment Bracket Mounting Bracket Fig 104: LH Config | Fig 105: RH Config |
| --- | --- | --- | --- |
| D0901551 AOSEM Alignment Assembly Lower Side of Intermediate Mass | Fig 106: D0901551 | Fig 107: LH Config | Fig 108: RH Config |
| D0901552 AOSEM Alignment Assembly Upper Side of Bottom Mass | Fig 109: D0901552 | Fig 110: LH Config | Fig 111: RH Config |
| D0901553 AOSEM Alignment Assembly Lower Side of Bottom Mass | Fig 112: D0901553 | Fig 113: LH Config | Fig 114: RH Config |

## Documents

D070442	HLTS Overall Assembly

## Materials

| D0902024 |  |  |  |
| --- | --- | --- | --- |
| Qty | U | ID | Description |
| 1 | Ea | D0902025 | HLTS AOSEM Mounting Bracket, Upper Side of Intermediate Mass |
| 1 | Ea | D0901492 | HLTS AOSEM Alignment Bracket, Intermediate Mass |
| 1 | Ea | D0901065 | AOSEM Assembly |
| 1 | Ea | D1002858 | AOSEM Adjustment Collar, Reverse Side |
| 2 | Ea | D1000659 | Adjuster Shaft |
| 2 | Ea | D1002865 | Adjustment Nut, Thick |
| 2 | Ea | NA | Socket Head Cap Screw, 8-32 x 0.625”, SSTL |
| 2 | Ea | 1185-2EN246 | Helicoil, 8-32 x 0.246” |
| 2 | Ea | NA | Flat Washer, Vented, #8 |
| 1 | Ea | NA | Socket Head Cap Screw, 2-56 x 0.375” SSTL |
| 1 | Ea | D030021 | Socket Head Cap Screw, ¼-20 x 2.5”, Round Tip SSTL |
| 1 | Ea | 1185-4EN250 | Helicoil, ¼-20 x 0.25” |
| 1 | Ea | NA | Hex Nut, ¼-20 AgPlated |
| D0901551 |  |  |  |
| Qty | U | ID | Description |
| 1 | Ea | D0901493 | Lower AOSEM Mounting Bracket, Intermediate Mass |
| 1 | Ea | D0901492 | HLTS AOSEM Alignment Bracket, Intermediate Mass |
| 1 | Ea | D0901065 | AOSEM Assembly |
| 1 | Ea | D1002858 | AOSEM Adjustment Collar, Reverse Side |
| 2 | Ea | D1000659 | Adjuster Shaft |
| 2 | Ea | D1002865 | Adjustment Nut, Thick |
| 2 | Ea | NA | Socket Head Cap Screw, 8-32 x 0.625”, SSTL |
| 2 | Ea | 1185-2EN246 | Helicoil, 8-32 x 0.246” |
| 2 | Ea | NA | Flat Washer, Vented, #8 |
| 1 | Ea | NA | Socket Head Cap Screw, 2-56 x 0.375” SSTL |
| 1 | Ea | D030021 | Socket Head Cap Screw, ¼-20 x 2.5”, Round Tip SSTL |
| 1 | Ea | 1185-4EN250 | Helicoil, ¼-20 x 0.25” |
| 1 | Ea | NA | Hex Nut, ¼-20 AgPlated |
| D0901552 |  |  |  |
| Qty | U | ID | Description |
| 1 | Ea | D0901550 | AOSEM Mounting Bracket, Upper Position of Bottom Mass |
| 1 | Ea | D0901549 | AOSEM Alignment Bracket, Bottom Mass |
| 1 | Ea | D0901065 | AOSEM Assembly |
| 1 | Ea | D1002858 | AOSEM Adjustment Collar, Reverse Side |
| 2 | Ea | D1000659 | Adjuster Shaft |
| 2 | Ea | D1002865 | Adjustment Nut, Thick |
| 1 | Ea | D080725 | EQ Stop for Metal or Glass |
| 2 | Ea | NA | Socket Head Cap Screw, 8-32 x 0.625”, SSTL |
| 2 | Ea | 1185-2EN246 | Helicoil, 8-32 x 0.246” |
| 2 | Ea | NA | Flat Washer, Vented, #8 |
| 1 | Ea | NA | Socket Head Cap Screw, 2-56 x 0.375” SSTL |
| 1 | Ea | 1185-4EN250 | Helicoil, ¼-20 x 0.25” |
| 1 | Ea | NA | Hex Nut, ¼-20 AgPlated |
| D0901553 |  |  |  |
| Qty | U | ID | Description |
| 1 | Ea | D0900340 | AOSEM Mounting Bracket, Lower Position of Bottom Mass |
| 1 | Ea | D0901549 | AOSEM Alignment Bracket, Bottom Mass |
| 1 | Ea | D0901065 | AOSEM Assembly |
| 1 | Ea | D1002858 | AOSEM Adjustment Collar, Reverse Side |
| 2 | Ea | D1000659 | Adjuster Shaft |
| 2 | Ea | D1002865 | Adjustment Nut, Thick |
| 1 | Ea | D080725 | EQ Stop for Metal or Glass |
| 2 | Ea | NA | Socket Head Cap Screw, 8-32 x 0.625”, SSTL |
| 2 | Ea | 1185-2EN246 | Helicoil, 8-32 x 0.246” |
| 2 | Ea | NA | Flat Washer, Vented, #8 |
| 1 | Ea | NA | Socket Head Cap Screw, 2-56 x 0.375” SSTL |
| 1 | Ea | 1185-4EN250 | Helicoil, ¼-20 x 0.25” |
| 1 | Ea | NA | Hex Nut, ¼-20 AgPlated |

## Procedure

Assembly procedure is similar for all 4 units, but varies by the Mount Bracket and Alignment Bracket.

| - Assemble Adjustment Collar to AOSEM - Assemble SHCS 2-56 to Collar - Assemble Adjuster Shafts to Collar - Assemble AOSEM w/Adjuster Shafts to Alignment Bracket, paying close attention to part number and orientation of Bracket - Assemble Mount to Alignment Bracket - Assemble EQ Stop to Alignment Bracket with a Hex Nut | 5 4 3 6 2 1 Fig 115: Assembly |
| --- | --- |

# Assembling the Rotational Adjusters

## Documents

D070326	HLTS Rotational Adjuster

E1000169	Blade Characterization Spreadsheet

T0900559	Blade Pairings Spreadsheet

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D080265 | HLTS Upper Blade Bake Fixture |
| 1 | Ea | D070327 | HLTS RA Base Plate |
| 1 | Ea | D070328 | HLTS RA Rotating Plate |
| 3 | Ea | NA | Socket Head Cap Screw ¼-20 x 0.625” SSTL |
| 3 | Ea | NA | Flat Washer ¼” |
| 1 | Ea | D070329 | HLTS RA Pull Plate |
| 2 | Ea | NA | Socket Head Cap Screw 8-32 x 1.0” AgPlated |
| 1 | Ea | D070330 | HLTS RA Push Plate |
| 2 | Ea | NA | Socket Head Cap Screw 8-32 x 1.25” AgPlated |
| 1 | Ea | NA | Socket Head Cap Screw 8-32 x 1.5” Fully-Threaded SSTL |
| 1 | Ea | NA | Flat Washer #8 Vented SSTL |
| 1 | Ea | NA | Socket Head Cap Screw 8-32 x 1.5” Full-Thread Round-Tip SSTL |
| 3 | Ea | NA | Socket Head Cap Screw 5/16-18 x 1.75” AgPlated |
| 3 | Ea | NA | Flat Washer 5/16” Vented SSTL |
| 1 | Ea | D0900665 | HLTS Upper Blade Library of Clamps |
| 2 | Ea | D020617 | Upper Blade |
| 2 | Ea | Various | Upper Blade Clamp Inside |
| 2 | Ea | D070331 | Upper Blade Clamp Shim, 1.0mm |
|  |  |  | Consult with a Suspension Design Engineer to verify that 1.0 is appropriate; Shim height needed depends upon actual Weldment height. |
| 1 | Ea | D020660 | Blade Pulldown Device |
| 1 | Ea | NA | Machinist’s Square |

## Procedure

| Prepare a D020660 Blade Pulldown Device with: 1 D0901814 Upper Clamp, Inside 1 D070341 Upper Clamp, Outside Socket Head Cap Screw - 8-32 x 0.5” AgPlated Flat Washer #8 Vented, SSTL 2’ of 0.024” Wire. | Fig 29: Blade Pulldown Device |
| --- | --- |
| - Remove the Bridges and Side Supports from the D080268 Upper Blade Bake Fixture Base Plate. Mount the Base Plate to an Optics Table, leaving one end of the Plate extending beyond the Table Edge to accommodate the Blade Pulldown Device.Assemble 2 D070328 Rotational Adjuster Rotating Plates to the Fixture Base Plate using: 4 Socket Head Cap Screws - ¼-20 x .625” SSTL 4 Flat Washers ¼” Vented SSTL - Tighten the Screws firmly. - Square the edge of the Rotational Adjuster Rotating Plate against the Fixture Stop Block. Assemble the D1003307 Hold Down Bar to the Fixture using: 2 Socket Head Cap Screws - ¼-20 x 1.00” SSTL 2 Flat Washer ¼” Vented SSTL - Tighten the Screws firmly. | Table Edge Bridges and Supports Stop Block Fig 30: Rotating Plates & Hold Down Bar |
| Select a matching pair of Upper Blades and Clamps per the E1000169 Blade Characterization Spreadsheet: Blade launch angle is set by Blade Clamps. These range from 0-3.5 deg. in .5 deg. increments. Select Clamps from the D0900665 HLTS Upper Blade Library of Clamps Select Clamps according to Blade Characterization data for stiffness and expected load. Select Blades in pairs according to Blade Characterization data. Align the Blade, Clamps and Shim using a Precision Square. Record the Blade serial numbers and Blade clamp angles and orientations within ICS. | Fig 31: Clamps Control Launch Angle Fig 32: Upper Blade Library of Clamps |
| Assemble to the Rotational Adjuster Blade Rotating Plate furthest from the Table edge: 3 Socket Head Cap Screws - 5/16-18 x 1.75” AgPlated 3 Flat Washers 5/16” Vented SSTL D09006XX Upper Blade Clamp Outside D020617 Upper Blade D09006XX Upper Blade Clamp Inside D070331 Upper Blade Clamp Shim 1.0mm - Consult with a Suspension Design Engineer to verify that 1.0mm is appropriate; Shim height needed depends upon actual Weldment height. - Using the Machinist’s Square, Square the Blade, Clamps, and Shim to - each other and to the Rotating Plate and hand-tighten the 3 Screws Hang the 18.252 kg Pulldown Weight from the Blade Tip Assemble to the Bake Fixture Base Plate: 2 D1003306 Side Support 1 D080267 Top Bridge 4 Socket Head Cap Screws - ¼-20 x 2.5” SSTL - Tighten the Screws firmly D0900999 Socket Head Cap Screw - ¼-20 x 2.0” Round-Tip Hex Nut ¼-20 AgPlated Turn down the Round Tip Screw until the weighted Blade tip is level with the Blade root. - Be careful not to damage the nickel plating on the blade | Table Edge Machinists Square Fig 33: Blade & Clamps Assembled |
| Leaving the Wire Clamp attached to the Blade, remove the rest of the Blade Pulldown Device. Tighten the 5/16” Screws to 184 in-lb. Re-attach the Blade Pulldown Device to the Wire Clamp. Turn back the Round Tip Screw and remove the Top Bridge and Side Supports. Slowly lift and then remove the Blade Pulldown Device, allowing the Blade to curve fully upward. Detach the Bake Fixture from the Optics Table, rotate it 180 deg., and remount it on the Optics Table, leaving one end of the Plate extending beyond the Table Edge to accommodate the Blade Pulldown Device. Repeat steps 7-15 to assemble the second Upper Blade to the Bake Fixture. Remove the Hold-down Bar. Disassemble both Rotational Adjusters from the Upper Blade Bake Fixture. |  |
| Assemble the D070328 Rotating Plate to the D070327 HLTS Rotational Adjuster Base Plate using: 3 Socket Head Cap Screw - ¼ – 20 x 0.625” SSTL 3 Flat Washer ¼” - Hand-tighten only, for now. | Base Plate Rotating Plate Fig 34: Attaching Base Plate |
| Assemble to the Base Plate: 1 D070329 Pull Plate 2 Socket Head Cap Screws - 8-32 x 1.0” AgPlated 1 D070330 Push Plate 2 Socket Head Cap Screws - 8-32 x 1.25” AgPlated 1 Socket Head Cap Screw - 8-32 x 1.5” Fully Threaded SSTL 1 Flat Washer #8 Vented SSTL 1 Socket Head Cap Screw - 8-32 x 1.5” Fully Threaded, Round Tip SSTL | 8-32 x 1.5” Round Tip 8-32 x 1.5” & Washer 8-32 x 1.25” 8-32 x 1.0” Pull Plate Push Plate Fig 35: Assembling Push & Pull Plates |

# Assembling Upper Mass and Coil Holder (M1)

## Documents

D070335	Upper Mass Assembly

D080677	Coil Holder Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D040259 | Upper Mass Jig |
| 1 | Ea | D020605 | HLTS Upper Mass Main Section |
| 1 | Ea | D020607 | Upper Mass T-Piece |
| 8 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.25” AgPlated SSTL |
| 1 | Ea | D1001669 | Pitch Insert |
| 1 | Ea | NA | Socket Set Screw ½-20 x 3.5” AgPlated |
| 2 | Ea | NA | Socket Set Screw ¼-20 x 2.0” Fully-Threaded SSTL |
| 2 | Ea | D080221 | HLTS Lower Blade Guards |
| 4 | Ea | D030025 | Socket Head Cap Screws 8-32 x 1.0” SSTL Fully-Threaded Round-Tip |
| 8 | Ea | NA | Hex Nuts 8-32 AgPlated |
| 4 | Ea | NA | Socket Head Cap Screws 4-40 x 0.375” AgPlated SSTL |
| 12 | Ea | NA | Flat Washers #8 Vented SSTL |
| 4 | Ea | Various | HLTS Blade Clamp, Lower Blade, Inside |
| 4 | Ea | D020615 | Lower Blade |
| 4 | Ea | Various | HLTS Blade Clamp, Lower Blade, Outside |
| 10 | Ea | NA | Flat Washer ¼” SSTL |
| 4 | Ea | D020653 | Screwdrive Block |
| 8 | Ea | NA | Socket Head Cap Screw 8-32 x 1.0” AgPlated |
| 4 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.125” AgPlated |
| 1 | Ea | D030139 | Roll Offset, Upper Mass T-Piece |
| 2 | Ea | NA | Socket Set Screw ¼-20 x 0.5” AgPlated |
| 4 | Ea | NA | Socket Head Cap Screw ¼-20 x 2.0” SSTL |
| 4 | Ea | NA | Hex Nut ¼-20 AgPlated |
| 9 | Ea | D1001699 | Magnet Holder, HLTS BOSEM |
| 18 | Ea | NA | Socket Head Cap Screw 4-40 x 0.625” AgPlated |
| 9 | Ea | D1001695 | Magnet / Flag Assembly |
| 2 | Ea | D070340 | HLTS Upper Wire Assemblies |
| 4 | Ea | D070393 | HLTS Intermediate Wire Assemblies |
| 1 | Ea | NA | Machinist’s Square |
|  |  |  | Components Added Only for Weighing |
| 2 | Ea | D020652 | Upper Mass C-Clamp |
| 2 | Ea | D020610 | Inside Upper Wire Lower Clamp |
| 2 | Ea | D020624 | Outside Upper Wire Lower Clamp |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” AgPlated |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 1.125” SSTL |
| 4 | Ea | NA | Flat Washer #8 SSTL |
| 1 | Ea | D070449 | HLTS Coil Holder |
| 8 | Ea | NA | Socket Head Cap Screw 4-40 x 0.25” AgPlated |
| 4 | Ea | NA | Flat Washer #4, Vented, SSTL |
| Qty | U | ID | Description |
| 2 | Ea | D020660 | Lower Blade Pulldown Device 6.092 kg Hang Weight Lower Blade Wire Clamp from Intermediate Wire Assembly 8 Socket Head Cap Screw - 4-40 x 0.25” AgPlated Music wire .008” dia. minimum, about 18” long, clamped at one end, and with a small loop tied in the other.) |

## Procedure

| - Assemble to the - D020605 Upper Mass Main Section: D020607 Upper Mass T-Piece 4 Socket Head Cap Screws - ¼-20 x 1.125” AgPlated - Torque to 100 in-lb Assemble to the D1001669 Pitch Insert: 1 Socket Set Screw - ½-20 x 3.5” AgPlated Assemble to the T-Piece: 1 D1001669 Pitch Insert - Center the Insert within the T-Piece 2 Socket Set Screws - ¼-20 x 1.5” AgPlated |  | Main Section ¼-20 x 1.125” T-Piece 1.5” Set Screws Pitch Insert 3.5” Set Screw Fig 39: T-Piece Assembled to Main Section |  |
| --- | --- | --- | --- |
| Attach the D040259 Upper Mass Jig to an Optics Table with a ¼-20 Ag-Plated Bolt. Thread the T-Piece onto the ¼-20 stud at the top of the Jig. - The Jig will not be shown for the remainder of the assembly steps, but is necessary to secure the Upper Mass during the assembly process. |  | Jig 53 Fig 40: Upper Mass mounted on Jig |  |
| Use safety glasses per E0900332. Note: Blades are shown here as flattened, but are actually curved until weighted. Per data in T0900559 Blade Pairings, retrieve: 4 matched D020615 Lower Blades. 4 pair (Upper/Lower) of Lower Blade Clamps, each pair with a Clamp angle corresponding to the Matched Blades chosen above. The D0900681 Library of Clamps lists all Blade Clamp angles and part numbers. - Identify the placement of the blades as follows: - Blade with highest tip in +X, +Y corner - Blade with next to highest tip in –X, +Y corner - Blade with next to lowest tip in +X, -Y corner - Blade with lowest tip in –X, -Y corner Attach the 2 Clamp/Wire Assemblies from the Lower Blade Pulldown Device to the tip of each Blade. 2 D0901815 Upper Clamp, Int. Wire, Inside 2 D0901813 Upper Clamp, Int. Wire, Outside 4 Socket Head Cap Screw - 4-40 x .375” AgPlated 4 Flat Washer #4 Vented SSTL 4 Socket Head Cap Screw - 4-40 x .25” AgPlated Assemble each Lower Blade as shown: 2 Socket Head Cap Screws - ¼-20 x 1.25” AgPlated SSTL 2 Flat Washer ¼ Vented SSTL 1 Lower Blade Clamp, Lower 1 D020615 Lower Blade 1 Lower Blade Clamp, Upper |  | Blade Clamp Lower Side Blade D0901815 & D0901813 Clamps Assembly ¼-20 x 1.25” AgPlated Blade Clamp Upper Side Fig 41: Weighting the Blades |  |
| Attach the 2 Blade assemblies to the Upper Mass Main Section and snug the bolts finger- tight. Ensure the Blades are aligned such that the Blade tips can move through the oval cutouts without touching the cutout walls. Use the Machinist’s Square to ensure the Blades and Blade Clamps are square with the Main Section. | Blade tips must not touch walls of oval cutouts as the Blades are straightened. Square these surfaces Machinists Square Fig 42: Attaching the Blades |  |  |
| Hang the 2 D020660 Pulldown Device Weights on the wire loops. The weight will pull the blade tip through the oval slot in the Upper Mass Main Section until the Blade is essentially flat. Torque the 4 Screws to 100 in-lb. - Be sure to Torque the Screws only after the Blades have been pulled down. |  |  | Machinists Square Fig 43: Adding Weight to the Blades |
| Visually verify that the profiles of pairs of Blades match + 0.5 mm. If the Blades are not of identical flatness within 1 mm, either the Blades or Blade Clamps must be replaced to achieve this flatness. Blade Clamps are available in angles from 0 – 3.5 deg. When using Blade Clamp pairs other than 0 deg., CAREFULLY INSPECT CLAMP ORIENTATION to ensure the pairs are aligned so that the Clamp sidewalls are parallel (and therefore, the bolt holes are concentric). |  | Fig 44: Verifying Blade Matching Sight across here Correct Incorrect or or Fig 45: Upper & Lower Blade Clamp Alignment |  |
| Assemble to 2 D080221 Lower Blade Guards: 4 D030025 Socket Head Cap Screws - 8-32 x 1.0” Round Tip SSTL 4 Hex Nuts 8-32 AgPlated Attach the Blade Guards to the D020605 Main Section using: 4 Socket Head Cap Screws - 4-40 x 0.375” AgPlated 4 Flat Washers #8 Vented SSTL |  | 4-40 x 0.375” & Vented Washer Blade Guard 8-32 x 1.0” Hex Nut Fig 46: Main Section with Blade Guards |  |
| Assemble 4 sets of Screw Drives to the Upper Mass Main Section: 1 D020653 Screwdrive Block 2 Socket Head Cap Screw - 8-32 x 1.00” AgPlated 2 Flat Washer #8 SSTL 1 Socket Head Cap Screw - 8-32 x 1.50” SSTL Round Tipped 1 Hex Nut 8-32 AgPlated Tighten Finger-Tight |  | Screwdrive Block 8-32 x 1.0” & Washer 8-32 x 1.5” Round Tipped & Hex Nut Fig 47: Screw Drives Installed |  |
| Assemble to the T-Piece: D030139 Roll Offset - Center the Offset within the T-Piece 2 Socket Set Screw - ¼-20 x 0.5”AgPlated D1001669 Pitch Insert 1 Socket Set Screw - ½-20 x 3.0” AgPlated 2 Socket Set Screw - ¼-20 x 1.5” AgPlated 2 Socket Head Cap Screw - ¼-20 x 2.0” SSTL 2 Hex Nut ¼-20 AgPlated |  | ¼-20 x 2.0” ¼-20 x 1.5” ½-20 x 3.0” ¼-20 x 0.5 D030139 Fig 48: Roll Offset and Pitch Insert added |  |

| Assemble 9 assemblies to the Upper Mass Main Section, each with: 1 D1001699 Magnet Holder 2 Socket Head Cap Screws - 4-40 x 0.625” AgPlated - Torque to 6 in-lb 1 D1001695 Magnet / Flag Assembly - Leave off one of the end Magnet Flag assemblies until the Coil Holder has been assembled to the Upper Mass. | Fig 49: Magnet Assemblies added |
| --- | --- |
| Weigh the Upper Mass Assembly (including all 9 Magnet/Flag assemblies) along with these additional mass components: - 2 Upper Wire, Lower Clamp Assemblies 2 D020652 Upper Mass C-Clamps 2 D020610 Inside Upper Wire Lwr Clamp 2 D020624 Outside Upper Wire Lwr Clamp 4 Socket Head Cap Screw - 8-32 x 0.5” AgPlated 4 Flat Washer #8 Vented SSTL 4 Socket Head Cap Screw - 8-32 x 1.125” SSTL 4 Flat Washer #8 SSTL - 4 Int. Wire, Upper Clamp Assemblies 4 D070585 Upper Clamp, Inside 4 D070394 Upper Clamp, Outside 8 Socket Head Cap Screws - 4-40 x .375” AgPlated 8 Flat Washer #4 Vented 8 Socket Head Cap Screws - 4-40 x 0.25” AgPlated - Hardware for Additional Mass Disks 2 Socket Head Cap Screw - ¼-20 x 2.0” SSTL 2 ¼-20 Hex Nut AgPlated 2 Flat Washer ¼” SSTL Additional Mass Disks sufficient to bring the total mass to 12.087 kg. The additional mass should be divided equally between top disks and bottom disks. | Fig 50: Weighing the Assembly |
| Record the actual mass, and the actual additional mass amount and locations, in ICS. |  |

| Attach 2 Upper Wire Assemblies. Grasp the Upper Clamp end of each D070340 Upper Wire Assembly and feed the Assemblies upwards through the outboard openings toward either end of the Upper Mass. Attach to the Upper Mass with: 4 Socket Head Cap Screws - 8-32 x 1.125” SSTL 4 D1100785 Flat Washer #8 Nitronic 60 - If any Wire becomes bent during assembly, replace with another Wire Assembly. Attach 4 D070393 Intermediate Wire Assemblies to the 4 Lower Blades, with: 8 Socket Head Cap Screw - 4-40 x 0.25” AgPlated - Torque to 6 in-lb - Fit the tab on the bottom of the Upper Clamp into a slot on the Lower Blade. - If any Wire becomes bent during assembly, replace with another Wire Assembly. | Fig 51: Intermediate Wire Assemblies added |
| --- | --- |
| Assemble 12 EQ Stops to the D070449 Coil Holder: 12 Socket Head Cap Screws - ¼-20 x 1.0” Fully-Threaded Round-Tip SSTL Hex Nut ¼-20 AgPlated Place the Coil Holder over the Upper Mass. Feed the Upper Wire Assemblies through the Coil Holder; feed the Intermediate Wires through the holes on the ends of the Coil Holder. Assemble to the Upper Mass: 2 Socket Head Cap Screws ¼-20 x 2.0” Fully-Threaded SSTL 4 Hex Nuts, ¼-20 AgPlated - Tighten the Hex Nuts to secure the Upper Mass to the Coil Holder. | 2 ¼-20 x 2.0” Feed Intermediate Wire ends through 4 end openings 12 EQ Stops Fig 52: Coil Holder added |

# Assembling Magnets – Upper Mass

## Documents

M0900034	RODA - Use of Magnets in Suspensions

D1001695	Magnet and Flag Assembly, BOSEM

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 9 | Ea | D1100573 | BOSEM Flat Magnet Flag |
| 9 | Ea | D1100574 | BOSEM Flat Magnet Flag Disk |
| 9 | Ea | NA | - Flat Head Socket Screw 4-40 x 0.1875” SSTL |
| 9 | Ea | DCNI 01888N | Sintered NdFeB Magnet, Ni Plated, 10mm x 5mm |
| 9 | Ea | D1001534 | BOSEM Magnetic Plug |
| 9 | Ea | D1001697 | HLTS BOSEM Magnet Retainer |

## Procedure – Assembly

| - Assemble 9 D1001695 BOSEM Magnet / Flag Assemblies, each with (shown left-to-right): D1100573 BOSEM Flat Magnet Flag D1100574 BOSEM Flat Magnet Flag Disk Flat Head Socket Screw - 4-40 x 0.1875” SSTL Magnet D394394N35UHP Sintered NdFeB Ni-Plated 10 mm x 10 mm D1001534 BOSEM Magnetic Plug D1001697 HLTS BOSEM Magnet Retainer | Magnet Retainer Plug Magnet Screw Disk Flag Fig 36: HLTS BOSEM Magnet/Flag Assembly |
| --- | --- |

The Magnets attach the Flags to the Upper Mass.  Since the Flags are aluminum and the Magnet Retainers are non-magnetic stainless steel, a magnetic 416 stainless steel Magnetic Plug is press-fit into each Flag and Magnet Retainer.  The press-fitting operation must be done after all parts are cleaned and baked.  The Press itself must be cleaned and/or wrapped in UHV foil.

## Procedure – Plug Insertion

| Procedure for assembling D1001534 Plug to Magnet Holder: - Heat Air Bake Oven to 70 c; Attach Magnet Holders to Heating Fixture with: - Socket Head Cap Screw - 8-32 x 0.3125” SSTL - Screws must be Class A or B clean Place Heating Fixture in Oven for 10 min. minimum; Remove Heating Fixture from Oven and inspect Magnet Holders for out-of-round condition, using tapered end of the Disk Insertion Tool to address any out-of-round conditions. Place Disk on a Magnet Holder, Place non-tapered end of Disk Insertion Tool on Disk, and tap Insertion Tool until Disk is fully seated within Holder. Return Heating Fixture to Oven for another 5 minutes, minimum. Remove Heating Fixture from Oven, and repeat Step 5, above. Remove Magnet Holders from Heating Fixture. | Fig 37: Heating Fixture with Holders Fig 38: Insertion Tool in position Note: Tapered end of Tool is up Note: Seated Disks on left 2 Holders |
| --- | --- |

# Assembling Intermediate Mass (M2)

## Documents

D070334	HLTS Intermediate Mass Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D070336 | Main Section, Intermediate Mass |
| 1 | Ea | D080181 | Changer Assembly, Intermediate Mass |
| 1 | Ea | D030155 | Center Offset, Intermediate Mass |
| 2 | Ea | NA | Socket Head Cap Screw 8-32 x 0.75” SSTL |
| 2 | Ea | NA | Socket Head Cap Screw 8-32 x 0.75” AgPlated |
| X | Ea | D1001230 | HLTS Additional Mass Disk 10g |
| X | Ea | D1001229 | HLTS Additional Mass Disk 20g |
| X | Ea | D0901405 | HLTS Additional Mass Disk 50g |
| X | Ea | D070333 | HLTS Additional Mass Disk 100g |
| 2 | Ea | D030156 | Side Offset, Intermediate Mass |
| 4 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.5” Fully-Threaded SSTL |
| 4 | Ea | NA | Hex Nut ¼-20 AgPlated |
| 4 | Ea | NA | Flat Washer ¼ Vented SSTL |
| 4 | Ea | NA | Socket Head Set Screw ¼-20 x 0.50” AgPlated |
| 2 | Ea | NA | Socket Head Set Screw ¼-20 x 0.25” AgPlated |
|  |  |  | Components for Weighing |
| 4 | Ea | D030149 | Breakoff, Intermediate Wire |
| 4 | Ea | D070406 | Lower Clamp, Intermediate Wire, Outside |
| 4 | Ea | D070405 | Lower Clamp, Intermediate Wire, Inside |
| 8 | Ea | NA | - Socket Head Cap Screw 8-32 x 0.625” SSTL |
| 8 | Ea | NA | Flat Washer #8 Vented SSTL |
| 2 | Ea | D030148 | Lower Wire Breakoff |
| 4 | Ea | D070438 | Upper Clamp, Lower Wire, Outside |
| 4 | Ea | NA | - Socket Head Cap Screw 8-32 x 0.5625” SSTL |
| 4 | Ea | NA | Flat Washer #8 Vented SSTL |

## Procedure

| - Assemble the D080181 Changer Assembly: D030155 Center Offset 300g nominal Upper Collar D080223 300g nominal Lower Collar D080232 - 100g Collars shown Socket Head Cap Screw - 8-32 x 0.75” SSTL or AgPlated - Use SSTL with Aluminum Collars - Torque SSTL Screws to 20 in-lb - Use AgPlated with SSTL Collars - Torque AgPlated Screws to 30 in-lb | 100g Collar Fig 53: Changer Assembly |  |
| --- | --- | --- |
| Assemble 2 Side Offset assemblies, each with: 1 Side Offset D030156 2 Mass Disk D070333 100g nominal 2 Flat Washer ¼” Vented SSTL 2 Hex Nut ¼-20 AgPlated 2 Socket Head Cap Screw - ¼-20 x 1.5” SSTL | Fig 54: 1 of 2 Side Offset Assemblies |  |
| Slide the Changer Assembly and 2 Side Offsets into the D070336 Main Section and secure with: 4 Socket Head Set Screw ¼-20 x 0.50” AgPlated 2 Socket Head Set Screw ¼-20 x 0.25” AgPlated Rotate Side Offsets so Mass Disks are toward the center, and Scribe Lines are aligned. |  |  |
| Fig 55: Changer and Offsets in Main Section | Scribe Lines Fig 56: Completed Intermediate Mass Assembly |  |
| Weigh the completed Assembly along with: 4 D030149 Breakoff, Intermediate Wire 4 D070406 Lower Clamp, Intermediate Wire, Outside 4 D070405 Lower Clamp, Intermediate Wire, Inside 8 Socket Head Cap Screw - 8-32 x 0.625” SSTL 8 Flat Washer #8 Vented SSTL 12 Socket Head Cap Screw - 8-32 x .625” SSTL AgPlated 2 D030148 Lower Wire Breakoff 4 D070438 Upper Clamp, Lower Wire, Outside 4 Socket Head Cap Screw - 8-32 x 0.5625” SSTL 4 Flat Washer #8 Vented SSTL 6 Socket Head Cap Screw - 8-32 x .625” SSTL AgPlated 4 D0901927 Magnets 4 D970075 Magnet Standoffs |  | Fig 57: Items to Weigh |
| Swap Collars and Disks (red) to reach a total weight of 12.227 kg. Record the total weight, noting changes in Disks or Collars. Record the size and location of the added weights within ICS. |  |  |

# Assembling Bottom Mass and Optic (M3)

## Documents

D070337	HLTS Bottom Mass Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D070338 | Metal Bottom Mass |
| 1 | Ea | D0902658 | Optic Holder |
| 1 | Ea | D0902661 | Sapphire Prism Bonding Jig |
| 4 | Ea | D980184 | LOS Clamps, Long |
| 2 | Ea | D1100197 | Spacer |
| 2 | Ea | 033-0280 | Optosigma Mirrors |
| 8 | Ea | NA | Socket Head Cap Screw 4-40 x 0.5” Vented SSTL |
| 2 | Ea | D080124 | Prism Breakoff, Lower Wire, Lower Mass |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.75” SSTL |
| 12 | Ea | NA | Flat Washer #8 Vented SSTL |
| 8 | Ea | NA | Flat Washer #4 Vented SSTL |
| 2 | Ea | D0901286 | Secondary Metal Prism Breakoff |
| 2 |  | D070441 | Sapphire Prism Breakoff |

## Procedure

| - Place the D070338 Metal Bottom Mass, with the inscribed arrow on top of the Mass, pointing away from the glued-on Dumbbell Magnets, into the D0902658 Optic Holder. |  | Fig 58: Optic Holder and Bottom Mass |  |
| --- | --- | --- | --- |
| Assemble to the Bottom Mass: 2 D1100197 Spacers 2 033-0280 Optosigma Mirrors 8 Socket Head Cap Screws - 4-40 x 0. 5” Vented SSTL 8 Flat Washer # 4 Vented SSTL - Torque to 7 in-lb 2 D080124 Prism Breakoffs 4 Socket Head Cap Screws - 8-32 x 0.75 SSTL 4 Flat Washers #8 Vented SSTL - Torque to 25 in-lb - The Spacer wings must be aligned with the Screw locations. - Mirror Arrow must face outwards. - Be especially careful not to damage the glued magnet/dumbbell assemblies. |  | 8-32 4-40 Secondary Prism Breakoff Prism Breakoff Spacer Mirror Fig 59: Bottom Mass Fig 60: Orientation of Spacer |  |
| The fixture is used to attach the Sapphire Prism Breakoff (D070441) to the fused silica optics. The Sapphire Prism Placement Fixture is made up of the Optic Holder (D0902658) and the Sapphire Prism Bonding Jig (D0902661). The Sapphire Prism Bonding Jig itself is made up of the Sapphire Prism Bonding Fixture (D0902662) and the Sapphire Prism Holder Assembly (D0902663). | D0902662 D0902663 Fig 61: Sapphire Prism Bonding Jig |  |
| With the assembly process complete, weigh the Bottom Mass Assembly with the D0901286 Secondary Metal Prism Breakoffs; the combined weight should be 12.142 kg. Record this value in ICS. The Bottom Mass is not designed to be weight-adjusted; weight is added to or subtracted from the Intermediate Mass. So adjusting Bottom Mass weight is actually adjusting the combined weight of the Intermediate and Bottom Masses, a total of 12.227 + 12.142 kg = 24.369 kg. |  |  |

# Installing Top Blade Guards

# Installing Rotational Adjusters and Top Blade Guards

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 2 | Ea | D020660 | Blade Pulldown Device |
| 2 | Ea | D0901814 | Upper Clamp, Inside, Class B Clean |
| 2 | Ea | D070341 | Upper Clamp, Outside, Class B Clean |
| 1 | Roll | NA | Music Wire 0.024” dia. |
| 1 | Btl | NA | Methanol |
| 1 | Btl | NA | Acetone |
| 1 | Btl | NA | Isopropanol |
| 1 | Spl | NA | Steel Music Wire, 0.0106” dia. |
| 1 | Pkg | NA | Lint-Free Wipes |
| 8 | Ea | NA | Socket Head Cap Screw 8-32 x 0.50” AgPlated |
| 4 | Ea | NA | Flat Washer #8 Vented SSTL |
| 2 | Ea | D1102119 | Blade Pulldown Support |
| 1 | Roll | NA | UHV Aluminum Foil |
| 2 | Ea | D070308 | HLTS Top Blade Guard Assembly |
| 16 | Ea | NA | Socket Head Cap Screw 8-32 x 0.50” AgPlated |
| 16 | Ea | NA | Flat Washer #8 Vented SSTL |
| 2 | Ea | D070326 | HLTS Rotational Adjuster |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.75” SSTL HoloKrome |
| 4 | Ea | NA | Flat Washer #8 Vented SSTL |

## Procedure

| - Prepare 2 D020660 Blade Pulldown Devices, each with: 1 D0901814 Upper Clamp, Inside 1 D070341 Upper Clamp, Outside 2 Socket Head Cap Screw - 8-32 x 0.5” AgPlated 2 Flat Washer #8 Vented, SSTL 3’ of 0.024” Wire. - Clean the Wire per Section 12.4. | Fig 77: Blade Pulldown Device |
| --- | --- |
| Assemble the D070308 Top Blade Guard Assembly to the D070442 Structural Weldment using: 16 Socket Head Cap Screws - 8-32 x 0.5” AgPlated 16 Flat Washers #8 Vented, SSTL - Torque to 30 in-lb | Fig 78: Top Blade Guards installed |
| Remove the D070310 Bars. | Fig 79: Bars removed |
| Install the 2 D070326 Rotational Adjusters to the Mounting Pads, using for each: 4 Socket Head Cap Screws - 8-32 x 0.75” SSTL - Use HoloKrome Screws 4 Flat Washer #8 Vented, SSTL - Torque to 30 in-lb - Record the serial number and location of both Upper Blades in ICS in the RA assembly load. - The Blades are shown here as flat, but are actually curved upward. | Fig 80: Rotational Adjusters Installed |
| - Wear proper safety glasses per E1000043. Attach the 2 D1102119 Blade Pulldown Supports to the Weldment location shown; that is, in the center of the end Weldment cross member, with the Clevis extending outboard. Cover each end of the Weldment Structure and surrounding Optical Table areas with UHV Aluminum Foil, to protect them from the dirty Pulldown Device. - 2 workers required: 1st person holds the Pulldown Weight. 2nd person passes Wire Clamp of the Pulldown Device through the Weldment side opening, up toward the Upper Blade Tip, then attaches the Clamp to the Blade tip with: 2 Socket Head Cap Screws - 8-32 x 0.50” AgPlated 1st person gently drapes the wire over the Clevis, and slowly releases the Weight. Repeat Steps 7-9 for the second Pulldown Device. | Fig 81: Blade Pulldown Support Attach Blade Pulldown Supports here Fig 82: Location of Blade Pulldown Support |
| Re-Assemble the 2 D070310 Top Blade Guard Bars to the Risers, using for each: 4 Socket Head Cap Screws - ¼-20 x 1.0” SSTL - Torque to 75 in-lb - Ensure the Bars are oriented with the EQ Stop Screws directly over the Blades. - The EQ Stop Screws should be adjusted so the Blades are flat. Once adjusted, the Screws should be secured with the Hex Nuts. Carefully remove the 2 Blade Pulldown Devices. Remove the 2 Blade Pulldown Supports. | Fig 83: Top View: Screws centered over Blades Fig 84: End View: Screws adjusted and secured |

# Installing EQ Stop Mounts and Brackets

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 10 | Ea | D080726 | HLTS EQ Stop, Long Mount |
| 8 | Ea | D080727 | HLTS EQ Stop, Long Bracket |
| 2 | Ea | D080728 | EQ Stop Assembly, Upper, Bottom Mass |
| 4 | Ea | D1002821 | EQ Stop Assembly, Bottom Mass, Lower |
| 4 | Ea | D1102071 | EQ Stop Assembly, Bridge, Upper |
| 52 | Ea | NA | Socket Head Cap Screws 8-32 x 0.5” AgPlated |
| 52 | Ea | NA | Flat Washers #8 SSTL |
| 4 | Ea | NA | Socket Head Cap Screws 8-32 x 0.5” Vented AgPlated |
| 4 | Ea | NA | Flat Washers #8 Vented SSTL |

## Procedure

| - Attach 2 D1002821 Short Brkt EQ Stops to the inside front of the Weldment, using: 4 Socket Head Cap Screw - 8-32 x .5” Vented, AgPlated 4 Flat Washer #8 SSTL - Torque to 30 in-lb Attach the following, using: 52 Socket Head Cap Screws 8-32 x 0.5” AgPlated 52 Flat Washers #8 SSTL - Torque to 30 in-lb 6 D080726 Long Mount EQ Stops to the inside front of the Weldment 2 D080726 Long Mount EQ Stops to the inside rear of the Weldment 4 D080727 Long Brkt EQ Stops to the inside front of the Weldment 4 D080727 Long Brkt EQ Stops to the inside rear of the Weldment 2 D080728 Bottom Mass Upper EQ Stops to the inside front of the Weldment (replacing the lowest 2X D080726 shown in Figure 75); NOTE: these assemblies are canted outboard by 5 degrees, as shown by the threaded holes in the Structure. 4 D1102071 Bridge EQ Stops to the Weldment. At the Bridge EQ Stops, torque the 8 ¼-20 Screws (assembled earlier) that attach the Crossbars to the Bridge pieces, to 75 in-lb. | D080726 D080727 D080727 D1102071 D1002821 D080727 D080727 Fig 75: Facing Front from Inside of Weldment D080726 D1102071 D080727 Fig 76: Facing Rear from Inside of Weldment |
| --- | --- |

# Assembling Intermediate Wires

## Documents

D070393	HLTS Intermediate Wire Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D0900630 | HLTS Intermediate Wire Jig |
| 4 | Ea | D980184 | LOS Clamps, Long |
| 1 | Ea | D070585 | - Inside Upper Clamp |
| 1 | Ea | D070394 | Outside Intermediate Wire Upper Clamp |
| 1 | Ea | D030149 | Intermediate Wire Breakoff |
| 1 | Ea | D070405 | Inside Intermediate Wire Lower Clamp |
| 1 | Ea | D070406 | Outside Intermediate Wire Lower Clamp |
| 2 | Ea | NA | Socket Head Cap Screw 4-40 x 0.5” SSTL |
| 3 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” SSTL |
| 2 | Ea | NA | Socket Head Cap Screw 4-40 x 0.375 AgPlated |
| 2 | Ea | NA | Flat Washer #4 Vented SSTL |
| 2 | Ea | NA | Socket Head Cap Screw 8-32 x 0.625” SSTL |
| 2 | Ea | NA | Flat Washer #8 Vented SSTL |
| 2 | Ea | 1185-2EN164 | Helicoil 8-32 x .164” Nitronic 60 |
| 2 | Ea | 1185-04EN224 | Helicoil 4-40 x .224” Nitronic 60 |
| 1 | Btl | NA | Methanol |
| 1 | Btl | NA | Acetone |
| 1 | Btl | NA | Isopropanol |
| 1 | Spl | NA | Steel Music Wire, 0.0134” dia. |
| 1 | Pkg | NA | Lint-Free Wipes |
| 1 | Ea | NA | Class B Wire Cutters |
| 1 | Ea | NA | Hang Weight, 1 Kg, or D020660 Blade Pulldown Device |
| 4 | Ea | NA | Large Test Weights, 10, 5, 2, 1 Kg |
| 1 | Ea | NA | Small Test Weight Set, 1 to 500 g |

## Procedure

| - 4 Assemblies are required per HLTS. - Use safety glasses per E0900332. - Attach the D0900630 Jig to an Optical Table using the 4 D980184 LOS Clamps. The end of the Jig with the Clevis and Pin must hang over the edge of the Table enough to provide clearance for the hanging weight. | Edge of Table Fig 13: Jig |  |
| --- | --- | --- |
| Assemble all the Clamps you’ll need; do not tighten the clamps. - Intermediate Wire Upper Clamp includes: 1 D070585 Intermediate Wire Upper Clamp, Inside 2 1185-04EN224 Helicoil 4-40 x .224” 1 D070394 Intermediate Wire Upper Clamp, Outside Socket Head Cap Screw - 2 4-40 x 0.375” AgPlated 2 Flat Washer #4 SSTL - Intermediate Wire Breakoff Assembly includes: 1 D030149 Intermediate Wire Breakoff 1 D070405 Intermediate Wire Lower Clamp, Inside 2 1185-2EN164 Helicoil 8-32 x .164” 1 D070406 Intermediate Wire Lower Clamp, Outside 2 Socket Head Cap Screw - 8-32 x 0.625” SSTL 2 Flat Washer #8 SSTL |  | D070394 D070585 (2) Fig 14: Upper Clamp D070406 (2) D070405 (2) D030149 (2) Fig 15: Breakoff |
| Attach to the Jig: 1 Upper Clamp Assembly 2 Socket Head Cap Screw - 4-40 x 0.5” SSTL 1 Breakoff Assembly 3 Socket Head Cap Screw - 8-32 x 0.5” SSTL | Upper Clamp Assembly Breakoff Assembly Fig 16: Clamps Attached to Jig |  |
| Cut 24” of 0.0134” diameter Steel Music Wire. Clean the Wire per Section 12.4. Feed the Wire through each set of 4 Clamps. Allow the extra Wire to hang over the edge of the Table. Tighten the Screws on the Wire Start Clamp farthest from the Table edge. - Ensure Wire is secured within the Clamp groove. | Edge of Table Fig 17: Wire Path |  |
| Tie a small loop in the end of the Wire hanging off the Table and hang the 1 kg Hang Weight or Pulldown Device from the loop. Add additional weights to the Hanger for a total of 6.092 kg. Then change gloves, since the weights are not clean. | Fig 18: Hang Weight in use |  |
| Wait 5 minutes for the wire to come to equilibrium. Tighten the screws on each of the 3 remaining wire clamps in the order shown. 4-40 Screws: Torque to 6 in-lb 8-32 Screws: Torque to 30 in-lb - During tightening: Ensure Wire is secured within the grooves of each Clamp. Ensure Clamp halves are aligned with each other by using a Precision Square. Remove the Hang Weight and cut the Wire in the locations shown. Change gloves since the weights are not clean. Loosen the Screws on the Wire Clamps on the jig and discard the leftover Wire. | Precision Square Cut here Cut here 1 3 2 Fig 19: Cutting the Wires |  |
| Remove the screws holding the Intermediate Wire Assembly to the Intermediate Wire Jig. Bend back the ends of the wire to eliminate sharp points. Measure the Wire length between Clamps to confirm that it is 203.60mm. | Fig 20: Completed Intermediate Wire Assembly |  |

# Assembling Upper Wires

## Documents

D070340	HLTS Upper Wire Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D0900594 | HLTS Upper Wire Assembly Jig |
| 4 | Ea | D980184 | LOS Clamps, Long |
| 1 | Ea | D020611 | Upper Clamp, Upper Wire, Inside |
| 1 | Ea | D020652 | C-Clamp, Upper Mass |
| 1 | Ea | D070341 | Upper Clamp, Upper Wire, Outside |
| 1 | Ea | D020610 | Upper Wire, Lower Clamp, Inside |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.625” SSTL |
| 4 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” AgPlated |
| 4 | Ea | NA | Flat Washer #8 Vented SSTL |
| 4 | Ea | 1185-2EN246 | Helicoil 8-32 x .246” |
| 1 | Ea | NA | Hang Weight, 1 Kg, or D020660 Blade Pulldown Device |
| 4 | Ea | NA | Large Test Weights, 10, 5, 2, 1 Kg |
| 1 | Ea | NA | Small Test Weight Set, 1 to 500 g |
| 1 | Btl | NA | Methanol |
| 1 | Btl | NA | Acetone |
| 1 | Btl | NA | Isopropanol |
|  |  |  | Steel Music Wire, 0.024” dia. |
|  |  |  | Lint-Free Wipes |
|  |  |  | Class B Wire Cutters |

LOS Clamp

Procedure

| - 2 Assemblies are required per HLTS. - Use safety glasses per E0900332. - Ensure the D0900594 Assembly Jig has been cleaned Class B per E960022. - Ensure Jig is fully assembled per the drawing. - Attach the D0900594 Jig to an Optical Table using the 4 D980184 LOS Clamps. The end of the Jig with the Clevis and Pin must hang over the edge of the Table enough to provide clearance for the hanging weight. | Clevis Class B Class A Table Edge Fig 5: Jig |  |
| --- | --- | --- |
| - Assemble all the Clamps you’ll need; do not tighten the clamps. - Upper Wire Upper Clamp includes: 1 D020611 Upper Wire Upper Clamp, Inside 2 1185-2EN246 Helicoils 8-32 x .246” 1 D070341 Upper Wire Upper Clamp, Outside Socket Head Cap Screw - 2 8-32 x 0.5” AgPlated 2 Flat Washer #8 Vented SSTL - Upper Wire C-Clamp Assembly includes: 1 D020652 Upper Mass C-Clamp 1 D020610 Upper Wire Lower Clamp, Inside 2 1185-2EN246 Helicoil 8-32 x .246” 1 D020624 Upper Wire Lower Clamp, Outside 2 Socket Head Cap Screws - 8-32 x 0.5625” SSTL 2 Flat Washer #8 Vented SSTL |  | D070341 (2) D020611 Fig 6: Upper Clamp D020624 (2) D020652 D020610 (2) Fig 7: C-Clamp |
| - Attach to the Jig: 1 Upper Clamp Assembly 2 Socket Head Cap Screws - 8-32 x 0.625” SSTL 1 C-Clamp Assembly with 2 Socket Head Cap Screws - 8-32 x 0.625” SSTL | Upper Clamp C-Clamp Fig 8: Clamps on Jig |  |
| - Cut a 36” of 0.024” diameter Wire from the spool. - Clean the Wire per Section 12.4. - Feed the Wire through the Clevis and Clamps, as shown: ### Over Clevis Pin ### 1st Start Post; ### L-Clamp; ### C-Clamp; ### 2nd Start Post - Tighten the 2nd Start Post Clamp, after feeding through about ½” of wire. - Ensure Wire is secured within the Clamp groove. | Tighten first Fig 9: Wire Path |  |
| - Tie a small loop in the end of the wire hanging off the Table and hang the 1 kg Hang Weight. Add large and small weights for a total of 18.252 kg. Then change gloves, since the weights are not clean. | Fig 10: Hang Weight in use |  |
| - Wait 5 minutes for the wire to come to equilibrium. Tighten the screws on each of the 3 remaining wire clamps in the order shown. Torque to 30 in-lb. - During tightening: Ensure Wire is secured within the grooves of each Clamp. Ensure Clamp halves are aligned with each other by using a Precision Square. - Remove the Hang Weight and cut the Wire in the locations shown. Change gloves since the weights are not clean. - Loosen the Screws on the Wire Clamps on the jig and discard the leftover Wire. | Precision Square Cut here Cut here 3 2 1 Fig 11: Cutting the Wires |  |
| - Remove the Screws holding the Upper Wire Assembly to the Upper Wire Jig. Bend back the ends of the wire to eliminate sharp points. - Measure the Wire length between Clamps to confirm that it is 202.50mm. | Fig 12: Upper Wire Assembly complete |  |

# Installing Upper Mass and Coil Holder

## Documents

D070442	HLTS Overall Assembly

D1101493	HSTS / HLTS / OMCS OSEM Orientations

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D040259 | Upper Mass Jig |
| 1 | Ea | NA | HLTS Upper Mass and Coil Holder Assembly from section 17, above |
| 12 | Ea | NA | Socket Head Cap Screws ¼-20 x 1.0” Round-Tip SSTL |
| 14 | Ea | NA | Hex Nut ¼-20 AgPlated |
| 5 | Ea | Several | HLTS Additional Mass Disks (10g, 20g 50g 100g) |
| 2 | Ea | NA | Socket Head Cap Screw ¼-20 x 2.0” SSTL |
| 2 | Ea | NA | Flat Washer ¼” Vented SSTL |
| NA | Ea | D1002133 | HLTS Coil Holder Bracket Assemblies |
| 4 | Ea | D1002134 | Bracket Lateral Section |
| 4 | Ea | D1002135 | Bracket Longitudinal Section |
| 8 | Ea | NA | Socket Head Cap Screws 4-40 0.75” SSTL |
| 8 | Ea | NA | Flat Washer #4 Vented SSTL |
| 12 | Ea | NA | Socket Head Cap Screw 8-32 x 0.75” AgPlated |
| 12 | Ea | NA | Socket Head Cap Screw 8-32 x 0.625” SSTL |
| 24 | Ea | NA | Flat Washer #8 Vented SSTL |

## Procedure

| The corner Brackets are shown here only for orientation and are not yet attached. Note the (2) Bracket Screws used for vertical adjustments. - Attach the D040259 Jig to an Optics Table. Remove but keep all Add-On hardware from the bottom of the Upper Mass T-Piece. Attach the previously-assembled Upper Mass and Coil Holder (with 6 Wire assemblies) to the Jig. (Jig not shown here, but in section on Assembling Upper Mass). Using the pair of ¼-20 Screws at the top center, raise the Upper Mass completely into the Coil Holder, then tighten the Hex Nuts. Remove the 4 side EQ Stop Screws and Hex Nuts. Lock down the Upper Mass with the remaining 8 EQ Stop Screws and Hex Nuts. | Lock down Upper Mass Remove 4 Side Screws Bracket with 2 vertical adjustment Screws Raise the Upper Mass Screws for vertical adjustment Fig 85: D1002133 Coil Holder Brackets (4) |  |
| --- | --- | --- |
| Assemble 4 D1002133 HLTS Coil Holder Bracket Assemblies, each with: 1 D1002134 Lateral Section 1 D1002135 Longitudinal Section 2 Socket Head Cap Screws - 4-40 0.75” SSTL 2 Flat Washer #4 Vented SSTL Torque to 5 in-lb |  | Fig 86: Bracket Assembly |
| Attach the 4 Brackets to the Weldment, using for each: 3 Socket Head Cap Screws - 8-32 x 0.75” AgPlated 3 Flat Washer #8 Vented SSTL - Torque Screws after Coil Holder is installed and aligned. |  | Fig 87: Coil Holder Brackets Installed |
| Insert the Upper Mass / Coil Holder Assembly through either narrow opening in the top section of the Weldment. The pair of BOSEMS on top of the Assembly must be at the –Y side of the Suspension per D1101493 OSEM Orientations. |  |  |
| Assemble the Coil Holder to the 4 Brackets using at each Bracket: 3 Socket Head Cap Screws - 8-32 x 0.625” SSTL 3 Flat Washer #8 Vented SSTL - Torque to 20 in-lb Back out each of the 4 lower Clamps of the Intermediate Wire Assemblies, back through the 4 end openings in the Coil Holder. Re-assemble 4 EQ Stop Screws to the long sides of the Coil Holder, each using: 1 Socket Head Cap Screw - ¼-20 x 1.0” Round Tip SSTL 1 Hex Nut ¼-20 AgPlated |  | Side EQ Stop Screw Fig 88: Upper Mass & Coil Holder Installed |
| Install the Additional Mass Disks on the top and bottom of the Upper Mass, as equally divided as possible; tighten the Hex Nuts to secure the Disks. Use: 2 Socket Head Cap Screw - ¼-20 x 2.0” SSTL 2 ¼-20 Hex Nut AgPlated 2 Flat Washer ¼” SSTL |  | Fig 89: Additional Mass Disks Added |

# Installing the Intermediate Mass

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | ea | D070334 | HLTS Intermediate Mass |
| 1 | Ea | D1102344 | HLTS Intermediate Mass Lifting Plate Assembly |

## Procedure

| - Raise the 4 bottom EQ Stops as far as possible. - Back off the 12 top, front and side EQ Stops | Fig 90: Rear View; EQ Stops prepared |
| --- | --- |
| - Using the D1102344 Lifting Plate Assembly, insert the D070334 Intermediate Mass assembly through the rear of the Weldment, ensuring the 4 Magnet assemblies are facing the rear. Center the assembly front-to-back and left-to-right on the bottom EQ Stops. | Fig 91: Lifting Plate Fig 92: Rear View; Magnets facing rear |
| - Lock down the Intermediate Mass assembly using the 8 top and side EQ Stops. | Fig 93: Rear View; Mass locked down |

# Assembling Lower Loop Wire

## Documents

D070436	HLTS Lower Loop Wire Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D0901419 | HLTS Lower Loop Wire Jig |
| 4 | Ea | D980184 | LOS Clamps, Long |
| 2 | Ea | D030148 | - Lower Loop Wire Intermediate Mass Breakoffs |
| 6 | Ea | NA | Socket Head Cap Screw 8-32 x 0.5” SSTL |
| 4 | Ea | D070438 | - Lower Loop Wire Clamps |
| 4 | Ea | NA | - Socket Head Cap Screw 8-32 x 0.5625” SSTL |
| 4 | Ea | NA | - Flat Washer #8 Vented SSTL |
| 4 | Ea | 1185-2EN164 | Helicoil 8-32 x .164” |
| 1 | Btl | NA | Methanol |
| 1 | Btl | NA | Acetone |
| 1 | Btl | NA | Isopropanol |
| 1 | Spl | NA | Steel Music Wire, 0.0106” dia. |
| 1 | Pkg | NA | Lint-Free Wipes |
| 1 | Ea | NA | Wire Cutters, cleaned Class B |
| 1 | Ea | NA | Hang Weight, 1 Kg, or D020660 Blade Pulldown Device |
| 4 | Ea | NA | Large Test Weights, 10, 5, 2, 1 Kg |
| 1 | Ea | NA | Small Test Weight Set, 1 to 500 g |

## Procedure

| - One Assembly is required per HLTS. - Use safety glasses per E0900332. - Attach the D0901419 Lower Loop Wire Jig to an Optical Table using 4 D980184 Long LOS Clamps arranged so both ends of the Jig hang over the edges of the Table. |  | Table Edge Fig 21: Jig |  |
| --- | --- | --- | --- |
| Assemble loosely the Lower Loop Breakoffs, each using: 1 D030148 Breakoff - Ensure the vent groove faces downward 2 D070438 Lower Loop Wire Clamp 2 Socket Head Cap Screw - 8-32 x .5625” SSTL 2 Flat Washer #8 Vented SSTL |  |  | Vent Groove (2) D030148 (2) D070438 (2) Fig 22: Breakoff |
| Attach the D030148 Lower Loop Wire Intermediate Mass Breakoffs, vent grooves facing down. 6 Socket Head Cap Screw - 8-32 x 0.5” SSTL - Tighten the screws. |  | Breakoff Breakoff Fig 23: Attaching the Breakoffs |  |
| Cut one length of 0.0106” diameter Steel Music Wire. Clean the Wire per Section 12.4. Feed the Wire through a set of Wire Clamps in the order shown. Allow the long end of the wire to hang over the edge of the Table passing over the Wire Jig Pin Support. Tighten Wire Start Clamp farthest from the corresponding Wire Jig Pin Support. - Ensure the wire is secured within the groove of the clamp. Repeat Steps 4-7 for the second wire, but feed the Wire through the Clamp in the opposite direction. |  | 3 4 2 1 Fig 24: Loading the Wires |  |
| Tie a small loop in each Wire, at the end of the wire hanging off the Table. Hang a 1 kg Hang Weight from each loop. Add more weights to each hanger for a total of 6.071 kg per Wire. Then Change gloves since the weights are not clean. | Fig 25: Hang Weight in position |  |  |
| Wait 5 minutes for the wires to come to equilibrium. Tighten the Screws on the Lower Loop Wire Clamps, then tighten the screws on the remaining Wire Start Clamps. Torque to 30 in-lb. - During tightening: Ensure Wire is secured within the grooves of each Clamp. Ensure Clamp halves are aligned with each other by using a Precision Square. Cut the wires between the Lower Loop Wire Clamps and the Wire Start Clamps. Remove the weights on both ends. Change gloves since the weights are not clean. |  | Precision Square Fig 26: Tightening and Cutting the Wires |  |
| Loosen the screws on the Wire Start Clamps on the Jig and discard the leftover pieces of wire. Remove the Screws holding the Lower Loop Wire Assembly to the Jig. Bend back the ends of the wires to eliminate sharp points. |  | Fig 27: Lower Loop Wire complete Assembly |  |

# Installing Wire Assemblies

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 4 | Ea | NA | Socket Head Cap Screws 8-32 x .50” AgPlated |
| 18 | Ea | NA | Socket Head Cap Screws 8-32 x .625” AgPlated |

## Procedure

| Upper Wire Assemblies fully installed. |  |  |  | Fig 94: Upper Wire Assemblies installed |
| --- | --- | --- | --- | --- |
| - Raise the Upper Mass/Tablecloth Assembly on its brackets as high as possible and secure it in place. |  | Loosen Screws to allow vertical adjustment Fig 95: Raising Coil Holder on its Brackets |  |  |
| - Be careful not to kink or twist the wires. - Attach the 2 Upper Wire Upper Clamps to the 2 Upper Blades using: 4 Socket Head Cap Screws - 8-32 x 0.50” AgPlated - Torque to 30 in-lb |  |  | Fig 96: Upper Clamps attached to Upper Blades |  |
| - Center the C-Clamps using the 4 Screwdrive System Screws, then tighten the Hex Nuts to lock the Screws. - Torque the eight 8-32 x 1.125” C-Clamp Screws to 30 in-lb. |  |  |  | Fig 97: Centering C-Clamps with Screw Drives |
| - Adjust the height of the Upper Mass/Tablecloth Assembly on its brackets to remove slack in the Upper Wires. | Loosen Screws to allow vertical adjustment Fig 98: Lowering Coil Holder to eliminate slack in Wires |  |  |  |
| - Attach the 4 Intermediate Wire Lower Clamps to each end of the Intermediate Mass Assembly using: 12 Socket Head Cap Screws - 8-32 x .625” AgPlated - Torque to 30 in-lb |  |  |  | Fig 99: End of Intermediate Mass |
| - Remove the 2 Screws and Hex Nuts at the top center of the Upper Mass Assembly that are holding the Upper Mass Assembly against the Coil Holder. The Upper Mass Assembly should still be locked relative to the Coil Holder using the top and side earthquake stops. |  |  |  | Fig 100: Screws to remove |
| - Adjust any or all of the 16 top, bottom and side intermediate EQ Stops to remove slack in the Intermediate Wires. |  |  |  | Fig 101: Front View, EQ Stops |
| - Attach the Lower Loop Wire to the Intermediate Mass Assembly using: 6 Socket Head Cap Screws - 8-32 x .625” AgPlated - Torque to 30 in-lb - Be careful not to kink, twist or tangle the wires. Use silver-plated SHCS. |  |  |  | Fig 102: End View of Intermediate Mass |

# Installing Bottom Mass

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 2 | Ea | D0901286 | HLTS Secondary Metal Prism Breakoff |

## Procedure

| - Adjust the Lower Bridge EQ Stops (4 Screws) as far in as possible. Back off the Upper Bridge EQ Stops (4 Screws) and Front EQ Stops (4 Screws). | Fig 116: EQ Stops Adjusted |
| --- | --- |
| Install the Bottom Mass Assembly through the back of the HLTS Structure. Place the Bottom Mass Assembly above the Lower Wire Assembly, resting on the Lower Bridge EQ Stops, making sure that the 2 Wires are straight and fit in the 2 grooves of each Prism Breakoff. - Be sure not to kink, twist or tangle the wires. Install the 2 D0901286 Secondary Metal Prism Breakoffs. Insert the breakoffs between the Bottom Mass Assembly and the Lower Loop Wire at the points where the breakoffs just touch the wires. | Fig 117: Lower Mass Placed on Bridge EQ Stops |

# Suspension and Alignment of Masses

Note:  “Level” as referenced below, is defined as being within the visual indicators on the Bubble Levels used (Carpenter’s Level for Optical Table, Single Bubble for Suspension Masses).

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | NA | Bubble Level |
| 1 | Ea | NA | Optical Level |
| 1 | Ea | NA | Lower Loop Wire Comb |

## Procedure

| - Unlock all Stops to allow the Suspension to hang freely: Bottom Mass 4 Face EQ Stops 8 Bridge EQ Stops Intermediate Mass 16 EQ Stops Upper Mass 4 Lower Blade Stops 12 Coil Holder Stops Upper Blades 4 Upper Blade Stops |  | Bridge EQ Stops Face EQ Stops EQ Stops Coil Holder Stops Lower Blade Stops Upper Blade Stops Fig 118: Unlock all Stops |
| --- | --- | --- |
| Lock the Intermediate Mass Lock the Intermediate Mass with the 16 EQ Stops such that the Mass is level. Use a Bubble Level, and then an Optical Level. The Mass height is unimportant at this point. Adjust Bottom Mass Roll Using the Lower Loop Wire Comb, ensure the Lower Wires are parallel and 10mm apart. Measure and compare the heights of the bottom of the side bores, on the with side of the front of the Bottom Mass. Rotate the Mass until the heights are equal. Adjust Bottom Mass Pitch Measure and compare the heights of the bottom of the bores on the front and back of the Mass. If unequal, first repeat Step 3, above. Then verify that the Lower Wire Clamps are attached squarely to the Intermediate Mass. If the heights remain unequal, the lengths of the Lower Wires are unequal, and need to be replaced. Unlock the Intermediate Mass Once the Bottom Mass is level, unlock the Intermediate Mass. |  | 2X 2X 2X 2X 2X 2X 4X Fig 119: Intermediate Mass Leveling Fig 120: Adjust Roll and Pitch |
| Adjust Yaw Using a Ruler while adjusting the Corner Brackets, center the Coil Holder within the Weldment, in the X direction (front-to-back). Adjust the Rotational Adjusters as needed to center the Magnet Holders at either end of the Upper Mass within the Coil Holder Openings. To adjust the Rotational Adjusters, loosen the 3 screws attaching the RA to its base, then adjust the Push and Pull Screws, then tighten the 3 screws. |  | Loosen 3 Screws Center Magnet Holders within opening Push & Pull Screws Fig 121: Adjusting Yaw |
| Adjust Upper Mass Pitch Measure the heights of either the bottom of the T-Piece or the bottoms of the Screw Drive Blocks, front and back (+X & - X). To adjust Pitch, manipulate the Screw Drives to shift the location of the Upper Wire Lower Clamps. Diagram shows Upper Mass without Coil Holder nor Blades, for visibility. |  | Screw Drive Measuring Points Fig 122: Upper Mass Pitch, End View |
| Measure heights in mm from the Optical Table surface of these points. *Not measurable with Optical Level | Upper Blade Wire Breakoffs (2); NOT Blade tip; 806.120. Coil Holder; Upper surface; 658.587. Upper Mass – Screwdrive Blocks (4); Bottom surface; 613.181. Upper Mass – T-Piece; Bottom surface; 552.228. Lower Blade Wire Breakoff (4); NOT Blade tip; 608.838. Intermediate Mass; Upper surface; 461.620. Intermediate Mass; Lower surface; 371.620. Bottom Mass Metal – Side Bores (2); Bottom; 130.243. Bottom Mass Metal or Optic; Centerline*; 158.500. Bottom Mass Metal; Bottom; 25.984. * Not measurable using an Optical Level |
| - Perform the following on the 28 ¼-20 EQ Stop Screws shown at right: Adjust each Screw so that it contacts the Mass, but applies no pressure. Rotate each Screw ¾ turn counterclockwise to leave a 1 mm gap between the Screw Tip and the Mass. Tighten the Hex Nuts at each of the Screws, to lock the Screws in position. |  | 4X 4X 2X 2X 2X 4X 2X 2X 2X 2X 2X Fig 123: 1 mm gap at EQ Stops |
| Adjust the 4 Lower Blade Guard Screws so they contact the Lower Blades but apply no pressure. Rotate the Screws 1 ¼ turns counterclockwise, to leave a 1 mm gap at the Lower Blades. Repeat these 2 steps with the 2 Upper Blade Guard Screws, but with only a ¾ turn counterclockwise. |  | Fig 124: Blade Guard Screws |

# Assembling Magnets To Intermediate Mass

## Documents

E990196	Magnet / Standoff Assembly Preparation

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D0901461 | HLTS Magnet Placement Fixture, Intermediate Mass |
| 4 | Ea | D980184 | LOS Clamps |
| 4 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.5” AgPlated |
| 4 | Ea | D020661 | Magnet/Standoff Assemblies, 2 N and 2 S configurations |
| 1 | Ea | NA | Machinist Square, approx. 6” in length |
| 1 | Ea | TBD | Epoxy Gun Applicator, MasterBond |
| 1 | Ea | EP30-2 | Epoxy, Double Barrel Cartridge with Mix Tube, MasterBond |
| 1 | Ea | NA | Tweezers |
| 1 | Btl | NA | Isopropanol |
| X | Ea | NA | Lint Free Wipes |
| X | Ea | TBD | Sewing Needle |
| X | Ea | TBD | Razor Blade |
| X | Roll | NA | UHV Aluminum Foil |
| 1 | Ea | NA | Heat Lamp, 120w Bulb |

## Procedure

| - Prepare 2 “N” and 2 “S” D020661 Magnet/Standoff assemblies per E990196 Preparation procedure. |  | Fig 125: D020661 Magnet/Standoff Assembly |
| --- | --- | --- |
| Place the Mass or Optic Mount the Base Plate of the D0901461 Magnet Placement Fixture on an Optics Table using the 4 D980184 LOS Clamps. Place the Intermediate Mass Assembly on the Base Plate with the back side facing up. The Intermediate Mass is symmetric; front/back is identified only after the Mass is balanced within the Weldment. Place the Positioning Standoff over the Mass. Using a Machinist’s Square, carefully align the engraved markings on the D0902452 Side Plates with those on the main Standoff Plate D0902445 (see red dashed lines at right). The Sides and Standoff Plate together form the D0902444 Positioning Standoff. | Positioning Standoff ### Base Plate Fig 126: Magnet Placement Fixture Machinist’s Square Fig 127: Aligning Side Plates with Main Plate |  |  |
| - Load Plungers Wipe the counterbore end of each plunger with Isopropanol and a Wipe. Using the Tweezers, load 4 Magnet/Standoff assemblies into the 4 Plungers, 2 North Magnets and 2 South Magnets. The Magnet end of each assembly rests within the Plunger counterbore. - The Magnet/Standoffs are held to the Plungers magnetically. |  | Counterbore Assembly held to Plunger Magnetically Fig 128: Plungers Empty and Loaded |
| Lay the loaded Plungers on the Optics Table around the Fixture per the Magnetic Polarity layout shown. |  | Bottom of Mass Top of Mass N S N S Fig 129: Magnet Polarity Layout |
| - Bond Magnets to Mass Load the EP30-2 Cartridge with Mix Tube attached, into the Gun Applicator. Pull the trigger on the Gun Applicator 1 full stroke, to purge the Mix Tube of under-mixed adhesive. Dispense a “quarter-sized” pool of Adhesive onto a small piece of clean UHV aluminum foil. Pick up a Plunger loaded with a Magnet/Standoff assembly and hold it vertically, with the Magnet/Standoff end facing up. Clean the Standoff with Isopropanol and a Wipe. Dip the end of a Sewing Needle in the pool of Epoxy and withdraw it, leaving a tiny drop on the Needle tip. Apply approximately ½ mm of Epoxy to the center of the Standoff end. Load the Plunger, Magnet/Standoff down, into the appropriate Bushing in the MPF. Slide the Plunger down within the Bushing until the Standoff contacts the Mass. Press down on the Plunger lightly with one finger for about 2 seconds, then release. Repeat steps 11-13 to load all 4 Plungers into the Placement Fixture. Allow the Epoxy to cure within the Fixture at room temperature for 24 hours. |  | Plunger in Bushing Fig 130: Loading Magnets into Fixture |
| Carefully remove the 4 Plungers from the Magnet Placement Fixture, and remove the Positioning Standoff from the Mass. Center the Heat Lamp over the Fixture and adjust the height such that the Fixture surface is receiving 60°C, then allow the adhesive to cure for 4hr. |  | Fig 131: Heat Lamp Cure |

# Assembling Magnets To Bottom Mass

## Documents

E990196	Magnet / Standoff Assembly Preparation

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 1 | Ea | D0901460 | HLTS Magnet Placement Fixture, Bottom Mass |
| 4 | Ea | D980184 | LOS Clamps |
| 4 | Ea | NA | Socket Head Cap Screw ¼-20 x 1.5” AgPlated |
| 1 | Ea | D070338 | HLTS Bottom Mass |
| 1 | Ea | Various | HLTS Optic |
| 4 | Ea | D0902432 | Magnet/Standoff Assemblies, 2 N and 2 S configurations |
| 1 | Ea | NA | Machinist Square, approx. 6” in length |
| 1 | Ea | NA | Depth Gage; either Vernier Calipers or Spring-Type Needle Gage |
| 1 | Ea | EP30-2 | Epoxy, Double Barrel Cartridge with Mix Tube, MasterBond |
| 1 | Ea | TBD | Gun Applicator, MasterBond |
| 1 | Ea | NA | Tweezers |
| 1 | Btl | NA | Isopropanol |
| X | Ea | NA | Lint Free Wipes |
| X | Ea | TBD | Sewing Needle |
| X | Ea | TBD | Razor Blade |
| X | Roll | NA | UHV Aluminum Foil |
| 1 | Ea | NA | Heat Lamp, 120w Bulb |

## Procedure

- Glass Optics and Metal Masses will not be Air Baked.

- Glue Magnets before gluing Prisms (primary and secondary).

- Ensure the Main Section of the Mass has been cleaned and baked before attaching the Magnet/Dumbbell assemblies.

- Thoroughly Class B clean all parts of the Magnet Placement Fixture.

| - Prepare 2 “N” and 2 “S” D0902432 Magnet/Standoff assemblies per E990196 Preparation procedure. | Fig 132: D0902432 Magnet/Standoff Assembly |
| --- | --- |
| - Place the Mass or Optic Mount the Base Plate of the D0901460 HLTS Magnet Placement Fixture on an Optics Table using the 4 D980184 LOS Clamps. Place the Test Mass or Optic on the Base Plate. - The Mass/Optic Arrow must be facing downwards. | Arrow Fig 133: Base Plate Mounted on Optics Table |
| Place the Magnet Positioning Fixture (MPF) on top of the Mass / Optic. - For the Optic, ensure the PFA440HP-tipped Stops are installed on the Stop Screws | MPF PFA Tip Fig 134: MPF in place |
| - Align the MPF and Mass / Optic scribe lines: - For clarity, the Base Plate is not shown. - The MPF Scribe Lines are not visible in the drawings; they fall within 4 flat areas on the MPF perimeter, about 15° clockwise from each Screw. Center the MPF on the Mass / Optic by obtaining equidistant readings between opposing pairs of MPF Scribe Lines and the Mass Optic perimeter, using a Depth Gage. The MPF Screw tips must barely contact and not “clamp” the Mass/Optic. For a Metal Mass, rotate the Mass while using a Machinist Square to align the Mass and Fixture scribe lines. Align at 2 Lines 90° apart. - For an Optic, sight across (through) the glass through 2 opposing scribe lines, then rotate the Optic while using a Machinist Square to align the MPF and Optic scribe lines. | Equal measurements at opposing scribe lines Machinist’s Square Fig 135: Centering the MPF on the Mass / Optic |
| - Load Plungers Wipe the counterbore end of each plunger with Isopropanol and a Wipe. Using the Tweezers, load 4 Magnet/Standoff assemblies into the 4 Plungers, 2 North Magnets and 2 South Magnets. The Magnet end of each assembly rests within the Plunger counterbore. - The Magnet/Standoffs are held to the Plungers magnetically. | Counterbore Assembly held to Plunger Magnetically Fig 136: Plungers Empty and Loaded |
| Lay the loaded Plungers on the Optics Table around the Fixture per the Magnetic Polarity layout shown. - Note that the “X” pattern is rectangular and not square; this pattern defines the Magnet layout. | S N S N Fig 137: Magnet Polarity Layout |
| - Bond Magnets to Mass/Optic Load the EP30-2 Cartridge with Mix Tube attached, into the Gun Applicator. Pull the trigger on the Gun Applicator 1 full stroke, to purge the Mix Tube of under-mixed adhesive. Dispense a “quarter-sized” pool of Adhesive onto a small piece of clean UHV aluminum foil. Hold a Plunger loaded with a Magnet/Standoff assembly vertically, with the Magnet/Standoff end facing up. Clean the Standoff with Isopropanol and a Wipe. Dip the end of a Sewing Needle in the pool of Epoxy and withdraw it, leaving a tiny drop on the Needle tip. Apply approximately ½ mm of Epoxy to the center of the Standoff end. | Plunger in Bushing Fig 138: Loading Magnets into Fixture |
| Load the Plunger, Magnet/Standoff down, into the appropriate Bushing in the MPF. Slide the Plunger down within the Bushing until the Standoff contacts the Mass/Optic. Press down on the Plunger lightly with one finger for about 2 seconds, then release. Repeat steps 11-13 to load all 4 Plungers into the Placement Fixture. Allow the Epoxy to cure within the Fixture at room temperature for 24 hours. |  |
| Carefully remove the 4 Plungers from the MPF, and remove the MPF from the Mass/Optic. Center the Heat Lamp over the Fixture and adjust the height such that the Fixture surface is receiving 60°C, then allow the adhesive to cure for 4hr. | Fig 139: Heat Lamp Cure |

# Installing AOSEM Brackets

## Procedure

- Adjust the push and pull screws on each AOSEM bracket so that the front face of each AOSEM is as far back as possible.  Back off the earthquake stop on each AOSEM bracket.

- Attach the AOSEM brackets (as shown in Figure 37) to the HLTS Structure in this order:  Bottom Mass Lower AOSEM Alignment Assembly (D0901551), Bottom Mass Upper AOSEM Alignment Assembly (D0901552), Intermediate Mass Lower AOSEM Alignment Assembly (D0901551), Intermediate Mass Upper AOSEM Alignment Assembly (D0902024).  Record the serial number and position of each AOSEM on the HLTS Process Traveler.

Fig 141: AOSEM Bracket Arrangement

- Align each AOSEM vertically and laterally using the slots on the AOSEM brackets.  Look through the barrel of the AOSEM and align the magnet/dumbbell assembly with the center axis of the AOSEM.

- Adjust the earthquake stop on each bracket to a distance of approximately 0.75 mm (between ½ and ¾ of a turn of the ¼-20 earthquake stop screw) from the corresponding mass.  Tighten the hex nuts on the bracket to lock the earthquake stop in place.

- After all AOSEM brackets are installed, connect the in-vacuum cabling and related cable connectors and clamps.

# Installing the BOSEMs

## Documents

D070442	HLTS Overall Assembly

## Materials

| Qty | U | ID | Description |
| --- | --- | --- | --- |
| 6 | Ea | D060218 | BOSEM Assembly |
| 24 | Ea | NA | Socket Head Cap Screw 4-40 x 7/8” SSTL |
| 24 | Ea | NA | Flat Washer #4 SSTL |

## Procedure

| - Install the 6 BOSEMs in the arrangement shown, using: 24 Socket Head Cap Screws - 4-40 x 7/8” SSTL 24 Flat Washers #4 SSTL Torque to 5 in-lb - It may be necessary to install the end BOSEM opposite from the end shown, due to space requirements inside the HAM chamber. Use a non-magnetic hex key to avoid attracting the hex key to the Magnet and Flag Assembly. - Be careful not to bend the Magnet Flags. - Record the serial number and position of each BOSEM in ICS. | Fig 140: BOSEMs Installed |
| --- | --- |

- After all 6 BOSEMs are attached, install the in-vacuum cabling and related cable connectors and clamps.  Use silver-plated SHCS in threaded holes in the HLTS Structure.

# Aligning OSEMs

## Procedure

- Connect the in-vacuum cabling for the AOSEMs and BOSEMs to the Triples Test Stand.  Confirm that each cable is connected correctly.

- Using the MEDM screens on the Triples Test Stand, read the open light voltage for each BOSEM.  Adjust the BOSEM using the SHCS connecting it to the tablecloth for rough positioning.  Use the PEEK hex nuts for fine adjustment.  Position the BOSEM so that the Triples Test Stand indicates 50% of the open light value.

- Using the MEDM screens on the Triples Test Stand, read the open light voltage for each AOSEM, starting with those for the Intermediate Mass.  Use the push and pull screws on each bracket to adjust the position of each AOSEM.  Position each AOSEM so that the Triples Test Stand indicates 50% of the open light value.