LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T1000256-v1	LIGO	14/5/10

Test Procedure for RF Distribution Amplifier

Paul Schwinberg and Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure descrbes the test of proper operation of the RF Distribution Amplifier. The unused outputs should always be properly terminated.

# Test Equipment

- Voltmeter

- Oscilloscope

- Stanford Research SR785 analyzer

- Tektronix AFG3101 function generator

- RF Power Meter HP E4418A

- Board Schematics--http://ilog.ligo-wa.caltech.edu:7285/advligo/RfDistributionAmplifier1U

# Tests

The RF Distribution Amplifier comes with a number of different power supply boards so I will assume that we are using the latest which is the Low Noise Power Module (D0901846) with the RF Distribution  Amplifier :Interface (D1000064).

- Verify the proper current draw.  Using a bench DC supply apply +- 24Volts to P7 and +- 17 Volts to P6 of the low noise power Module (D0901846).  Measure the current draw of the board.

+24 Volt current ___________0.1 A Nom.  	 	-24 Volt current _________0.0 A Nom.

+17 Volt current ______________     less than 1.1 A

-17  Volt current ______________      less than 0.01 A

- On the low noise power module check the voltage on TP 1-13.

TP1 ( +17V ) ______________		TP2 (-17V ) _______________

TP3 , 4 ( GND )				TP5 (+ 5V)________________

TP6 (-15V ) ________________		TP7 (+24V ) ________________

TP8 ( GND )				TP9 ( -24V ) ____________________

TP10 ( GND )				TP11 (+15V ) ____________________

TP12 (+VREF ) _____________		TP13  (-VREF ) ____________________

- If  TP 1 , 2, 7 , 9 and 8 are correct then pin 5 on U1 and U7,  ( OK  ) should be

Logic high ~3Volts.	Confirm._____________________

- The noise on TP 12, 13, 11 and 6 should be measured with a SR785 using an

rms power spectrum.

TP12  noise ____________________less than 200 nVrms/sqrt Hz at 140 Hz

TP13 noise _____________________less than 900 nVrms/sqrt Hz at 140 Hz

TP11 noise _____________________ less than 1 uVrms/sqrt Hz at 140 Hz

TP6 noise ______________________ less than 1 uVrms/sqrt Hz at 140 Hz.

- Test the power monitors by applying a 30 MHz, 10 dBm rf signal through an attenuator to J1.  Measure the output voltages mon1, mon2 and mon3  and with a RF power meter measure the RF power at the 13 dBm nominal outputs.

Mon1

| Nom output pwr | Input pwr dBm | Mon volt (M) | Measured volt. | Measured Pwr |
| --- | --- | --- | --- | --- |
| 13 dBm |  | 2.9V (0.725) |  |  |
| 10 dBm |  | 3.2V (0.800) |  |  |
| 7 dBm |  | 3.5V (0.875) |  |  |
| 0 dBm |  | 4.2V (1.05) |  |  |
| -10 dBm |  | 5.2V (1.30) |  |  |
| -20 dBm |  | 6.2V (1.55) |  |  |

Mon 2

| Nom output pwr | Input pwr dBm | Mon volt (M) | Measured volt | Measured Pwr |
| --- | --- | --- | --- | --- |
| 13 dBm |  | 2.9V (0.725) |  |  |
| 10 dBm |  | 3.2V (0.800) |  |  |
| 7 dBm |  | 3.5V (0.875) |  |  |
| 0 dBm |  | 4.2V (1.05) |  |  |
| -10 dBm |  | 5.2V (1.30) |  |  |
| -20 dBm |  | 6.2V (1.55) |  |  |

Mon3

| Nom output pwr | Input pwr dBm | Mon volt (M) | Measured volt | Measured Pwr |
| --- | --- | --- | --- | --- |
| 13 dBm |  | 2.9V (0.725) |  |  |
| 10 dBm |  | 3.2V (0.800) |  |  |
| 7 dBm |  | 3.5V (0.875) |  |  |
| 0 dBm |  | 4.2V (1.05) |  |  |
| -10 dBm |  | 5.2V (1.30) |  |  |
| -20 dBm |  | 6.2V (1.55) |  |  |

- Test the noise on the RF power monitors by applying a 30 MHz, 10 dBm rf  signal to J1 and measure the noise on mon1, mon2, and mon3 with  a SR785 using an rms power spectrum.  Power monitor sensitivity is -100 mV/dBm. Compare with power monitor driven directly by 30 MHz, RF signal, remember to adjust the power to give the same dc output,  the noise should be within 3dB.

Mon 1

| Offset in Hz | AM spec. | DD RF pwr | DC out | DD measured | Measured |
| --- | --- | --- | --- | --- | --- |
| 10 Hz | -140 dBc/Hz |  |  |  |  |
| 100 Hz | -150 dBc/Hz |  |  |  |  |
| 1 kHz | -150 dBc/Hz |  |  |  |  |

Mon 2

| Offset | AM spec. | DD RF pwr | DC out | DD measured | Measured |
| --- | --- | --- | --- | --- | --- |
| 10 Hz | -140 dBc/Hz |  |  |  |  |
| 100 Hz | -150 dBc/Hz |  |  |  |  |
| 1 kHz | -150 dBc/Hz |  |  |  |  |

Mon 3

| Offset | AM spec. | DD RF pwr | DC out | DD measured | Measured |
| --- | --- | --- | --- | --- | --- |
| 10 Hz | -140 dBc/Hz |  |  |  |  |
| 100 Hz | -150 dBc/Hz |  |  |  |  |
| 1 kHz | -150 dBc/Hz |  |  |  |  |

7) Measure  the Phase noise of the RF Oscillator Source driving the RF Distribution Amplifier , with   1PPS locking using the Wenzel single channel phase noise measurement technique (3.5.3), Figure 3.5.2-1, which can be found  at

http://www.wenzel.com/pdffiles1/BP1000Manual/BP_1000_v101_2_.pdf .

A reasonable FFT analyzer is the SR785, which can be set to measure power units if you start in Display Setup.  A Reference Source must be provided which can be just a Wenzel crystal oscillator of frequency close enough to lock, properly powered and connected to the Wenzel phase noise measurement system. The output of the RF Distribution Amplifier will need to be attenuated to the amplitude needed by the wenzel phase noise measurement system (about 10 dBm ). Test all the outputs that have different amplifiers. Compare to the Phase noise of the RF Oscillator Source alone, it should be within 3dB.

J2-J7

| Offset freq. Hz | Phase noise spec. | RF osc. phase noise | RF osc mit amp noise |
| --- | --- | --- | --- |
| 10 Hz | -110 dBc/Hz |  |  |
| 100 Hz | -140 dBc/Hz |  |  |
| 1 kHz | -160 dBc/Hz |  |  |

J8

| Offset freq. Hz | Phase noise spec. | RF osc. phase noise | RF osc mit amp noise |
| --- | --- | --- | --- |
| 10 Hz | -110 dBc/Hz |  |  |
| 100 Hz | -140 dBc/Hz |  |  |
| 1 kHz | -160 dBc/Hz |  |  |

J9

| Offset freq. Hz | Phase noise spec. | RF osc. phase noise | RF osc mit amp noise |
| --- | --- | --- | --- |
| 10 Hz | -110 dBc/Hz |  |  |
| 100 Hz | -140 dBc/Hz |  |  |
| 1 kHz | -160 dBc/Hz |  |  |