## Sheet1
Pressure Contribution from Flag Hydrocarbons
C62910 |  |  |  | 40M Lab RGA Scan Results
|  | Description: |  | Picomotors | Date: | 2010-06-29
|  |  | Oven Used: | C
AMU 41 | 5.73e-15 | amps | from RGA scan listing
AMU 43 | 1.57e-15 | amps | from RGA scan listing
AMU 53 | 1.01e-15 | amps | from RGA scan listing
AMU 55 | 3.03e-15 | amps | from RGA scan listing
AMU 57 | 7.83e-16 | amps | from RGA scan listing
Sum Flag H/C AMUs | 1.2123e-14 | amps
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
AMU 40 (w/leak open) | 1e-14 | amps
AMU 40 (background) | 7e-15 | amps
Calib leak contributes | 3e-15 | amps | = (w/leak open) - (background)
Flag H/C Outgassing | 9.5368e-10 | torr l/s | = (Sum Flag H/C AMUs) x (Calib leak rate)/(Calib leak contrib.)
Test item surf area | 8 | ea
Normalized outgassing | 1.1921e-10 | torr l/s-cm2 |  | = Flag H/C Outgassing/Test item surf area
Full description: | 8 New focus 830-X UHV Picomotors
Pre-scan bake: | 100C for 48 Hrs.