## Outgassing Rates
Master Bond Epoxy Adhesive EP30-2
Species | AMU | Outgassing Rate (Torr-L/s/cm^2)
H | 2 | 1.2214e-12
He | 4 | 7.7545e-14
C+ | 12 | 1.7949e-13
N+, N2+ | 14 | 6.1419e-13
water (O+) | 16 | 1.004e-12
water (HO+) | 17 | 2.2296e-12
water (H2O+) | 18 | 1.7949e-12
CO2++ | 22 | 2.9307e-14
N2 | 28 | 2.0894e-11
O2 | 32 | 6.1559e-12
Ar | 40 | 5.9035e-13
hydrocarbon | 41 | 1.9071e-13
hydrocarbon | 43 | 3.113e-13
CO2 | 44 | 3.099e-11
hydrocarbon | 53 | 1.4023e-17
hydrocarbon | 55 | 1.4023e-17
hydrocarbon | 57 | 1.4023e-17
Sum HC |  | 5.0205e-13
*caveat: a peak detection algorithm is not used. The values above assume that there isn't a significant shift in the peak from the nominal AMU value.
outgassing rates for the residual gas estimate (E0900398)
| AMU |  | Outgassing Rate (Torr-L/s/cm^2)
hydrogen (H2) | 2 |  | 1.2214e-12
water (H2O) at 1000 hr | max of 16, 17, 18 |  | 2.2296e-12
air (N2) | 28 |  | 2.0894e-11
CO2 | 44 |  | 3.099e-11
hydrocarbons | sum of 41, 43, 53, 55, 57 |  | 5.0205e-13
Note that this approximation is not based on a least squares fit to the cracked patterns of multiple gases. It is simply an approximate assignment of the largest peak in a cracking pattern to the most likely gas source.

## EP30-2 Applications
APPROVED APPLICATIONS FOR MASTER BOND EP30-2 ADHESIVE IN ALIGO
[10 empty] | #suspensions per Vacuum Volume |  |  | Directly Exposed to the UHV System | [5 empty] | for use in E1000453
[13 empty] | Adhesive Volume (mm^3) |  |  | Adhesive Surface (mm^2) |  |  | Bond Area (mm^2)
dwg | Title | Notes on bond | Suspension | Qty per Susp | bond area ea (mm^2) | area each (mm^2) | thickness (mm) | mm^3 per sus | mm^2 per sus | Vertex | Diagonal | End | Vertex | Diagonal | End | Vertex | Diagonal | End | Vertex | Diagonal | End
D080479-v3 | ETM/ITM PM Break-Off Prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | ITM, ETM | 2 | 240 | 10 | 0.05 | 0.5 | 10 | 2 | 2 | 1 | 1 | 1 | 0.5 | 20 | 20 | 10 | 960 | 960 | 480
D080750-v1 | ERM/CP prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | ITM, ETM | 2 | 100 | 10 | 0.05 | 0.5 | 10 | 2 | 2 | 1 | 1 | 1 | 0.5 | 20 | 20 | 10 | 400 | 400 | 200
D080765-v2 | FM/BS 1st prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | BS, FM | 2 | 290 | 10 | 0.05 | 20 | 40 | 1 | 3 | 0 | 20 | 60 | 0 | 40 | 120 | 0 | 580 | 1740 | 0
D0902368-v2 | FM/BS 2nd prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | BS, FM | 2 | 129 | 10 | 0.05 | 20 | 40 | 1 | 3 | 0 | 20 | 60 | 0 | 40 | 120 | 0 | 258 | 774 | 0
D070234-v2 | OSEM Magnet Assy, Penultimate Reaction Mass |  | ITM, ETM, BS, FM | 4 | 5.3564 |  | -- | 12.169 | 243.3792 | 3 | 5 | 1 | 36.5069 | 60.8448 | 12.169 | 730.1375 | 1216.8959 | 243.3792 | 64.277 | 107.1283 | 21.4257
| D070237-v2, magnet base | .05 mm thick over entire base area |  | 4 |  | 52.8102 | 0.05
| D070238-v2, steel disk | .05 mm thick over one end of the cylinder |  | 12 |  | 2.6782 | 0.05
D070441-v4 | HLTS prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | HLTS | 2 | 187.44 | 10 | 0.05 | 0.5 | 10 | 1 | 1 | 0 | 0.5 | 0.5 | 0 | 10 | 10 | 0 | 374.88 | 374.88 | 0
D0901286 | HLTS secondary metal prism | 0.5 mm wide by 5 mm long at 4 corners of each prism | HLTS | 2 | 57.912 | 10 | 0.05 | 0.5 | 10 | 1 | 1 | 0 | 0.5 | 0.5 | 0 | 10 | 10 | 0 | 115.824 | 115.824 | 0
D0810033-v4 | HSTS prism attachment | 0.5 mm wide by 5 mm long at 4 corners of each prism | HSTS | 2 | 113.2 | 10 | 0.05 | 0.5 | 10 | 7 | 7 | 0 | 3.5 | 3.5 | 0 | 70 | 70 | 0 | 1584.8 | 1584.8 | 0
D0901278 | HSTS secondary metal prism | 0.5 mm wide by 5 mm long at 4 corners of each prism | HSTS | 2 | 46.2 | 10 | 0.05 | 0.5 | 10 | 7 | 7 | 0 | 3.5 | 3.5 | 0 | 70 | 70 | 0 | 646.8 | 646.8 | 0
D0902432 | Magnet StandoffAssy | .05 mm thick over entire circular area | HLTS, HSTS | 4 | 3.1624 | 6.3247 | 0.05 | 0.31624 | 6.3247 | 8 | 8 | 0 | 2.5299 | 2.5299 | 0 | 50.5978 | 50.5978 | 0 | 101.1956 | 101.1956 | 0
D0901066 | AOSEM optical filters | ~1.4 mm dia. x .05 thick at 1 spot | ITM, ETM | 4 | 1 | 1.5394 | 0.05 | 0.076969 | 1.5394 | 2 | 2 | 1 | 0.15394 | 0.15394 | 0.076969 | 3.0788 | 3.0788 | 1.5394 | 8 | 8 | 4
|  |  | HLTS | 8 | 1 | 1.5394 | 0.05 | 0.076969 | 1.5394 | 1 | 1 | 0 | 0.076969 | 0.076969 | 0 | 1.5394 | 1.5394 | 0 | 8 | 8 | 0
|  |  | HSTS | 8 | 1 | 1.5394 | 0.05 | 0.076969 | 1.5394 | 7 | 7 | 0 | 0.53878 | 0.53878 | 0 | 10.7757 | 10.7757 | 0 | 56 | 56 | 0
TBD | Acoustic Mode Damper (AMD) attachment | 10 x 10 x .05 | ITM, ETM | 4 | 100 | 100 | 0.05 | 5 | 100 | 2 | 2 | 1 | 10 | 10 | 5 | 200 | 200 | 100 | 800 | 800 | 400
[12 empty] | TOTALS (cm^3 or cm^2): | 0.099806 | 0.20414 | 0.018246 | 12.7613 | 19.2289 | 3.6492 | 59.5778 | 76.7663 | 11.0543
Note that these adhesive surface area amounts in each of the LIGO vacuum volumes contributes approximately the following partial pressures which are all acceptable:
|  |  |  | Vertex | End | Diagonal
|  |  | Hydrogen | 9.2329e-16 | 1.0469e-15 | 1.3958e-15
|  |  | Water (1000h) | 9.9703e-17 | 5.6525e-17 | 1.5073e-16
|  |  | Air | 3.3169e-14 | 3.7609e-14 | 5.0145e-14
|  |  | CO2 | 4.9196e-14 | 5.5782e-14 | 7.4376e-14
|  |  | hydrocarbons | 9.3765e-16 | 1.0632e-15 | 1.134e-15 | N.B.: projection from a background limited RGA measurement

## EP30-2 RGA scans ASCII
2010-11-04 | 2009 01:22:44 PM
Residual Gas Analyzer Software
Software Version | 3.005
Analog Scan Setup:
Data Points in Scan | 991
Units | Amps
Noise Floor | 0
CEM Status | OFF
Points Per AMU | 10
Scan Start Mass | 1 | amu
Scan Stop Mass | 100 | amu
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.1
Filament Current | 1 | mAmps |  | Amps | T-L/s/cm^2
|  |  | HC AMUs | bias
|  | 1e-19 |  | -9.3728e-16
1 | 3.96e-16 | 3.96e-16 |  | 1.3333e-15 | 1.8696e-13
1.1 | 1.67e-15 | 1.67e-15 |  | 2.6073e-15 | 3.6561e-13
1.2 | 7.13e-15 | 7.13e-15 |  | 8.0673e-15 | 1.1312e-12
1.3 | 1.5e-13 | 1.5e-13 |  | 1.5094e-13 | 2.1165e-11
1.4 | 3.88e-12 | 3.88e-12 |  | 3.8809e-12 | 5.4421e-10
1.5 | 6.14e-12 | 6.14e-12 |  | 6.1409e-12 | 8.6112e-10
1.6 | 4.52e-12 | 4.52e-12 |  | 4.5209e-12 | 6.3395e-10
1.7 | 3.03e-12 | 3.03e-12 |  | 3.0309e-12 | 4.2502e-10
1.8 | 7.33e-13 | 7.33e-13 |  | 7.3394e-13 | 1.0292e-10
1.9 | 4.12e-14 | 4.12e-14 |  | 4.2137e-14 | 5.9087e-12
2 | 8.71e-15 | 8.71e-15 |  | 9.6473e-15 | 1.3528e-12
2.1 | 2.96e-15 | 2.96e-15 |  | 3.8973e-15 | 5.465e-13
2.2 | 1.06e-15 | 1.06e-15 |  | 1.9973e-15 | 2.8007e-13
2.3 | -2.23e-16 | 1e-19 |  | 7.1428e-16 | 1.0016e-13
2.4 | -1.35e-16 | 1e-19 |  | 8.0228e-16 | 1.125e-13
2.5 | -2.97e-16 | 1e-19 |  | 6.4028e-16 | 8.9784e-14
2.6 | -2.98e-16 | 1e-19 |  | 6.3928e-16 | 8.9644e-14
2.7 | 6.53e-17 | 6.53e-17 |  | 1.0026e-15 | 1.4059e-13
2.8 | 2.36e-16 | 2.36e-16 |  | 1.1733e-15 | 1.6452e-13
2.9 | 4.93e-16 | 4.93e-16 |  | 1.4303e-15 | 2.0056e-13
3 | 7.94e-16 | 7.94e-16 |  | 1.7313e-15 | 2.4277e-13
3.1 | 3.18e-15 | 3.18e-15 |  | 4.1173e-15 | 5.7735e-13
3.2 | 7.86e-15 | 7.86e-15 |  | 8.7973e-15 | 1.2336e-12
3.3 | 1.32e-14 | 1.32e-14 |  | 1.4137e-14 | 1.9824e-12
3.4 | 1.73e-14 | 1.73e-14 |  | 1.8237e-14 | 2.5573e-12
3.5 | 1.75e-14 | 1.75e-14 |  | 1.8437e-14 | 2.5854e-12
3.6 | 1.57e-14 | 1.57e-14 |  | 1.6637e-14 | 2.333e-12
3.7 | 1.18e-14 | 1.18e-14 |  | 1.2737e-14 | 1.7861e-12
3.8 | 6.35e-15 | 6.35e-15 |  | 7.2873e-15 | 1.0219e-12
3.9 | 1.94e-15 | 1.94e-15 |  | 2.8773e-15 | 4.0347e-13
4 | 5.53e-16 | 5.53e-16 |  | 1.4903e-15 | 2.0898e-13
4.1 | -1.93e-16 | 1e-19 |  | 7.4428e-16 | 1.0437e-13
4.2 | -2.08e-16 | 1e-19 |  | 7.2928e-16 | 1.0226e-13
4.3 | -2.42e-17 | 1e-19 |  | 9.1308e-16 | 1.2804e-13
4.4 | -3.07e-16 | 1e-19 |  | 6.3028e-16 | 8.8382e-14
4.5 | -3.16e-16 | 1e-19 |  | 6.2128e-16 | 8.7119e-14
4.6 | -3.09e-16 | 1e-19 |  | 6.2828e-16 | 8.8101e-14
4.7 | -1.81e-16 | 1e-19 |  | 7.5628e-16 | 1.0605e-13
4.8 | -3.69e-17 | 1e-19 |  | 9.0038e-16 | 1.2626e-13
4.9 | -8.26e-17 | 1e-19 |  | 8.5468e-16 | 1.1985e-13
5 | -6.82e-17 | 1e-19 |  | 8.6908e-16 | 1.2187e-13
5.1 | -3.69e-16 | 1e-19 |  | 5.6828e-16 | 7.9688e-14
5.2 | -1.23e-16 | 1e-19 |  | 8.1428e-16 | 1.1418e-13
5.3 | -1.23e-16 | 1e-19 |  | 8.1428e-16 | 1.1418e-13
5.4 | -5.04e-16 | 1e-19 |  | 4.3328e-16 | 6.0757e-14
5.5 | -9.08e-17 | 1e-19 |  | 8.4648e-16 | 1.187e-13
5.6 | -2.53e-16 | 1e-19 |  | 6.8428e-16 | 9.5954e-14
5.7 | -1.42e-16 | 1e-19 |  | 7.9528e-16 | 1.1152e-13
5.8 | -2.15e-16 | 1e-19 |  | 7.2228e-16 | 1.0128e-13
5.9 | -2.51e-16 | 1e-19 |  | 6.8628e-16 | 9.6234e-14
6 | -1.03e-16 | 1e-19 |  | 8.3428e-16 | 1.1699e-13
6.1 | -1.23e-16 | 1e-19 |  | 8.1428e-16 | 1.1418e-13
6.2 | -1.69e-16 | 1e-19 |  | 7.6828e-16 | 1.0773e-13
6.3 | -8.61e-16 | 1e-19 |  | 7.628e-17 | 1.0696e-14
6.4 | -4.13e-16 | 1e-19 |  | 5.2428e-16 | 7.3518e-14
6.5 | -2.95e-16 | 1e-19 |  | 6.4228e-16 | 9.0064e-14
6.6 | -2.7e-16 | 1e-19 |  | 6.6728e-16 | 9.357e-14
6.7 | -2.21e-16 | 1e-19 |  | 7.1628e-16 | 1.0044e-13
6.8 | -3.44e-16 | 1e-19 |  | 5.9328e-16 | 8.3193e-14
6.9 | -2.39e-16 | 1e-19 |  | 6.9828e-16 | 9.7917e-14
7 | -3.44e-16 | 1e-19 |  | 5.9328e-16 | 8.3193e-14
7.1 | -1.67e-16 | 1e-19 |  | 7.7028e-16 | 1.0801e-13
7.2 | -3.4e-16 | 1e-19 |  | 5.9728e-16 | 8.3754e-14
7.3 | -7.96e-17 | 1e-19 |  | 8.5768e-16 | 1.2027e-13
7.4 | -4.53e-17 | 1e-19 |  | 8.9198e-16 | 1.2508e-13
7.5 | -2.56e-16 | 1e-19 |  | 6.8128e-16 | 9.5533e-14
7.6 | 1.66e-17 | 1.66e-17 |  | 9.5388e-16 | 1.3376e-13
7.7 | -2.63e-16 | 1e-19 |  | 6.7428e-16 | 9.4551e-14
7.8 | -3.88e-16 | 1e-19 |  | 5.4928e-16 | 7.7023e-14
7.9 | -3.6e-16 | 1e-19 |  | 5.7728e-16 | 8.095e-14
8 | -9.08e-17 | 1e-19 |  | 8.4648e-16 | 1.187e-13
8.1 | -1.98e-16 | 1e-19 |  | 7.3928e-16 | 1.0367e-13
8.2 | -2.79e-16 | 1e-19 |  | 6.5828e-16 | 9.2308e-14
8.3 | -1.86e-16 | 1e-19 |  | 7.5128e-16 | 1.0535e-13
8.4 | -2.45e-16 | 1e-19 |  | 6.9228e-16 | 9.7076e-14
8.5 | -4.12e-16 | 1e-19 |  | 5.2528e-16 | 7.3658e-14
8.6 | -2.32e-17 | 1e-19 |  | 9.1408e-16 | 1.2818e-13
8.7 | -2.86e-16 | 1e-19 |  | 6.5128e-16 | 9.1326e-14
8.8 | -2.17e-16 | 1e-19 |  | 7.2028e-16 | 1.01e-13
8.9 | -2.34e-16 | 1e-19 |  | 7.0328e-16 | 9.8618e-14
9 | -3.06e-16 | 1e-19 |  | 6.3128e-16 | 8.8522e-14
9.1 | -5.37e-16 | 1e-19 |  | 4.0028e-16 | 5.613e-14
9.2 | -2.38e-17 | 1e-19 |  | 9.1348e-16 | 1.2809e-13
9.3 | -4.63e-16 | 1e-19 |  | 4.7428e-16 | 6.6506e-14
9.4 | -2.63e-16 | 1e-19 |  | 6.7428e-16 | 9.4551e-14
9.5 | -3.53e-16 | 1e-19 |  | 5.8428e-16 | 8.1931e-14
9.6 | -2.22e-16 | 1e-19 |  | 7.1528e-16 | 1.003e-13
9.7 | -3.52e-16 | 1e-19 |  | 5.8528e-16 | 8.2071e-14
9.8 | -1.12e-16 | 1e-19 |  | 8.2528e-16 | 1.1573e-13
9.9 | -2.29e-16 | 1e-19 |  | 7.0828e-16 | 9.9319e-14
10 | -2.8e-16 | 1e-19 |  | 6.5728e-16 | 9.2168e-14
10.1 | -4.08e-16 | 1e-19 |  | 5.2928e-16 | 7.4219e-14
10.2 | -5.03e-17 | 1e-19 |  | 8.8698e-16 | 1.2438e-13
10.3 | -3.4e-16 | 1e-19 |  | 5.9728e-16 | 8.3754e-14
10.4 | 1.5e-17 | 1.5e-17 |  | 9.5228e-16 | 1.3353e-13
10.5 | -8.11e-17 | 1e-19 |  | 8.5618e-16 | 1.2006e-13
10.6 | -5.42e-16 | 1e-19 |  | 3.9528e-16 | 5.5428e-14
10.7 | -2.87e-16 | 1e-19 |  | 6.5028e-16 | 9.1186e-14
10.8 | 3.52e-17 | 3.52e-17 |  | 9.7248e-16 | 1.3637e-13
10.9 | 3.69e-17 | 3.69e-17 |  | 9.7418e-16 | 1.3661e-13
11 | 5.54e-16 | 5.54e-16 |  | 1.4913e-15 | 2.0912e-13
11.1 | 1.16e-15 | 1.16e-15 |  | 2.0973e-15 | 2.9409e-13
11.2 | 5.23e-15 | 5.23e-15 |  | 6.1673e-15 | 8.6481e-13
11.3 | 1.28e-14 | 1.28e-14 |  | 1.3737e-14 | 1.9263e-12
11.4 | 2e-14 | 2e-14 |  | 2.0937e-14 | 2.9359e-12
11.5 | 2.33e-14 | 2.33e-14 |  | 2.4237e-14 | 3.3987e-12
11.6 | 2.27e-14 | 2.27e-14 |  | 2.3637e-14 | 3.3146e-12
11.7 | 1.87e-14 | 1.87e-14 |  | 1.9637e-14 | 2.7537e-12
11.8 | 1.18e-14 | 1.18e-14 |  | 1.2737e-14 | 1.7861e-12
11.9 | 4.8e-15 | 4.8e-15 |  | 5.7373e-15 | 8.0451e-13
12 | 1.28e-15 | 1.28e-15 |  | 2.2173e-15 | 3.1092e-13
12.1 | 5.77e-16 | 5.77e-16 |  | 1.5143e-15 | 2.1234e-13
12.2 | 1.71e-16 | 1.71e-16 |  | 1.1083e-15 | 1.5541e-13
12.3 | 7.21e-16 | 7.21e-16 |  | 1.6583e-15 | 2.3253e-13
12.4 | 8.01e-16 | 8.01e-16 |  | 1.7383e-15 | 2.4375e-13
12.5 | 1e-15 | 1e-15 |  | 1.9373e-15 | 2.7166e-13
12.6 | 7.46e-16 | 7.46e-16 |  | 1.6833e-15 | 2.3604e-13
12.7 | 7.18e-16 | 7.18e-16 |  | 1.6553e-15 | 2.3211e-13
12.8 | 6.26e-16 | 6.26e-16 |  | 1.5633e-15 | 2.1921e-13
12.9 | 6.57e-16 | 6.57e-16 |  | 1.5943e-15 | 2.2356e-13
13 | 1.25e-15 | 1.25e-15 |  | 2.1873e-15 | 3.0671e-13
13.1 | 3.73e-15 | 3.73e-15 |  | 4.6673e-15 | 6.5447e-13
13.2 | 1.27e-14 | 1.27e-14 |  | 1.3637e-14 | 1.9123e-12
13.3 | 4.64e-14 | 4.64e-14 |  | 4.7337e-14 | 6.6379e-12
13.4 | 8.39e-14 | 8.39e-14 |  | 8.4837e-14 | 1.1896e-11
13.5 | 9.36e-14 | 9.36e-14 |  | 9.4537e-14 | 1.3257e-11
13.6 | 8.81e-14 | 8.81e-14 |  | 8.9037e-14 | 1.2485e-11
13.7 | 7.26e-14 | 7.26e-14 |  | 7.3537e-14 | 1.0312e-11
13.8 | 4.66e-14 | 4.66e-14 |  | 4.7537e-14 | 6.666e-12
13.9 | 1.3e-14 | 1.3e-14 |  | 1.3937e-14 | 1.9544e-12
14 | 4.38e-15 | 4.38e-15 |  | 5.3173e-15 | 7.4562e-13
14.1 | 1.51e-15 | 1.51e-15 |  | 2.4473e-15 | 3.4317e-13
14.2 | 7.73e-16 | 7.73e-16 |  | 1.7103e-15 | 2.3983e-13
14.3 | 1.12e-15 | 1.12e-15 |  | 2.0573e-15 | 2.8848e-13
14.4 | 2.01e-15 | 2.01e-15 |  | 2.9473e-15 | 4.1328e-13
14.5 | 2.18e-15 | 2.18e-15 |  | 3.1173e-15 | 4.3712e-13
14.6 | 2.43e-15 | 2.43e-15 |  | 3.3673e-15 | 4.7218e-13
14.7 | 2.59e-15 | 2.59e-15 |  | 3.5273e-15 | 4.9462e-13
14.8 | 1.94e-15 | 1.94e-15 |  | 2.8773e-15 | 4.0347e-13
14.9 | 1.35e-15 | 1.35e-15 |  | 2.2873e-15 | 3.2074e-13
15 | 1.56e-15 | 1.56e-15 |  | 2.4973e-15 | 3.5018e-13
15.1 | 2.53e-15 | 2.53e-15 |  | 3.4673e-15 | 4.862e-13
15.2 | 7.51e-15 | 7.51e-15 |  | 8.4473e-15 | 1.1845e-12
15.3 | 2.41e-14 | 2.41e-14 |  | 2.5037e-14 | 3.5109e-12
15.4 | 5.22e-14 | 5.22e-14 |  | 5.3137e-14 | 7.4512e-12
15.5 | 5.92e-14 | 5.92e-14 |  | 6.0137e-14 | 8.4328e-12
15.6 | 5.78e-14 | 5.78e-14 |  | 5.8737e-14 | 8.2365e-12
15.7 | 5.04e-14 | 5.04e-14 |  | 5.1337e-14 | 7.1988e-12
15.8 | 3.68e-14 | 3.68e-14 |  | 3.7737e-14 | 5.2917e-12
15.9 | 1.37e-14 | 1.37e-14 |  | 1.4637e-14 | 2.0525e-12
16 | 7.16e-15 | 7.16e-15 |  | 8.0973e-15 | 1.1354e-12
16.1 | 7.63e-15 | 7.63e-15 |  | 8.5673e-15 | 1.2014e-12
16.2 | 2.09e-14 | 2.09e-14 |  | 2.1837e-14 | 3.0621e-12
16.3 | 8.72e-14 | 8.72e-14 |  | 8.8137e-14 | 1.2359e-11
16.4 | 1.74e-13 | 1.74e-13 |  | 1.7494e-13 | 2.4531e-11
16.5 | 2.03e-13 | 2.03e-13 |  | 2.0394e-13 | 2.8597e-11
16.6 | 1.81e-13 | 1.81e-13 |  | 1.8194e-13 | 2.5512e-11
16.7 | 1.61e-13 | 1.61e-13 |  | 1.6194e-13 | 2.2708e-11
16.8 | 1.16e-13 | 1.16e-13 |  | 1.1694e-13 | 1.6398e-11
16.9 | 3.61e-14 | 3.61e-14 |  | 3.7037e-14 | 5.1936e-12
17 | 1.59e-14 | 1.59e-14 |  | 1.6837e-14 | 2.361e-12
17.1 | 2.69e-14 | 2.69e-14 |  | 2.7837e-14 | 3.9035e-12
17.2 | 9.44e-14 | 9.44e-14 |  | 9.5337e-14 | 1.3369e-11
17.3 | 3.59e-13 | 3.59e-13 |  | 3.5994e-13 | 5.0472e-11
17.4 | 6.94e-13 | 6.94e-13 |  | 6.9494e-13 | 9.7448e-11
17.5 | 7.76e-13 | 7.76e-13 |  | 7.7694e-13 | 1.0895e-10
17.6 | 7.05e-13 | 7.05e-13 |  | 7.0594e-13 | 9.8991e-11
17.7 | 6.15e-13 | 6.15e-13 |  | 6.1594e-13 | 8.637e-11
17.8 | 4.3e-13 | 4.3e-13 |  | 4.3094e-13 | 6.0429e-11
17.9 | 1.08e-13 | 1.08e-13 |  | 1.0894e-13 | 1.5276e-11
18 | 1.28e-14 | 1.28e-14 |  | 1.3737e-14 | 1.9263e-12
18.1 | 5.17e-15 | 5.17e-15 |  | 6.1073e-15 | 8.564e-13
18.2 | 3.67e-15 | 3.67e-15 |  | 4.6073e-15 | 6.4606e-13
18.3 | 6.38e-15 | 6.38e-15 |  | 7.3173e-15 | 1.0261e-12
18.4 | 1.1e-14 | 1.1e-14 |  | 1.1937e-14 | 1.6739e-12
18.5 | 1.55e-14 | 1.55e-14 |  | 1.6437e-14 | 2.3049e-12
18.6 | 1.66e-14 | 1.66e-14 |  | 1.7537e-14 | 2.4592e-12
18.7 | 1.76e-14 | 1.76e-14 |  | 1.8537e-14 | 2.5994e-12
18.8 | 1.41e-14 | 1.41e-14 |  | 1.5037e-14 | 2.1086e-12
18.9 | 7.39e-15 | 7.39e-15 |  | 8.3273e-15 | 1.1677e-12
19 | 2.63e-15 | 2.63e-15 |  | 3.5673e-15 | 5.0022e-13
19.1 | 9.91e-16 | 9.91e-16 |  | 1.9283e-15 | 2.7039e-13
19.2 | 1.14e-15 | 1.14e-15 |  | 2.0773e-15 | 2.9129e-13
19.3 | 1.82e-15 | 1.82e-15 |  | 2.7573e-15 | 3.8664e-13
19.4 | 5.29e-15 | 5.29e-15 |  | 6.2273e-15 | 8.7323e-13
19.5 | 9.68e-15 | 9.68e-15 |  | 1.0617e-14 | 1.4888e-12
19.6 | 1.22e-14 | 1.22e-14 |  | 1.3137e-14 | 1.8422e-12
19.7 | 1.17e-14 | 1.17e-14 |  | 1.2637e-14 | 1.7721e-12
19.8 | 1.12e-14 | 1.12e-14 |  | 1.2137e-14 | 1.702e-12
19.9 | 7.14e-15 | 7.14e-15 |  | 8.0773e-15 | 1.1326e-12
20 | 2.24e-15 | 2.24e-15 |  | 3.1773e-15 | 4.4554e-13
20.1 | 5.47e-16 | 5.47e-16 |  | 1.4843e-15 | 2.0813e-13
20.2 | 6.3e-17 | 6.3e-17 |  | 1.0003e-15 | 1.4027e-13
20.3 | -1.76e-16 | 1e-19 |  | 7.6128e-16 | 1.0675e-13
20.4 | -4e-17 | 1e-19 |  | 8.9728e-16 | 1.2582e-13
20.5 | -3.8e-16 | 1e-19 |  | 5.5728e-16 | 7.8145e-14
20.6 | -1e-16 | 1e-19 |  | 8.3728e-16 | 1.1741e-13
20.7 | -4.28e-16 | 1e-19 |  | 5.0928e-16 | 7.1414e-14
20.8 | -2.85e-16 | 1e-19 |  | 6.5228e-16 | 9.1466e-14
20.9 | -9.44e-17 | 1e-19 |  | 8.4288e-16 | 1.1819e-13
21 | -1.68e-16 | 1e-19 |  | 7.6928e-16 | 1.0787e-13
21.1 | -3.42e-16 | 1e-19 |  | 5.9528e-16 | 8.3474e-14
21.2 | -2.01e-16 | 1e-19 |  | 7.3628e-16 | 1.0325e-13
21.3 | 5.24e-17 | 5.24e-17 |  | 9.8968e-16 | 1.3878e-13
21.4 | 6.21e-16 | 6.21e-16 |  | 1.5583e-15 | 2.1851e-13
21.5 | 7.56e-16 | 7.56e-16 |  | 1.6933e-15 | 2.3744e-13
21.6 | 8.53e-16 | 8.53e-16 |  | 1.7903e-15 | 2.5104e-13
21.7 | 1.12e-15 | 1.12e-15 |  | 2.0573e-15 | 2.8848e-13
21.8 | 1.1e-15 | 1.1e-15 |  | 2.0373e-15 | 2.8568e-13
21.9 | 6.14e-16 | 6.14e-16 |  | 1.5513e-15 | 2.1753e-13
22 | 2.09e-16 | 2.09e-16 |  | 1.1463e-15 | 1.6074e-13
22.1 | -1.73e-16 | 1e-19 |  | 7.6428e-16 | 1.0717e-13
22.2 | -1.04e-16 | 1e-19 |  | 8.3328e-16 | 1.1685e-13
22.3 | 1.35e-17 | 1.35e-17 |  | 9.5078e-16 | 1.3332e-13
22.4 | -2.58e-16 | 1e-19 |  | 6.7928e-16 | 9.5253e-14
22.5 | -5.14e-17 | 1e-19 |  | 8.8588e-16 | 1.2422e-13
22.6 | -3.38e-16 | 1e-19 |  | 5.9928e-16 | 8.4035e-14
22.7 | -2.05e-16 | 1e-19 |  | 7.3228e-16 | 1.0268e-13
22.8 | -4.68e-17 | 1e-19 |  | 8.9048e-16 | 1.2487e-13
22.9 | -1.83e-16 | 1e-19 |  | 7.5428e-16 | 1.0577e-13
23 | -2.94e-16 | 1e-19 |  | 6.4328e-16 | 9.0204e-14
23.1 | -1.9e-16 | 1e-19 |  | 7.4728e-16 | 1.0479e-13
23.2 | -1.78e-16 | 1e-19 |  | 7.5928e-16 | 1.0647e-13
23.3 | 9.27e-18 | 9.27e-18 |  | 9.4655e-16 | 1.3273e-13
23.4 | -3.29e-17 | 1e-19 |  | 9.0438e-16 | 1.2682e-13
23.5 | -1.48e-16 | 1e-19 |  | 7.8928e-16 | 1.1068e-13
23.6 | -6.09e-17 | 1e-19 |  | 8.7638e-16 | 1.2289e-13
23.7 | 6.85e-17 | 6.85e-17 |  | 1.0058e-15 | 1.4104e-13
23.8 | -3.81e-17 | 1e-19 |  | 8.9918e-16 | 1.2609e-13
23.9 | -1.57e-16 | 1e-19 |  | 7.8028e-16 | 1.0942e-13
24 | -2.25e-16 | 1e-19 |  | 7.1228e-16 | 9.988e-14
24.1 | -1.48e-16 | 1e-19 |  | 7.8928e-16 | 1.1068e-13
24.2 | -3.78e-16 | 1e-19 |  | 5.5928e-16 | 7.8425e-14
24.3 | 7.88e-17 | 7.88e-17 |  | 1.0161e-15 | 1.4248e-13
24.4 | -5.9e-18 | 1e-19 |  | 9.3138e-16 | 1.306e-13
24.5 | -2.59e-16 | 1e-19 |  | 6.7828e-16 | 9.5112e-14
24.6 | 9.06e-18 | 9.06e-18 |  | 9.4634e-16 | 1.327e-13
24.7 | 1.59e-16 | 1.59e-16 |  | 1.0963e-15 | 1.5373e-13
24.8 | 5.24e-16 | 5.24e-16 |  | 1.4613e-15 | 2.0491e-13
24.9 | 5.97e-16 | 5.97e-16 |  | 1.5343e-15 | 2.1515e-13
25 | 3.88e-16 | 3.88e-16 |  | 1.3253e-15 | 1.8584e-13
25.1 | 1.7e-16 | 1.7e-16 |  | 1.1073e-15 | 1.5527e-13
25.2 | 4e-16 | 4e-16 |  | 1.3373e-15 | 1.8752e-13
25.3 | -3.92e-17 | 1e-19 |  | 8.9808e-16 | 1.2593e-13
25.4 | 6.2e-16 | 6.2e-16 |  | 1.5573e-15 | 2.1837e-13
25.5 | 1.95e-15 | 1.95e-15 |  | 2.8873e-15 | 4.0487e-13
25.6 | 3.62e-15 | 3.62e-15 |  | 4.5573e-15 | 6.3905e-13
25.7 | 4.86e-15 | 4.86e-15 |  | 5.7973e-15 | 8.1293e-13
25.8 | 4.08e-15 | 4.08e-15 |  | 5.0173e-15 | 7.0355e-13
25.9 | 3.4e-15 | 3.4e-15 |  | 4.3373e-15 | 6.082e-13
26 | 1.95e-15 | 1.95e-15 |  | 2.8873e-15 | 4.0487e-13
26.1 | 1.08e-15 | 1.08e-15 |  | 2.0173e-15 | 2.8287e-13
26.2 | 7.19e-16 | 7.19e-16 |  | 1.6563e-15 | 2.3225e-13
26.3 | 1.36e-15 | 1.36e-15 |  | 2.2973e-15 | 3.2214e-13
26.4 | 2.83e-15 | 2.83e-15 |  | 3.7673e-15 | 5.2827e-13
26.5 | 7.28e-15 | 7.28e-15 |  | 8.2173e-15 | 1.1523e-12
26.6 | 1.43e-14 | 1.43e-14 |  | 1.5237e-14 | 2.1367e-12
26.7 | 1.97e-14 | 1.97e-14 |  | 2.0637e-14 | 2.8939e-12
26.8 | 2.12e-14 | 2.12e-14 |  | 2.2137e-14 | 3.1042e-12
26.9 | 1.97e-14 | 1.97e-14 |  | 2.0637e-14 | 2.8939e-12
27 | 1.85e-14 | 1.85e-14 |  | 1.9437e-14 | 2.7256e-12
27.1 | 2.64e-14 | 2.64e-14 |  | 2.7337e-14 | 3.8334e-12
27.2 | 4.65e-14 | 4.65e-14 |  | 4.7437e-14 | 6.6519e-12
27.3 | 1.04e-13 | 1.04e-13 |  | 1.0494e-13 | 1.4715e-11
27.4 | 3.06e-13 | 3.06e-13 |  | 3.0694e-13 | 4.3041e-11
27.5 | 8.23e-13 | 8.23e-13 |  | 8.2394e-13 | 1.1554e-10
27.6 | 1.44e-12 | 1.44e-12 |  | 1.4409e-12 | 2.0206e-10
27.7 | 1.41e-12 | 1.41e-12 |  | 1.4109e-12 | 1.9785e-10
27.8 | 1.19e-12 | 1.19e-12 |  | 1.1909e-12 | 1.67e-10
27.9 | 7.73e-13 | 7.73e-13 |  | 7.7394e-13 | 1.0853e-10
28 | 1.49e-13 | 1.49e-13 |  | 1.4994e-13 | 2.1025e-11
28.1 | 1.34e-14 | 1.34e-14 |  | 1.4337e-14 | 2.0105e-12
28.2 | 4.3e-15 | 4.3e-15 |  | 5.2373e-15 | 7.344e-13
28.3 | 1.35e-15 | 1.35e-15 |  | 2.2873e-15 | 3.2074e-13
28.4 | 1.5e-15 | 1.5e-15 |  | 2.4373e-15 | 3.4177e-13
28.5 | 3.41e-15 | 3.41e-15 |  | 4.3473e-15 | 6.096e-13
28.6 | 6.51e-15 | 6.51e-15 |  | 7.4473e-15 | 1.0443e-12
28.7 | 9.58e-15 | 9.58e-15 |  | 1.0517e-14 | 1.4748e-12
28.8 | 1.01e-14 | 1.01e-14 |  | 1.1037e-14 | 1.5477e-12
28.9 | 8.58e-15 | 8.58e-15 |  | 9.5173e-15 | 1.3346e-12
29 | 4.85e-15 | 4.85e-15 |  | 5.7873e-15 | 8.1153e-13
29.1 | 1.6e-15 | 1.6e-15 |  | 2.5373e-15 | 3.5579e-13
29.2 | 7.47e-16 | 7.47e-16 |  | 1.6843e-15 | 2.3618e-13
29.3 | 3.12e-16 | 3.12e-16 |  | 1.2493e-15 | 1.7518e-13
29.4 | 2.08e-16 | 2.08e-16 |  | 1.1453e-15 | 1.606e-13
29.5 | 1.26e-15 | 1.26e-15 |  | 2.1973e-15 | 3.0812e-13
29.6 | 3.08e-15 | 3.08e-15 |  | 4.0173e-15 | 5.6333e-13
29.7 | 4.62e-15 | 4.62e-15 |  | 5.5573e-15 | 7.7927e-13
29.8 | 4.47e-15 | 4.47e-15 |  | 5.4073e-15 | 7.5824e-13
29.9 | 4.04e-15 | 4.04e-15 |  | 4.9773e-15 | 6.9794e-13
30 | 2.62e-15 | 2.62e-15 |  | 3.5573e-15 | 4.9882e-13
30.1 | 5.96e-16 | 5.96e-16 |  | 1.5333e-15 | 2.1501e-13
30.2 | 3.66e-17 | 3.66e-17 |  | 9.7388e-16 | 1.3656e-13
30.3 | -1.01e-16 | 1e-19 |  | 8.3628e-16 | 1.1727e-13
30.4 | -2.18e-16 | 1e-19 |  | 7.1928e-16 | 1.0086e-13
30.5 | -1.47e-17 | 1e-19 |  | 9.2258e-16 | 1.2937e-13
30.6 | 3.81e-17 | 3.81e-17 |  | 9.7538e-16 | 1.3677e-13
30.7 | 1.63e-16 | 1.63e-16 |  | 1.1003e-15 | 1.5429e-13
30.8 | 6.26e-16 | 6.26e-16 |  | 1.5633e-15 | 2.1921e-13
30.9 | 1.01e-15 | 1.01e-15 |  | 1.9473e-15 | 2.7306e-13
31 | 1.12e-15 | 1.12e-15 |  | 2.0573e-15 | 2.8848e-13
31.1 | 2.35e-15 | 2.35e-15 |  | 3.2873e-15 | 4.6096e-13
31.2 | 3.54e-15 | 3.54e-15 |  | 4.4773e-15 | 6.2783e-13
31.3 | 5.21e-15 | 5.21e-15 |  | 6.1473e-15 | 8.6201e-13
31.4 | 1.11e-14 | 1.11e-14 |  | 1.2037e-14 | 1.6879e-12
31.5 | 3.06e-14 | 3.06e-14 |  | 3.1537e-14 | 4.4223e-12
31.6 | 9.28e-14 | 9.28e-14 |  | 9.3737e-14 | 1.3144e-11
31.7 | 1.42e-13 | 1.42e-13 |  | 1.4294e-13 | 2.0043e-11
31.8 | 1.39e-13 | 1.39e-13 |  | 1.3994e-13 | 1.9623e-11
31.9 | 1.02e-13 | 1.02e-13 |  | 1.0294e-13 | 1.4434e-11
32 | 4.39e-14 | 4.39e-14 |  | 4.4837e-14 | 6.2873e-12
32.1 | 1.01e-14 | 1.01e-14 |  | 1.1037e-14 | 1.5477e-12
32.2 | 3.26e-15 | 3.26e-15 |  | 4.1973e-15 | 5.8857e-13
32.3 | 1.19e-15 | 1.19e-15 |  | 2.1273e-15 | 2.983e-13
32.4 | 2.73e-16 | 2.73e-16 |  | 1.2103e-15 | 1.6971e-13
32.5 | 4.42e-17 | 4.42e-17 |  | 9.8148e-16 | 1.3763e-13
32.6 | -2.33e-16 | 1e-19 |  | 7.0428e-16 | 9.8758e-14
32.7 | -8.85e-17 | 1e-19 |  | 8.4878e-16 | 1.1902e-13
32.8 | 4.51e-17 | 4.51e-17 |  | 9.8238e-16 | 1.3776e-13
32.9 | -1.07e-16 | 1e-19 |  | 8.3028e-16 | 1.1643e-13
33 | -1.81e-16 | 1e-19 |  | 7.5628e-16 | 1.0605e-13
33.1 | -4.37e-16 | 1e-19 |  | 5.0028e-16 | 7.0152e-14
33.2 | -2.75e-16 | 1e-19 |  | 6.6228e-16 | 9.2869e-14
33.3 | 3.24e-17 | 3.24e-17 |  | 9.6968e-16 | 1.3597e-13
33.4 | -4.16e-16 | 1e-19 |  | 5.2128e-16 | 7.3097e-14
33.5 | -2.18e-16 | 1e-19 |  | 7.1928e-16 | 1.0086e-13
33.6 | 1.96e-17 | 1.96e-17 |  | 9.5688e-16 | 1.3418e-13
33.7 | 2.07e-16 | 2.07e-16 |  | 1.1443e-15 | 1.6046e-13
33.8 | 5.17e-16 | 5.17e-16 |  | 1.4543e-15 | 2.0393e-13
33.9 | 2.12e-16 | 2.12e-16 |  | 1.1493e-15 | 1.6116e-13
34 | 1.4e-16 | 1.4e-16 |  | 1.0773e-15 | 1.5106e-13
34.1 | -2.61e-17 | 1e-19 |  | 9.1118e-16 | 1.2777e-13
34.2 | -1.06e-16 | 1e-19 |  | 8.3128e-16 | 1.1657e-13
34.3 | -2.01e-16 | 1e-19 |  | 7.3628e-16 | 1.0325e-13
34.4 | -1.81e-16 | 1e-19 |  | 7.5628e-16 | 1.0605e-13
34.5 | -3.21e-16 | 1e-19 |  | 6.1628e-16 | 8.6418e-14
34.6 | 1.22e-16 | 1.22e-16 |  | 1.0593e-15 | 1.4854e-13
34.7 | 6.74e-18 | 6.74e-18 |  | 9.4402e-16 | 1.3238e-13
34.8 | -6.02e-17 | 1e-19 |  | 8.7708e-16 | 1.2299e-13
34.9 | 4.97e-17 | 4.97e-17 |  | 9.8698e-16 | 1.384e-13
35 | -3.65e-16 | 1e-19 |  | 5.7228e-16 | 8.0248e-14
35.1 | 8.09e-17 | 8.09e-17 |  | 1.0182e-15 | 1.4278e-13
35.2 | -1.2e-16 | 1e-19 |  | 8.1728e-16 | 1.146e-13
35.3 | -1.4e-16 | 1e-19 |  | 7.9728e-16 | 1.118e-13
35.4 | -4.4e-17 | 1e-19 |  | 8.9328e-16 | 1.2526e-13
35.5 | 1.86e-16 | 1.86e-16 |  | 1.1233e-15 | 1.5751e-13
35.6 | 6.73e-16 | 6.73e-16 |  | 1.6103e-15 | 2.258e-13
35.7 | 1.18e-15 | 1.18e-15 |  | 2.1173e-15 | 2.969e-13
35.8 | 1.92e-15 | 1.92e-15 |  | 2.8573e-15 | 4.0066e-13
35.9 | 1.8e-15 | 1.8e-15 |  | 2.7373e-15 | 3.8384e-13
36 | 1.17e-15 | 1.17e-15 |  | 2.1073e-15 | 2.9549e-13
36.1 | 4.78e-16 | 4.78e-16 |  | 1.4153e-15 | 1.9846e-13
36.2 | 1.64e-16 | 1.64e-16 |  | 1.1013e-15 | 1.5443e-13
36.3 | -1.61e-16 | 1e-19 |  | 7.7628e-16 | 1.0885e-13
36.4 | 2.49e-17 | 2.49e-17 |  | 9.6218e-16 | 1.3492e-13
36.5 | -3.05e-16 | 1e-19 |  | 6.3228e-16 | 8.8662e-14
36.6 | 7.1e-17 | 7.1e-17 |  | 1.0083e-15 | 1.4139e-13
36.7 | -1.32e-16 | 1e-19 |  | 8.0528e-16 | 1.1292e-13
36.8 | -1.47e-16 | 1e-19 |  | 7.9028e-16 | 1.1082e-13
36.9 | 7.2e-17 | 7.2e-17 |  | 1.0093e-15 | 1.4153e-13
37 | -3.17e-16 | 1e-19 |  | 6.2028e-16 | 8.6979e-14
37.1 | -1.71e-17 | 1e-19 |  | 9.2018e-16 | 1.2903e-13
37.2 | -7.94e-17 | 1e-19 |  | 8.5788e-16 | 1.203e-13
37.3 | -3.51e-16 | 1e-19 |  | 5.8628e-16 | 8.2212e-14
37.4 | 1.57e-16 | 1.57e-16 |  | 1.0943e-15 | 1.5345e-13
37.5 | -1.85e-16 | 1e-19 |  | 7.5228e-16 | 1.0549e-13
37.6 | -1.74e-16 | 1e-19 |  | 7.6328e-16 | 1.0703e-13
37.7 | 2.66e-16 | 2.66e-16 |  | 1.2033e-15 | 1.6873e-13
37.8 | 3e-16 | 3e-16 |  | 1.2373e-15 | 1.735e-13
37.9 | 1.16e-15 | 1.16e-15 |  | 2.0973e-15 | 2.9409e-13
38 | 6.45e-16 | 6.45e-16 |  | 1.5823e-15 | 2.2188e-13
38.1 | 4.99e-17 | 4.99e-17 |  | 9.8718e-16 | 1.3843e-13
38.2 | -1.56e-16 | 1e-19 |  | 7.8128e-16 | 1.0956e-13
38.3 | -1.5e-16 | 1e-19 |  | 7.8728e-16 | 1.104e-13
38.4 | -1.59e-16 | 1e-19 |  | 7.7828e-16 | 1.0913e-13
38.5 | -3.3e-16 | 1e-19 |  | 6.0728e-16 | 8.5156e-14
38.6 | 1.41e-16 | 1.41e-16 |  | 1.0783e-15 | 1.512e-13
38.7 | 3.54e-16 | 3.54e-16 |  | 1.2913e-15 | 1.8107e-13
38.8 | 4.04e-16 | 4.04e-16 |  | 1.3413e-15 | 1.8808e-13
38.9 | 7.57e-16 | 7.57e-16 |  | 1.6943e-15 | 2.3758e-13
39 | 7.75e-16 | 7.75e-16 |  | 1.7123e-15 | 2.4011e-13
39.1 | 6.28e-16 | 6.28e-16 |  | 1.5653e-15 | 2.1949e-13
39.2 | 2.9e-16 | 2.9e-16 |  | 1.2273e-15 | 1.721e-13
39.3 | 2.24e-16 | 2.24e-16 |  | 1.1613e-15 | 1.6284e-13
39.4 | -3.34e-16 | 1e-19 |  | 6.0328e-16 | 8.4595e-14
39.5 | -1.18e-16 | 1e-19 |  | 8.1928e-16 | 1.1488e-13
39.6 | 5.89e-16 | 5.89e-16 |  | 1.5263e-15 | 2.1402e-13
39.7 | 2.39e-15 | 2.39e-15 |  | 3.3273e-15 | 4.6657e-13
39.8 | 3.96e-15 | 3.96e-15 |  | 4.8973e-15 | 6.8672e-13
39.9 | 5.3e-15 | 5.3e-15 |  | 6.2373e-15 | 8.7463e-13
40 | 4.21e-15 | 4.21e-15 |  | 5.1473e-15 | 7.2178e-13
40.1 | 2.4e-15 | 2.4e-15 |  | 3.3373e-15 | 4.6797e-13
40.2 | 6.62e-16 | 6.62e-16 |  | 1.5993e-15 | 2.2426e-13
40.3 | 5.6e-17 | 5.6e-17 |  | 9.9328e-16 | 1.3928e-13
40.4 | -2.51e-16 | 1e-19 |  | 6.8628e-16 | 9.6234e-14
40.5 | -1.44e-16 | 1e-19 |  | 7.9328e-16 | 1.1124e-13
40.6 | -1.16e-17 | 1e-19 |  | 9.2568e-16 | 1.298e-13
40.7 | 2.05e-16 | 2.05e-16 |  | 1.1423e-15 | 1.6018e-13
40.8 | 5.77e-16 | 5.77e-16 |  | 1.5143e-15 | 2.1234e-13
40.9 | 9.02e-16 | 9.02e-16 |  | 1.8393e-15 | 2.5791e-13
41 | 1.36e-15 | 1.36e-15 | 1.36e-15 | 2.2973e-15 | 3.2214e-13
41.1 | 1.11e-15 | 1.11e-15 |  | 2.0473e-15 | 2.8708e-13
41.2 | 4.28e-16 | 4.28e-16 |  | 1.3653e-15 | 1.9145e-13
41.3 | -3.55e-16 | 1e-19 |  | 5.8228e-16 | 8.1651e-14
41.4 | -3.11e-16 | 1e-19 |  | 6.2628e-16 | 8.7821e-14
41.5 | -1.62e-17 | 1e-19 |  | 9.2108e-16 | 1.2916e-13
41.6 | 1.87e-16 | 1.87e-16 |  | 1.1243e-15 | 1.5765e-13
41.7 | -2.74e-18 | 1e-19 |  | 9.3454e-16 | 1.3105e-13
41.8 | 4.72e-16 | 4.72e-16 |  | 1.4093e-15 | 1.9762e-13
41.9 | 1.79e-16 | 1.79e-16 |  | 1.1163e-15 | 1.5653e-13
42 | 3.95e-16 | 3.95e-16 |  | 1.3323e-15 | 1.8682e-13
42.1 | 5.16e-16 | 5.16e-16 |  | 1.4533e-15 | 2.0379e-13
42.2 | 3.88e-16 | 3.88e-16 |  | 1.3253e-15 | 1.8584e-13
42.3 | -2.5e-16 | 1e-19 |  | 6.8728e-16 | 9.6374e-14
42.4 | 5.86e-17 | 5.86e-17 |  | 9.9588e-16 | 1.3965e-13
42.5 | -6.2e-17 | 1e-19 |  | 8.7528e-16 | 1.2274e-13
42.6 | -5.47e-17 | 1e-19 |  | 8.8258e-16 | 1.2376e-13
42.7 | 9.49e-17 | 9.49e-17 |  | 1.0322e-15 | 1.4474e-13
42.8 | 4.16e-16 | 4.16e-16 |  | 1.3533e-15 | 1.8976e-13
42.9 | 1.23e-15 | 1.23e-15 |  | 2.1673e-15 | 3.0391e-13
43 | 2.22e-15 | 2.22e-15 |  | 3.1573e-15 | 4.4273e-13
43.1 | 2.6e-15 | 2.6e-15 |  | 3.5373e-15 | 4.9602e-13
43.2 | 4.02e-15 | 4.02e-15 |  | 4.9573e-15 | 6.9514e-13
43.3 | 5.4e-15 | 5.4e-15 |  | 6.3373e-15 | 8.8865e-13
43.4 | 7.76e-15 | 7.76e-15 | 7.76e-15 | 8.6973e-15 | 1.2196e-12
43.5 | 1.26e-14 | 1.26e-14 |  | 1.3537e-14 | 1.8983e-12
43.6 | 2.38e-14 | 2.38e-14 |  | 2.4737e-14 | 3.4688e-12
43.7 | 7.98e-14 | 7.98e-14 |  | 8.0737e-14 | 1.1321e-11
43.8 | 1.5e-13 | 1.5e-13 |  | 1.5094e-13 | 2.1165e-11
43.9 | 2.55e-13 | 2.55e-13 |  | 2.5594e-13 | 3.5889e-11
44 | 2.21e-13 | 2.21e-13 |  | 2.2194e-13 | 3.1121e-11
44.1 | 1.2e-13 | 1.2e-13 |  | 1.2094e-13 | 1.6959e-11
44.2 | 2.3e-14 | 2.3e-14 |  | 2.3937e-14 | 3.3566e-12
44.3 | 5.9e-15 | 5.9e-15 |  | 6.8373e-15 | 9.5876e-13
44.4 | 2.18e-15 | 2.18e-15 |  | 3.1173e-15 | 4.3712e-13
44.5 | 9.47e-16 | 9.47e-16 |  | 1.8843e-15 | 2.6422e-13
44.6 | 3.27e-16 | 3.27e-16 |  | 1.2643e-15 | 1.7728e-13
44.7 | 3.88e-16 | 3.88e-16 |  | 1.3253e-15 | 1.8584e-13
44.8 | 1.2e-15 | 1.2e-15 |  | 2.1373e-15 | 2.997e-13
44.9 | 2.2e-15 | 2.2e-15 |  | 3.1373e-15 | 4.3993e-13
45 | 2.43e-15 | 2.43e-15 |  | 3.3673e-15 | 4.7218e-13
45.1 | 2.33e-15 | 2.33e-15 |  | 3.2673e-15 | 4.5816e-13
45.2 | 1.04e-15 | 1.04e-15 |  | 1.9773e-15 | 2.7727e-13
45.3 | 2.48e-16 | 2.48e-16 |  | 1.1853e-15 | 1.6621e-13
45.4 | 7.86e-17 | 7.86e-17 |  | 1.0159e-15 | 1.4245e-13
45.5 | -2.56e-16 | 1e-19 |  | 6.8128e-16 | 9.5533e-14
45.6 | -2.95e-16 | 1e-19 |  | 6.4228e-16 | 9.0064e-14
45.7 | -7.26e-17 | 1e-19 |  | 8.6468e-16 | 1.2125e-13
45.8 | -1.54e-16 | 1e-19 |  | 7.8328e-16 | 1.0984e-13
45.9 | 3.62e-16 | 3.62e-16 |  | 1.2993e-15 | 1.8219e-13
46 | 5.26e-16 | 5.26e-16 |  | 1.4633e-15 | 2.0519e-13
46.1 | 8.69e-16 | 8.69e-16 |  | 1.8063e-15 | 2.5329e-13
46.2 | 3.23e-16 | 3.23e-16 |  | 1.2603e-15 | 1.7672e-13
46.3 | -2.83e-16 | 1e-19 |  | 6.5428e-16 | 9.1747e-14
46.4 | -2.89e-16 | 1e-19 |  | 6.4828e-16 | 9.0906e-14
46.5 | -5.36e-16 | 1e-19 |  | 4.0128e-16 | 5.627e-14
46.6 | -2.23e-16 | 1e-19 |  | 7.1428e-16 | 1.0016e-13
46.7 | -2.55e-16 | 1e-19 |  | 6.8228e-16 | 9.5673e-14
46.8 | -1.89e-16 | 1e-19 |  | 7.4828e-16 | 1.0493e-13
46.9 | -3.51e-16 | 1e-19 |  | 5.8628e-16 | 8.2212e-14
47 | 1.11e-16 | 1.11e-16 |  | 1.0483e-15 | 1.47e-13
47.1 | 7.01e-17 | 7.01e-17 |  | 1.0074e-15 | 1.4126e-13
47.2 | -4.43e-16 | 1e-19 |  | 4.9428e-16 | 6.9311e-14
47.3 | -5.07e-16 | 1e-19 |  | 4.3028e-16 | 6.0336e-14
47.4 | -3.68e-16 | 1e-19 |  | 5.6928e-16 | 7.9828e-14
47.5 | -3.74e-16 | 1e-19 |  | 5.6328e-16 | 7.8986e-14
47.6 | -3.87e-16 | 1e-19 |  | 5.5028e-16 | 7.7163e-14
47.7 | -2.29e-16 | 1e-19 |  | 7.0828e-16 | 9.9319e-14
47.8 | 1.24e-17 | 1.24e-17 |  | 9.4968e-16 | 1.3317e-13
47.9 | -2.16e-17 | 1e-19 |  | 9.1568e-16 | 1.284e-13
48 | -1.85e-16 | 1e-19 |  | 7.5228e-16 | 1.0549e-13
48.1 | -3.92e-16 | 1e-19 |  | 5.4528e-16 | 7.6462e-14
48.2 | -5.63e-16 | 1e-19 |  | 3.7428e-16 | 5.2484e-14
48.3 | -1.47e-16 | 1e-19 |  | 7.9028e-16 | 1.1082e-13
48.4 | -1.32e-16 | 1e-19 |  | 8.0528e-16 | 1.1292e-13
48.5 | -1.67e-16 | 1e-19 |  | 7.7028e-16 | 1.0801e-13
48.6 | -2.09e-16 | 1e-19 |  | 7.2828e-16 | 1.0212e-13
48.7 | -3.51e-16 | 1e-19 |  | 5.8628e-16 | 8.2212e-14
48.8 | -2.43e-16 | 1e-19 |  | 6.9428e-16 | 9.7356e-14
48.9 | -2.4e-16 | 1e-19 |  | 6.9728e-16 | 9.7777e-14
49 | -1.96e-16 | 1e-19 |  | 7.4128e-16 | 1.0395e-13
49.1 | -4.18e-16 | 1e-19 |  | 5.1928e-16 | 7.2816e-14
49.2 | 3.58e-17 | 3.58e-17 |  | 9.7308e-16 | 1.3645e-13
49.3 | -2.71e-16 | 1e-19 |  | 6.6628e-16 | 9.343e-14
49.4 | -1.94e-16 | 1e-19 |  | 7.4328e-16 | 1.0423e-13
49.5 | -2.15e-16 | 1e-19 |  | 7.2228e-16 | 1.0128e-13
49.6 | -3.97e-16 | 1e-19 |  | 5.4028e-16 | 7.5761e-14
49.7 | -2.5e-16 | 1e-19 |  | 6.8728e-16 | 9.6374e-14
49.8 | -4.7e-16 | 1e-19 |  | 4.6728e-16 | 6.5525e-14
49.9 | -9.95e-17 | 1e-19 |  | 8.3778e-16 | 1.1748e-13
50 | 2.14e-16 | 2.14e-16 |  | 1.1513e-15 | 1.6144e-13
50.1 | 1.21e-16 | 1.21e-16 |  | 1.0583e-15 | 1.484e-13
50.2 | -7.51e-17 | 1e-19 |  | 8.6218e-16 | 1.209e-13
50.3 | -1.57e-16 | 1e-19 |  | 7.8028e-16 | 1.0942e-13
50.4 | -7.03e-17 | 1e-19 |  | 8.6698e-16 | 1.2157e-13
50.5 | -1.84e-16 | 1e-19 |  | 7.5328e-16 | 1.0563e-13
50.6 | -1.67e-16 | 1e-19 |  | 7.7028e-16 | 1.0801e-13
50.7 | -2.07e-16 | 1e-19 |  | 7.3028e-16 | 1.024e-13
50.8 | -4.86e-16 | 1e-19 |  | 4.5128e-16 | 6.3281e-14
50.9 | -4.03e-16 | 1e-19 |  | 5.3428e-16 | 7.492e-14
51 | -3.48e-16 | 1e-19 |  | 5.8928e-16 | 8.2632e-14
51.1 | 1.47e-16 | 1.47e-16 |  | 1.0843e-15 | 1.5204e-13
51.2 | 7.47e-17 | 7.47e-17 |  | 1.012e-15 | 1.4191e-13
51.3 | -6.25e-17 | 1e-19 |  | 8.7478e-16 | 1.2267e-13
51.4 | -3.39e-16 | 1e-19 |  | 5.9828e-16 | 8.3894e-14
51.5 | -2.35e-16 | 1e-19 |  | 7.0228e-16 | 9.8478e-14
51.6 | -2.1e-16 | 1e-19 |  | 7.2728e-16 | 1.0198e-13
51.7 | -4.05e-16 | 1e-19 |  | 5.3228e-16 | 7.4639e-14
51.8 | -7.12e-17 | 1e-19 |  | 8.6608e-16 | 1.2145e-13
51.9 | -1.63e-16 | 1e-19 |  | 7.7428e-16 | 1.0857e-13
52 | -1.45e-16 | 1e-19 |  | 7.9228e-16 | 1.111e-13
52.1 | -4e-17 | 1e-19 |  | 8.9728e-16 | 1.2582e-13
52.2 | -1.56e-16 | 1e-19 |  | 7.8128e-16 | 1.0956e-13
52.3 | -1.81e-16 | 1e-19 |  | 7.5628e-16 | 1.0605e-13
52.4 | -2.92e-16 | 1e-19 |  | 6.4528e-16 | 9.0485e-14
52.5 | -3.29e-17 | 1e-19 |  | 9.0438e-16 | 1.2682e-13
52.6 | -1.36e-16 | 1e-19 |  | 8.0128e-16 | 1.1236e-13
52.7 | -4.24e-16 | 1e-19 |  | 5.1328e-16 | 7.1975e-14
52.8 | -2.16e-16 | 1e-19 |  | 7.2128e-16 | 1.0114e-13
52.9 | -2.26e-16 | 1e-19 |  | 7.1128e-16 | 9.974e-14
53 | -3.68e-16 | 1e-19 |  | 5.6928e-16 | 7.9828e-14
53.1 | -1.82e-16 | 1e-19 |  | 7.5528e-16 | 1.0591e-13
53.2 | -2.04e-17 | 1e-19 |  | 9.1688e-16 | 1.2857e-13
53.3 | -6.47e-16 | 1e-19 |  | 2.9028e-16 | 4.0705e-14
53.4 | -2.18e-16 | 1e-19 | 1e-19 | 7.1928e-16 | 1.0086e-13
53.5 | -3.01e-16 | 1e-19 |  | 6.3628e-16 | 8.9223e-14
53.6 | -1.02e-16 | 1e-19 |  | 8.3528e-16 | 1.1713e-13
53.7 | -1.46e-16 | 1e-19 |  | 7.9128e-16 | 1.1096e-13
53.8 | -2.21e-16 | 1e-19 |  | 7.1628e-16 | 1.0044e-13
53.9 | -4.75e-16 | 1e-19 |  | 4.6228e-16 | 6.4824e-14
54 | -1.28e-16 | 1e-19 |  | 8.0928e-16 | 1.1348e-13
54.1 | -3.42e-16 | 1e-19 |  | 5.9528e-16 | 8.3474e-14
54.2 | -2.26e-16 | 1e-19 |  | 7.1128e-16 | 9.974e-14
54.3 | -3.81e-16 | 1e-19 |  | 5.5628e-16 | 7.8005e-14
54.4 | -5.72e-16 | 1e-19 |  | 3.6528e-16 | 5.1222e-14
54.5 | -4.16e-16 | 1e-19 |  | 5.2128e-16 | 7.3097e-14
54.6 | -3.31e-16 | 1e-19 |  | 6.0628e-16 | 8.5016e-14
54.7 | -3.23e-16 | 1e-19 |  | 6.1428e-16 | 8.6138e-14
54.8 | -1.89e-16 | 1e-19 |  | 7.4828e-16 | 1.0493e-13
54.9 | -3.45e-16 | 1e-19 |  | 5.9228e-16 | 8.3053e-14
55 | -5.35e-17 | 1e-19 |  | 8.8378e-16 | 1.2393e-13
55.1 | 1.03e-16 | 1.03e-16 |  | 1.0403e-15 | 1.4587e-13
55.2 | -4.6e-16 | 1e-19 |  | 4.7728e-16 | 6.6927e-14
55.3 | -1.74e-16 | 1e-19 |  | 7.6328e-16 | 1.0703e-13
55.4 | 2.23e-17 | 2.23e-17 | 2.23e-17 | 9.5958e-16 | 1.3456e-13
55.5 | -1.46e-16 | 1e-19 |  | 7.9128e-16 | 1.1096e-13
55.6 | -6.91e-16 | 1e-19 |  | 2.4628e-16 | 3.4535e-14
55.7 | -7.28e-17 | 1e-19 |  | 8.6448e-16 | 1.2122e-13
55.8 | 4.85e-17 | 4.85e-17 |  | 9.8578e-16 | 1.3823e-13
55.9 | -9.93e-17 | 1e-19 |  | 8.3798e-16 | 1.1751e-13
56 | -3.83e-16 | 1e-19 |  | 5.5428e-16 | 7.7724e-14
56.1 | -1.07e-16 | 1e-19 |  | 8.3028e-16 | 1.1643e-13
56.2 | -3.23e-16 | 1e-19 |  | 6.1428e-16 | 8.6138e-14
56.3 | -5.73e-16 | 1e-19 |  | 3.6428e-16 | 5.1081e-14
56.4 | -2.53e-16 | 1e-19 |  | 6.8428e-16 | 9.5954e-14
56.5 | -6.32e-16 | 1e-19 |  | 3.0528e-16 | 4.2808e-14
56.6 | -3.66e-16 | 1e-19 |  | 5.7128e-16 | 8.0108e-14
56.7 | -2.59e-16 | 1e-19 |  | 6.7828e-16 | 9.5112e-14
56.8 | -4.3e-16 | 1e-19 |  | 5.0728e-16 | 7.1134e-14
56.9 | -4.57e-16 | 1e-19 |  | 4.8028e-16 | 6.7348e-14
57 | -3.88e-16 | 1e-19 |  | 5.4928e-16 | 7.7023e-14
57.1 | -1.66e-16 | 1e-19 |  | 7.7128e-16 | 1.0815e-13
57.2 | -2.89e-16 | 1e-19 |  | 6.4828e-16 | 9.0906e-14
57.3 | -2.51e-16 | 1e-19 |  | 6.8628e-16 | 9.6234e-14
57.4 | -1.61e-16 | 1e-19 | 1e-19 | 7.7628e-16 | 1.0885e-13
57.5 | -2.85e-17 | 1e-19 |  | 9.0878e-16 | 1.2743e-13
57.6 | -3.37e-16 | 1e-19 |  | 6.0028e-16 | 8.4175e-14
57.7 | -4.31e-16 | 1e-19 |  | 5.0628e-16 | 7.0994e-14
57.8 | -3.75e-16 | 1e-19 |  | 5.6228e-16 | 7.8846e-14
57.9 | -4.24e-16 | 1e-19 |  | 5.1328e-16 | 7.1975e-14
58 | -9.1e-17 | 1e-19 |  | 8.4628e-16 | 1.1867e-13
58.1 | -1.35e-16 | 1e-19 |  | 8.0228e-16 | 1.125e-13
58.2 | -4.23e-17 | 1e-19 |  | 8.9498e-16 | 1.255e-13
58.3 | 7.72e-17 | 7.72e-17 |  | 1.0145e-15 | 1.4226e-13
58.4 | -2.6e-17 | 1e-19 |  | 9.1128e-16 | 1.2778e-13
58.5 | -3.12e-16 | 1e-19 |  | 6.2528e-16 | 8.768e-14
58.6 | -2.63e-16 | 1e-19 |  | 6.7428e-16 | 9.4551e-14
58.7 | -8.78e-17 | 1e-19 |  | 8.4948e-16 | 1.1912e-13
58.8 | 3.61e-17 | 3.61e-17 |  | 9.7338e-16 | 1.3649e-13
58.9 | -2.42e-16 | 1e-19 |  | 6.9528e-16 | 9.7496e-14
59 | -4.41e-17 | 1e-19 |  | 8.9318e-16 | 1.2525e-13
59.1 | -3.11e-16 | 1e-19 |  | 6.2628e-16 | 8.7821e-14
59.2 | -1.62e-16 | 1e-19 |  | 7.7528e-16 | 1.0871e-13
59.3 | -3.08e-16 | 1e-19 |  | 6.2928e-16 | 8.8241e-14
59.4 | -3.47e-16 | 1e-19 |  | 5.9028e-16 | 8.2772e-14
59.5 | -4.99e-16 | 1e-19 |  | 4.3828e-16 | 6.1458e-14
59.6 | -4.66e-16 | 1e-19 |  | 4.7128e-16 | 6.6086e-14
59.7 | -4.99e-16 | 1e-19 |  | 4.3828e-16 | 6.1458e-14
59.8 | -4.14e-17 | 1e-19 |  | 8.9588e-16 | 1.2563e-13
59.9 | -8.29e-17 | 1e-19 |  | 8.5438e-16 | 1.1981e-13
60 | -2.27e-16 | 1e-19 |  | 7.1028e-16 | 9.96e-14
60.1 | -3.64e-16 | 1e-19 |  | 5.7328e-16 | 8.0389e-14
60.2 | -5.63e-16 | 1e-19 |  | 3.7428e-16 | 5.2484e-14
60.3 | -1.78e-16 | 1e-19 |  | 7.5928e-16 | 1.0647e-13
60.4 | -1.3e-16 | 1e-19 |  | 8.0728e-16 | 1.132e-13
60.5 | -2.73e-16 | 1e-19 |  | 6.6428e-16 | 9.3149e-14
60.6 | -3.19e-16 | 1e-19 |  | 6.1828e-16 | 8.6699e-14
60.7 | 6.66e-18 | 6.66e-18 |  | 9.4394e-16 | 1.3236e-13
60.8 | -1.5e-16 | 1e-19 |  | 7.8728e-16 | 1.104e-13
60.9 | -1.9e-16 | 1e-19 |  | 7.4728e-16 | 1.0479e-13
61 | -1.92e-16 | 1e-19 |  | 7.4528e-16 | 1.0451e-13
61.1 | -2.11e-17 | 1e-19 |  | 9.1618e-16 | 1.2847e-13
61.2 | -2.06e-16 | 1e-19 |  | 7.3128e-16 | 1.0254e-13
61.3 | -1.13e-16 | 1e-19 |  | 8.2428e-16 | 1.1559e-13
61.4 | -1.78e-16 | 1e-19 |  | 7.5928e-16 | 1.0647e-13
61.5 | -2.69e-17 | 1e-19 |  | 9.1038e-16 | 1.2766e-13
61.6 | -4.5e-16 | 1e-19 |  | 4.8728e-16 | 6.8329e-14
61.7 | -3.42e-16 | 1e-19 |  | 5.9528e-16 | 8.3474e-14
61.8 | -3.29e-16 | 1e-19 |  | 6.0828e-16 | 8.5297e-14
61.9 | -2.15e-16 | 1e-19 |  | 7.2228e-16 | 1.0128e-13
62 | -6.51e-16 | 1e-19 |  | 2.8628e-16 | 4.0144e-14
62.1 | -3.67e-16 | 1e-19 |  | 5.7028e-16 | 7.9968e-14
62.2 | -2.78e-16 | 1e-19 |  | 6.5928e-16 | 9.2448e-14
62.3 | -1.9e-16 | 1e-19 |  | 7.4728e-16 | 1.0479e-13
62.4 | -3.3e-16 | 1e-19 |  | 6.0728e-16 | 8.5156e-14
62.5 | -1.61e-16 | 1e-19 |  | 7.7628e-16 | 1.0885e-13
62.6 | -6.43e-17 | 1e-19 |  | 8.7298e-16 | 1.2241e-13
62.7 | -4.06e-16 | 1e-19 |  | 5.3128e-16 | 7.4499e-14
62.8 | -3.21e-16 | 1e-19 |  | 6.1628e-16 | 8.6418e-14
62.9 | -3.06e-16 | 1e-19 |  | 6.3128e-16 | 8.8522e-14
63 | -2.17e-16 | 1e-19 |  | 7.2028e-16 | 1.01e-13
63.1 | -4.04e-16 | 1e-19 |  | 5.3328e-16 | 7.478e-14
63.2 | -1.77e-16 | 1e-19 |  | 7.6028e-16 | 1.0661e-13
63.3 | -2.57e-16 | 1e-19 |  | 6.8028e-16 | 9.5393e-14
63.4 | -4.04e-16 | 1e-19 |  | 5.3328e-16 | 7.478e-14
63.5 | -3.12e-16 | 1e-19 |  | 6.2528e-16 | 8.768e-14
63.6 | -2.26e-16 | 1e-19 |  | 7.1128e-16 | 9.974e-14
63.7 | -2.28e-16 | 1e-19 |  | 7.0928e-16 | 9.9459e-14
63.8 | -2.03e-16 | 1e-19 |  | 7.3428e-16 | 1.0296e-13
63.9 | -2.03e-16 | 1e-19 |  | 7.3428e-16 | 1.0296e-13
64 | -3.04e-16 | 1e-19 |  | 6.3328e-16 | 8.8802e-14
64.1 | -2.38e-16 | 1e-19 |  | 6.9928e-16 | 9.8057e-14
64.2 | -3.34e-16 | 1e-19 |  | 6.0328e-16 | 8.4595e-14
64.3 | -2.68e-16 | 1e-19 |  | 6.6928e-16 | 9.385e-14
64.4 | 6.62e-17 | 6.62e-17 |  | 1.0035e-15 | 1.4071e-13
64.5 | -3.72e-16 | 1e-19 |  | 5.6528e-16 | 7.9267e-14
64.6 | -5.42e-16 | 1e-19 |  | 3.9528e-16 | 5.5428e-14
64.7 | -2.16e-16 | 1e-19 |  | 7.2128e-16 | 1.0114e-13
64.8 | 3.33e-17 | 3.33e-17 |  | 9.7058e-16 | 1.361e-13
64.9 | -2.85e-16 | 1e-19 |  | 6.5228e-16 | 9.1466e-14
65 | -2.84e-16 | 1e-19 |  | 6.5328e-16 | 9.1607e-14
65.1 | -2.58e-16 | 1e-19 |  | 6.7928e-16 | 9.5253e-14
65.2 | -1.83e-16 | 1e-19 |  | 7.5428e-16 | 1.0577e-13
65.3 | -3.25e-16 | 1e-19 |  | 6.1228e-16 | 8.5857e-14
65.4 | -1.06e-16 | 1e-19 |  | 8.3128e-16 | 1.1657e-13
65.5 | -3.11e-16 | 1e-19 |  | 6.2628e-16 | 8.7821e-14
65.6 | -2.41e-16 | 1e-19 |  | 6.9628e-16 | 9.7636e-14
65.7 | -3.79e-16 | 1e-19 |  | 5.5828e-16 | 7.8285e-14
65.8 | -3.63e-16 | 1e-19 |  | 5.7428e-16 | 8.0529e-14
65.9 | -3.49e-16 | 1e-19 |  | 5.8828e-16 | 8.2492e-14
66 | -2.94e-16 | 1e-19 |  | 6.4328e-16 | 9.0204e-14
66.1 | -4.2e-16 | 1e-19 |  | 5.1728e-16 | 7.2536e-14
66.2 | -3.52e-16 | 1e-19 |  | 5.8528e-16 | 8.2071e-14
66.3 | -2.44e-16 | 1e-19 |  | 6.9328e-16 | 9.7216e-14
66.4 | -3.97e-16 | 1e-19 |  | 5.4028e-16 | 7.5761e-14
66.5 | -1.44e-16 | 1e-19 |  | 7.9328e-16 | 1.1124e-13
66.6 | -3.68e-16 | 1e-19 |  | 5.6928e-16 | 7.9828e-14
66.7 | -2.83e-16 | 1e-19 |  | 6.5428e-16 | 9.1747e-14
66.8 | -4.61e-16 | 1e-19 |  | 4.7628e-16 | 6.6787e-14
66.9 | -2.58e-16 | 1e-19 |  | 6.7928e-16 | 9.5253e-14
67 | -3.32e-16 | 1e-19 |  | 6.0528e-16 | 8.4876e-14
67.1 | -4.31e-16 | 1e-19 |  | 5.0628e-16 | 7.0994e-14
67.2 | -5.3e-16 | 1e-19 |  | 4.0728e-16 | 5.7111e-14
67.3 | -1.79e-16 | 1e-19 |  | 7.5828e-16 | 1.0633e-13
67.4 | -2.59e-16 | 1e-19 |  | 6.7828e-16 | 9.5112e-14
67.5 | -3.13e-16 | 1e-19 |  | 6.2428e-16 | 8.754e-14
67.6 | -2.41e-16 | 1e-19 |  | 6.9628e-16 | 9.7636e-14
67.7 | -7.15e-17 | 1e-19 |  | 8.6578e-16 | 1.214e-13
67.8 | -1.42e-16 | 1e-19 |  | 7.9528e-16 | 1.1152e-13
67.9 | -4.02e-16 | 1e-19 |  | 5.3528e-16 | 7.506e-14
68 | -3.72e-16 | 1e-19 |  | 5.6528e-16 | 7.9267e-14
68.1 | -3.89e-16 | 1e-19 |  | 5.4828e-16 | 7.6883e-14
68.2 | -2.54e-16 | 1e-19 |  | 6.8328e-16 | 9.5813e-14
68.3 | -2.06e-16 | 1e-19 |  | 7.3128e-16 | 1.0254e-13
68.4 | -1.93e-16 | 1e-19 |  | 7.4428e-16 | 1.0437e-13
68.5 | -1.7e-17 | 1e-19 |  | 9.2028e-16 | 1.2905e-13
68.6 | -4.5e-17 | 1e-19 |  | 8.9228e-16 | 1.2512e-13
68.7 | -5.18e-16 | 1e-19 |  | 4.1928e-16 | 5.8794e-14
68.8 | -4.61e-16 | 1e-19 |  | 4.7628e-16 | 6.6787e-14
68.9 | -3.28e-16 | 1e-19 |  | 6.0928e-16 | 8.5437e-14
69 | -2.25e-16 | 1e-19 |  | 7.1228e-16 | 9.988e-14
69.1 | -2.46e-16 | 1e-19 |  | 6.9128e-16 | 9.6935e-14
69.2 | -9.86e-17 | 1e-19 |  | 8.3868e-16 | 1.176e-13
69.3 | -5.47e-17 | 1e-19 |  | 8.8258e-16 | 1.2376e-13
69.4 | 6.43e-17 | 6.43e-17 |  | 1.0016e-15 | 1.4045e-13
69.5 | 5.4e-17 | 5.4e-17 |  | 9.9128e-16 | 1.39e-13
69.6 | -2.36e-16 | 1e-19 |  | 7.0128e-16 | 9.8338e-14
69.7 | -2.52e-16 | 1e-19 |  | 6.8528e-16 | 9.6094e-14
69.8 | -2.54e-16 | 1e-19 |  | 6.8328e-16 | 9.5813e-14
69.9 | -2.81e-16 | 1e-19 |  | 6.5628e-16 | 9.2027e-14
70 | -3.42e-16 | 1e-19 |  | 5.9528e-16 | 8.3474e-14
70.1 | -3.02e-16 | 1e-19 |  | 6.3528e-16 | 8.9083e-14
70.2 | -3.8e-16 | 1e-19 |  | 5.5728e-16 | 7.8145e-14
70.3 | -3.2e-16 | 1e-19 |  | 6.1728e-16 | 8.6559e-14
70.4 | -3.04e-16 | 1e-19 |  | 6.3328e-16 | 8.8802e-14
70.5 | -3e-16 | 1e-19 |  | 6.3728e-16 | 8.9363e-14
70.6 | -3.31e-16 | 1e-19 |  | 6.0628e-16 | 8.5016e-14
70.7 | -2.67e-16 | 1e-19 |  | 6.7028e-16 | 9.3991e-14
70.8 | -3.34e-16 | 1e-19 |  | 6.0328e-16 | 8.4595e-14
70.9 | -2.26e-16 | 1e-19 |  | 7.1128e-16 | 9.974e-14
71 | -3.45e-16 | 1e-19 |  | 5.9228e-16 | 8.3053e-14
71.1 | -3.4e-16 | 1e-19 |  | 5.9728e-16 | 8.3754e-14
71.2 | -6.14e-16 | 1e-19 |  | 3.2328e-16 | 4.5332e-14
71.3 | -2.92e-16 | 1e-19 |  | 6.4528e-16 | 9.0485e-14
71.4 | -2.53e-16 | 1e-19 |  | 6.8428e-16 | 9.5954e-14
71.5 | -1.49e-17 | 1e-19 |  | 9.2238e-16 | 1.2934e-13
71.6 | -3.22e-16 | 1e-19 |  | 6.1528e-16 | 8.6278e-14
71.7 | -5.68e-17 | 1e-19 |  | 8.8048e-16 | 1.2347e-13
71.8 | -2.57e-16 | 1e-19 |  | 6.8028e-16 | 9.5393e-14
71.9 | -3.11e-16 | 1e-19 |  | 6.2628e-16 | 8.7821e-14
72 | 9.51e-17 | 9.51e-17 |  | 1.0324e-15 | 1.4477e-13
72.1 | -1.33e-16 | 1e-19 |  | 8.0428e-16 | 1.1278e-13
72.2 | -3.47e-16 | 1e-19 |  | 5.9028e-16 | 8.2772e-14
72.3 | -4.49e-16 | 1e-19 |  | 4.8828e-16 | 6.8469e-14
72.4 | -5.68e-16 | 1e-19 |  | 3.6928e-16 | 5.1783e-14
72.5 | -3.34e-16 | 1e-19 |  | 6.0328e-16 | 8.4595e-14
72.6 | -4.39e-16 | 1e-19 |  | 4.9828e-16 | 6.9872e-14
72.7 | -3.24e-16 | 1e-19 |  | 6.1328e-16 | 8.5998e-14
72.8 | -3.6e-16 | 1e-19 |  | 5.7728e-16 | 8.095e-14
72.9 | -3.76e-16 | 1e-19 |  | 5.6128e-16 | 7.8706e-14
73 | -2.36e-16 | 1e-19 |  | 7.0128e-16 | 9.8338e-14
73.1 | -2.36e-16 | 1e-19 |  | 7.0128e-16 | 9.8338e-14
73.2 | -2.8e-16 | 1e-19 |  | 6.5728e-16 | 9.2168e-14
73.3 | -2.52e-16 | 1e-19 |  | 6.8528e-16 | 9.6094e-14
73.4 | -3.48e-16 | 1e-19 |  | 5.8928e-16 | 8.2632e-14
73.5 | -2.53e-16 | 1e-19 |  | 6.8428e-16 | 9.5954e-14
73.6 | -3.93e-16 | 1e-19 |  | 5.4428e-16 | 7.6322e-14
73.7 | -2.32e-16 | 1e-19 |  | 7.0528e-16 | 9.8898e-14
73.8 | -2.23e-16 | 1e-19 |  | 7.1428e-16 | 1.0016e-13
73.9 | -1.47e-16 | 1e-19 |  | 7.9028e-16 | 1.1082e-13
74 | -4.78e-16 | 1e-19 |  | 4.5928e-16 | 6.4403e-14
74.1 | -3.88e-16 | 1e-19 |  | 5.4928e-16 | 7.7023e-14
74.2 | -2.83e-16 | 1e-19 |  | 6.5428e-16 | 9.1747e-14
74.3 | -2.38e-16 | 1e-19 |  | 6.9928e-16 | 9.8057e-14
74.4 | -2.05e-16 | 1e-19 |  | 7.3228e-16 | 1.0268e-13
74.5 | -2.03e-16 | 1e-19 |  | 7.3428e-16 | 1.0296e-13
74.6 | -3.2e-16 | 1e-19 |  | 6.1728e-16 | 8.6559e-14
74.7 | -3.72e-16 | 1e-19 |  | 5.6528e-16 | 7.9267e-14
74.8 | -1.74e-16 | 1e-19 |  | 7.6328e-16 | 1.0703e-13
74.9 | -3.28e-16 | 1e-19 |  | 6.0928e-16 | 8.5437e-14
75 | -2.56e-16 | 1e-19 |  | 6.8128e-16 | 9.5533e-14
75.1 | -1.93e-16 | 1e-19 |  | 7.4428e-16 | 1.0437e-13
75.2 | -2.5e-16 | 1e-19 |  | 6.8728e-16 | 9.6374e-14
75.3 | -2.04e-16 | 1e-19 |  | 7.3328e-16 | 1.0282e-13
75.4 | -1.89e-16 | 1e-19 |  | 7.4828e-16 | 1.0493e-13
75.5 | -4.57e-16 | 1e-19 |  | 4.8028e-16 | 6.7348e-14
75.6 | -2.81e-16 | 1e-19 |  | 6.5628e-16 | 9.2027e-14
75.7 | -2.43e-16 | 1e-19 |  | 6.9428e-16 | 9.7356e-14
75.8 | -4.69e-16 | 1e-19 |  | 4.6828e-16 | 6.5665e-14
75.9 | -3.2e-16 | 1e-19 |  | 6.1728e-16 | 8.6559e-14
76 | -3.27e-16 | 1e-19 |  | 6.1028e-16 | 8.5577e-14
76.1 | -3.09e-16 | 1e-19 |  | 6.2828e-16 | 8.8101e-14
76.2 | -2.77e-16 | 1e-19 |  | 6.6028e-16 | 9.2588e-14
76.3 | -2.5e-16 | 1e-19 |  | 6.8728e-16 | 9.6374e-14
76.4 | -2.71e-16 | 1e-19 |  | 6.6628e-16 | 9.343e-14
76.5 | -3.18e-16 | 1e-19 |  | 6.1928e-16 | 8.6839e-14
76.6 | -2.22e-16 | 1e-19 |  | 7.1528e-16 | 1.003e-13
76.7 | -3.69e-16 | 1e-19 |  | 5.6828e-16 | 7.9688e-14
76.8 | -3.54e-16 | 1e-19 |  | 5.8328e-16 | 8.1791e-14
76.9 | -2.31e-16 | 1e-19 |  | 7.0628e-16 | 9.9039e-14
77 | -5.22e-16 | 1e-19 |  | 4.1528e-16 | 5.8233e-14
77.1 | -4.24e-16 | 1e-19 |  | 5.1328e-16 | 7.1975e-14
77.2 | -3.26e-17 | 1e-19 |  | 9.0468e-16 | 1.2686e-13
77.3 | 2.76e-18 | 2.76e-18 |  | 9.4004e-16 | 1.3182e-13
77.4 | -2.58e-16 | 1e-19 |  | 6.7928e-16 | 9.5253e-14
77.5 | -5.23e-16 | 1e-19 |  | 4.1428e-16 | 5.8093e-14
77.6 | -2.11e-16 | 1e-19 |  | 7.2628e-16 | 1.0184e-13
77.7 | -3.92e-16 | 1e-19 |  | 5.4528e-16 | 7.6462e-14
77.8 | -3.58e-16 | 1e-19 |  | 5.7928e-16 | 8.123e-14
77.9 | -3.25e-16 | 1e-19 |  | 6.1228e-16 | 8.5857e-14
78 | -2.29e-16 | 1e-19 |  | 7.0828e-16 | 9.9319e-14
78.1 | -3.78e-16 | 1e-19 |  | 5.5928e-16 | 7.8425e-14
78.2 | -1.74e-16 | 1e-19 |  | 7.6328e-16 | 1.0703e-13
78.3 | -4.66e-16 | 1e-19 |  | 4.7128e-16 | 6.6086e-14
78.4 | -2.25e-16 | 1e-19 |  | 7.1228e-16 | 9.988e-14
78.5 | -3.09e-16 | 1e-19 |  | 6.2828e-16 | 8.8101e-14
78.6 | -2.29e-16 | 1e-19 |  | 7.0828e-16 | 9.9319e-14
78.7 | -3.58e-16 | 1e-19 |  | 5.7928e-16 | 8.123e-14
78.8 | -2.73e-16 | 1e-19 |  | 6.6428e-16 | 9.3149e-14
78.9 | -2.34e-16 | 1e-19 |  | 7.0328e-16 | 9.8618e-14
79 | -2.81e-16 | 1e-19 |  | 6.5628e-16 | 9.2027e-14
79.1 | -4.05e-16 | 1e-19 |  | 5.3228e-16 | 7.4639e-14
79.2 | -4.1e-16 | 1e-19 |  | 5.2728e-16 | 7.3938e-14
79.3 | -3.32e-16 | 1e-19 |  | 6.0528e-16 | 8.4876e-14
79.4 | -3.47e-16 | 1e-19 |  | 5.9028e-16 | 8.2772e-14
79.5 | -3.54e-16 | 1e-19 |  | 5.8328e-16 | 8.1791e-14
79.6 | 1.24e-17 | 1.24e-17 |  | 9.4968e-16 | 1.3317e-13
79.7 | -6.94e-17 | 1e-19 |  | 8.6788e-16 | 1.217e-13
79.8 | -8.52e-17 | 1e-19 |  | 8.5208e-16 | 1.1948e-13
79.9 | -2.71e-16 | 1e-19 |  | 6.6628e-16 | 9.343e-14
80 | -5.92e-16 | 1e-19 |  | 3.4528e-16 | 4.8417e-14
80.1 | -3.28e-16 | 1e-19 |  | 6.0928e-16 | 8.5437e-14
80.2 | -2.34e-16 | 1e-19 |  | 7.0328e-16 | 9.8618e-14
80.3 | -3.91e-16 | 1e-19 |  | 5.4628e-16 | 7.6603e-14
80.4 | -3.25e-16 | 1e-19 |  | 6.1228e-16 | 8.5857e-14
80.5 | -5.51e-17 | 1e-19 |  | 8.8218e-16 | 1.237e-13
80.6 | -2.02e-16 | 1e-19 |  | 7.3528e-16 | 1.0311e-13
80.7 | -2.11e-16 | 1e-19 |  | 7.2628e-16 | 1.0184e-13
80.8 | -1.08e-16 | 1e-19 |  | 8.2928e-16 | 1.1629e-13
80.9 | -1.98e-16 | 1e-19 |  | 7.3928e-16 | 1.0367e-13
81 | -2.38e-16 | 1e-19 |  | 6.9928e-16 | 9.8057e-14
81.1 | 1.08e-17 | 1.08e-17 |  | 9.4808e-16 | 1.3295e-13
81.2 | -2.44e-16 | 1e-19 |  | 6.9328e-16 | 9.7216e-14
81.3 | -9.28e-16 | 1e-19 |  | 9.28e-18 | 1.3013e-15
81.4 | -4.29e-16 | 1e-19 |  | 5.0828e-16 | 7.1274e-14
81.5 | -3.93e-16 | 1e-19 |  | 5.4428e-16 | 7.6322e-14
81.6 | -2.51e-16 | 1e-19 |  | 6.8628e-16 | 9.6234e-14
81.7 | -1.07e-16 | 1e-19 |  | 8.3028e-16 | 1.1643e-13
81.8 | -9.1e-17 | 1e-19 |  | 8.4628e-16 | 1.1867e-13
81.9 | -6.48e-17 | 1e-19 |  | 8.7248e-16 | 1.2234e-13
82 | -2.85e-16 | 1e-19 |  | 6.5228e-16 | 9.1466e-14
82.1 | -4.12e-16 | 1e-19 |  | 5.2528e-16 | 7.3658e-14
82.2 | -3.26e-16 | 1e-19 |  | 6.1128e-16 | 8.5717e-14
82.3 | -2.46e-16 | 1e-19 |  | 6.9128e-16 | 9.6935e-14
82.4 | -1.34e-16 | 1e-19 |  | 8.0328e-16 | 1.1264e-13
82.5 | -6.03e-16 | 1e-19 |  | 3.3428e-16 | 4.6875e-14
82.6 | -4.94e-16 | 1e-19 |  | 4.4328e-16 | 6.2159e-14
82.7 | -1.79e-16 | 1e-19 |  | 7.5828e-16 | 1.0633e-13
82.8 | -5.65e-17 | 1e-19 |  | 8.8078e-16 | 1.2351e-13
82.9 | -1.83e-16 | 1e-19 |  | 7.5428e-16 | 1.0577e-13
83 | -3.8e-16 | 1e-19 |  | 5.5728e-16 | 7.8145e-14
83.1 | -3.15e-16 | 1e-19 |  | 6.2228e-16 | 8.726e-14
83.2 | -2.35e-16 | 1e-19 |  | 7.0228e-16 | 9.8478e-14
83.3 | -3.7e-16 | 1e-19 |  | 5.6728e-16 | 7.9547e-14
83.4 | -3.24e-16 | 1e-19 |  | 6.1328e-16 | 8.5998e-14
83.5 | -2.41e-16 | 1e-19 |  | 6.9628e-16 | 9.7636e-14
83.6 | -1.85e-16 | 1e-19 |  | 7.5228e-16 | 1.0549e-13
83.7 | -2.44e-16 | 1e-19 |  | 6.9328e-16 | 9.7216e-14
83.8 | -1.82e-16 | 1e-19 |  | 7.5528e-16 | 1.0591e-13
83.9 | -2.69e-16 | 1e-19 |  | 6.6828e-16 | 9.371e-14
84 | -3.61e-16 | 1e-19 |  | 5.7628e-16 | 8.0809e-14
84.1 | -4.86e-16 | 1e-19 |  | 4.5128e-16 | 6.3281e-14
84.2 | -2.83e-16 | 1e-19 |  | 6.5428e-16 | 9.1747e-14
84.3 | -4.9e-16 | 1e-19 |  | 4.4728e-16 | 6.272e-14
84.4 | -4.63e-16 | 1e-19 |  | 4.7428e-16 | 6.6506e-14
84.5 | -2.19e-16 | 1e-19 |  | 7.1828e-16 | 1.0072e-13
84.6 | -9.56e-17 | 1e-19 |  | 8.4168e-16 | 1.1803e-13
84.7 | -1.67e-16 | 1e-19 |  | 7.7028e-16 | 1.0801e-13
84.8 | -2.3e-19 | 1e-19 |  | 9.3705e-16 | 1.314e-13
84.9 | -1.72e-16 | 1e-19 |  | 7.6528e-16 | 1.0731e-13
85 | -1.28e-16 | 1e-19 |  | 8.0928e-16 | 1.1348e-13
85.1 | -2.68e-16 | 1e-19 |  | 6.6928e-16 | 9.385e-14
85.2 | -3.08e-16 | 1e-19 |  | 6.2928e-16 | 8.8241e-14
85.3 | -2.64e-16 | 1e-19 |  | 6.7328e-16 | 9.4411e-14
85.4 | -4.89e-16 | 1e-19 |  | 4.4828e-16 | 6.286e-14
85.5 | -3.62e-16 | 1e-19 |  | 5.7528e-16 | 8.0669e-14
85.6 | -1.88e-16 | 1e-19 |  | 7.4928e-16 | 1.0507e-13
85.7 | -2.08e-16 | 1e-19 |  | 7.2928e-16 | 1.0226e-13
85.8 | -1.59e-16 | 1e-19 |  | 7.7828e-16 | 1.0913e-13
85.9 | -3.03e-16 | 1e-19 |  | 6.3428e-16 | 8.8942e-14
86 | -2.35e-16 | 1e-19 |  | 7.0228e-16 | 9.8478e-14
86.1 | -3.19e-16 | 1e-19 |  | 6.1828e-16 | 8.6699e-14
86.2 | -2.33e-16 | 1e-19 |  | 7.0428e-16 | 9.8758e-14
86.3 | -3.12e-16 | 1e-19 |  | 6.2528e-16 | 8.768e-14
86.4 | -4.28e-16 | 1e-19 |  | 5.0928e-16 | 7.1414e-14
86.5 | -2.64e-16 | 1e-19 |  | 6.7328e-16 | 9.4411e-14
86.6 | -3.01e-16 | 1e-19 |  | 6.3628e-16 | 8.9223e-14
86.7 | -5.1e-16 | 1e-19 |  | 4.2728e-16 | 5.9916e-14
86.8 | -4.17e-16 | 1e-19 |  | 5.2028e-16 | 7.2957e-14
86.9 | -4.35e-16 | 1e-19 |  | 5.0228e-16 | 7.0433e-14
87 | -2.8e-16 | 1e-19 |  | 6.5728e-16 | 9.2168e-14
87.1 | -3.68e-16 | 1e-19 |  | 5.6928e-16 | 7.9828e-14
87.2 | -1.36e-16 | 1e-19 |  | 8.0128e-16 | 1.1236e-13
87.3 | -5.44e-16 | 1e-19 |  | 3.9328e-16 | 5.5148e-14
87.4 | -3.79e-17 | 1e-19 |  | 8.9938e-16 | 1.2612e-13
87.5 | -6.71e-17 | 1e-19 |  | 8.7018e-16 | 1.2202e-13
87.6 | -3.87e-16 | 1e-19 |  | 5.5028e-16 | 7.7163e-14
87.7 | -2.87e-16 | 1e-19 |  | 6.5028e-16 | 9.1186e-14
87.8 | -3.22e-17 | 1e-19 |  | 9.0508e-16 | 1.2692e-13
87.9 | -1.52e-16 | 1e-19 |  | 7.8528e-16 | 1.1012e-13
88 | -3.57e-16 | 1e-19 |  | 5.8028e-16 | 8.137e-14
88.1 | -4.57e-16 | 1e-19 |  | 4.8028e-16 | 6.7348e-14
88.2 | -2.73e-17 | 1e-19 |  | 9.0998e-16 | 1.276e-13
88.3 | -4.45e-16 | 1e-19 |  | 4.9228e-16 | 6.903e-14
88.4 | -3.41e-16 | 1e-19 |  | 5.9628e-16 | 8.3614e-14
88.5 | -3.63e-16 | 1e-19 |  | 5.7428e-16 | 8.0529e-14
88.6 | -2.07e-16 | 1e-19 |  | 7.3028e-16 | 1.024e-13
88.7 | 1.45e-16 | 1.45e-16 |  | 1.0823e-15 | 1.5176e-13
88.8 | -2.54e-16 | 1e-19 |  | 6.8328e-16 | 9.5813e-14
88.9 | -3.52e-16 | 1e-19 |  | 5.8528e-16 | 8.2071e-14
89 | -2.8e-16 | 1e-19 |  | 6.5728e-16 | 9.2168e-14
89.1 | -2.24e-16 | 1e-19 |  | 7.1328e-16 | 1.0002e-13
89.2 | -1.33e-16 | 1e-19 |  | 8.0428e-16 | 1.1278e-13
89.3 | -3.25e-16 | 1e-19 |  | 6.1228e-16 | 8.5857e-14
89.4 | -1.76e-16 | 1e-19 |  | 7.6128e-16 | 1.0675e-13
89.5 | -3.55e-16 | 1e-19 |  | 5.8228e-16 | 8.1651e-14
89.6 | -4.54e-16 | 1e-19 |  | 4.8328e-16 | 6.7768e-14
89.7 | -7.39e-16 | 1e-19 |  | 1.9828e-16 | 2.7804e-14
89.8 | -4.46e-16 | 1e-19 |  | 4.9128e-16 | 6.889e-14
89.9 | -4.4e-16 | 1e-19 |  | 4.9728e-16 | 6.9731e-14
90 | -3.97e-16 | 1e-19 |  | 5.4028e-16 | 7.5761e-14
90.1 | -3.91e-16 | 1e-19 |  | 5.4628e-16 | 7.6603e-14
90.2 | -3.01e-17 | 1e-19 |  | 9.0718e-16 | 1.2721e-13
90.3 | 3.01e-17 | 3.01e-17 |  | 9.6738e-16 | 1.3565e-13
90.4 | -5.74e-16 | 1e-19 |  | 3.6328e-16 | 5.0941e-14
90.5 | -5.01e-17 | 1e-19 |  | 8.8718e-16 | 1.2441e-13
90.6 | -7.7e-17 | 1e-19 |  | 8.6028e-16 | 1.2063e-13
90.7 | -2.09e-16 | 1e-19 |  | 7.2828e-16 | 1.0212e-13
90.8 | -4.41e-17 | 1e-19 |  | 8.9318e-16 | 1.2525e-13
90.9 | -3.47e-16 | 1e-19 |  | 5.9028e-16 | 8.2772e-14
91 | -1.04e-16 | 1e-19 |  | 8.3328e-16 | 1.1685e-13
91.1 | -2.75e-16 | 1e-19 |  | 6.6228e-16 | 9.2869e-14
91.2 | -4.32e-16 | 1e-19 |  | 5.0528e-16 | 7.0853e-14
91.3 | -5.33e-17 | 1e-19 |  | 8.8398e-16 | 1.2396e-13
91.4 | -3.31e-16 | 1e-19 |  | 6.0628e-16 | 8.5016e-14
91.5 | -2.76e-16 | 1e-19 |  | 6.6128e-16 | 9.2729e-14
91.6 | -3.18e-16 | 1e-19 |  | 6.1928e-16 | 8.6839e-14
91.7 | -3.6e-16 | 1e-19 |  | 5.7728e-16 | 8.095e-14
91.8 | -3.41e-16 | 1e-19 |  | 5.9628e-16 | 8.3614e-14
91.9 | -2.05e-16 | 1e-19 |  | 7.3228e-16 | 1.0268e-13
92 | -3.09e-16 | 1e-19 |  | 6.2828e-16 | 8.8101e-14
92.1 | -1.25e-16 | 1e-19 |  | 8.1228e-16 | 1.139e-13
92.2 | 5.33e-17 | 5.33e-17 |  | 9.9058e-16 | 1.389e-13
92.3 | -1.56e-16 | 1e-19 |  | 7.8128e-16 | 1.0956e-13
92.4 | -2.67e-16 | 1e-19 |  | 6.7028e-16 | 9.3991e-14
92.5 | -2.96e-16 | 1e-19 |  | 6.4128e-16 | 8.9924e-14
92.6 | -2.15e-16 | 1e-19 |  | 7.2228e-16 | 1.0128e-13
92.7 | -3.06e-16 | 1e-19 |  | 6.3128e-16 | 8.8522e-14
92.8 | -2.03e-16 | 1e-19 |  | 7.3428e-16 | 1.0296e-13
92.9 | -1.33e-16 | 1e-19 |  | 8.0428e-16 | 1.1278e-13
93 | -2.11e-16 | 1e-19 |  | 7.2628e-16 | 1.0184e-13
93.1 | -1.54e-16 | 1e-19 |  | 7.8328e-16 | 1.0984e-13
93.2 | -4.41e-16 | 1e-19 |  | 4.9628e-16 | 6.9591e-14
93.3 | -3.51e-16 | 1e-19 |  | 5.8628e-16 | 8.2212e-14
93.4 | -4.89e-16 | 1e-19 |  | 4.4828e-16 | 6.286e-14
93.5 | -1.19e-16 | 1e-19 |  | 8.1828e-16 | 1.1474e-13
93.6 | -3.63e-16 | 1e-19 |  | 5.7428e-16 | 8.0529e-14
93.7 | -8.66e-17 | 1e-19 |  | 8.5068e-16 | 1.1929e-13
93.8 | -4.77e-16 | 1e-19 |  | 4.6028e-16 | 6.4543e-14
93.9 | -4.81e-16 | 1e-19 |  | 4.5628e-16 | 6.3982e-14
94 | -1.33e-16 | 1e-19 |  | 8.0428e-16 | 1.1278e-13
94.1 | 9.88e-18 | 9.88e-18 |  | 9.4716e-16 | 1.3282e-13
94.2 | -4.57e-16 | 1e-19 |  | 4.8028e-16 | 6.7348e-14
94.3 | -7.61e-17 | 1e-19 |  | 8.6118e-16 | 1.2076e-13
94.4 | -8.5e-17 | 1e-19 |  | 8.5228e-16 | 1.1951e-13
94.5 | -3.85e-16 | 1e-19 |  | 5.5228e-16 | 7.7444e-14
94.6 | -1.39e-16 | 1e-19 |  | 7.9828e-16 | 1.1194e-13
94.7 | -6.05e-16 | 1e-19 |  | 3.3228e-16 | 4.6594e-14
94.8 | -4.04e-16 | 1e-19 |  | 5.3328e-16 | 7.478e-14
94.9 | -4.2e-17 | 1e-19 |  | 8.9528e-16 | 1.2554e-13
95 | -1.93e-16 | 1e-19 |  | 7.4428e-16 | 1.0437e-13
95.1 | -2.34e-16 | 1e-19 |  | 7.0328e-16 | 9.8618e-14
95.2 | -1.03e-16 | 1e-19 |  | 8.3428e-16 | 1.1699e-13
95.3 | -6.2e-16 | 1e-19 |  | 3.1728e-16 | 4.4491e-14
95.4 | -2.87e-16 | 1e-19 |  | 6.5028e-16 | 9.1186e-14
95.5 | -4.76e-16 | 1e-19 |  | 4.6128e-16 | 6.4683e-14
95.6 | -3.3e-16 | 1e-19 |  | 6.0728e-16 | 8.5156e-14
95.7 | -2.19e-16 | 1e-19 |  | 7.1828e-16 | 1.0072e-13
95.8 | -3.63e-16 | 1e-19 |  | 5.7428e-16 | 8.0529e-14
95.9 | -1.74e-16 | 1e-19 |  | 7.6328e-16 | 1.0703e-13
96 | -4.57e-16 | 1e-19 |  | 4.8028e-16 | 6.7348e-14
96.1 | -2.11e-16 | 1e-19 |  | 7.2628e-16 | 1.0184e-13
96.2 | -2.42e-16 | 1e-19 |  | 6.9528e-16 | 9.7496e-14
96.3 | -2.66e-16 | 1e-19 |  | 6.7128e-16 | 9.4131e-14
96.4 | -1.39e-16 | 1e-19 |  | 7.9828e-16 | 1.1194e-13
96.5 | -3.48e-16 | 1e-19 |  | 5.8928e-16 | 8.2632e-14
96.6 | -6e-16 | 1e-19 |  | 3.3728e-16 | 4.7295e-14
96.7 | -3.72e-16 | 1e-19 |  | 5.6528e-16 | 7.9267e-14
96.8 | -4.63e-16 | 1e-19 |  | 4.7428e-16 | 6.6506e-14
96.9 | -3.37e-16 | 1e-19 |  | 6.0028e-16 | 8.4175e-14
97 | -2.98e-16 | 1e-19 |  | 6.3928e-16 | 8.9644e-14
97.1 | -1.35e-16 | 1e-19 |  | 8.0228e-16 | 1.125e-13
97.2 | -1.83e-16 | 1e-19 |  | 7.5428e-16 | 1.0577e-13
97.3 | -8.75e-17 | 1e-19 |  | 8.4978e-16 | 1.1916e-13
97.4 | -1.56e-16 | 1e-19 |  | 7.8128e-16 | 1.0956e-13
97.5 | -1.96e-16 | 1e-19 |  | 7.4128e-16 | 1.0395e-13
97.6 | -3.17e-16 | 1e-19 |  | 6.2028e-16 | 8.6979e-14
97.7 | -2.71e-16 | 1e-19 |  | 6.6628e-16 | 9.343e-14
97.8 | -2.46e-17 | 1e-19 |  | 9.1268e-16 | 1.2798e-13
97.9 | 2.97e-16 | 2.97e-16 |  | 1.2343e-15 | 1.7308e-13
98 | -9.42e-17 | 1e-19 |  | 8.4308e-16 | 1.1822e-13
98.1 | -4.72e-16 | 1e-19 |  | 4.6528e-16 | 6.5244e-14
98.2 | -7.19e-17 | 1e-19 |  | 8.6538e-16 | 1.2135e-13
98.3 | -5.1e-17 | 1e-19 |  | 8.8628e-16 | 1.2428e-13
98.4 | -1.25e-16 | 1e-19 |  | 8.1228e-16 | 1.139e-13
98.5 | -1.86e-16 | 1e-19 |  | 7.5128e-16 | 1.0535e-13
98.6 | -5.44e-16 | 1e-19 |  | 3.9328e-16 | 5.5148e-14
98.7 | -2.58e-16 | 1e-19 |  | 6.7928e-16 | 9.5253e-14
98.8 | -4.38e-16 | 1e-19 |  | 4.9928e-16 | 7.0012e-14
98.9 | -1.88e-17 | 1e-19 |  | 9.1848e-16 | 1.2879e-13
99 | -2.86e-16 | 1e-19 |  | 6.5128e-16 | 9.1326e-14
99.1 | -4.82e-16 | 1e-19 |  | 4.5528e-16 | 6.3842e-14
99.2 | -3.4e-16 | 1e-19 |  | 5.9728e-16 | 8.3754e-14
99.3 | -2.87e-16 | 1e-19 |  | 6.5028e-16 | 9.1186e-14
99.4 | -1.07e-16 | 1e-19 |  | 8.3028e-16 | 1.1643e-13
99.5 | -1.87e-16 | 1e-19 |  | 7.5028e-16 | 1.0521e-13
99.6 | -1.73e-16 | 1e-19 |  | 7.6428e-16 | 1.0717e-13
99.7 | -4.59e-16 | 1e-19 |  | 4.7828e-16 | 6.7067e-14
99.8 | -3.83e-16 | 1e-19 |  | 5.5428e-16 | 7.7724e-14
99.9 | -2.77e-16 | 1e-19 |  | 6.6028e-16 | 9.2588e-14
100 | 1.36e-16 | 1.36e-16 |  | 1.0733e-15 | 1.505e-13
|  | 3.2579e-11 | 9.1425e-15
min | -9.28e-16 | 1e-19 |  | 9.28e-18

## cal
Pressure Contribution from Flag Hydrocarbons
40M Lab RGA Scan Results
Job #C110409 |  | Description: |  | Masterbond EP30-2 Epoxy Adhesive Samples on glass slides | Date: | 2009-11-04
|  |  | Oven Used: | C
AMU 41 | 1.36e-15 | amps | from RGA scan listing
AMU 43 | 7.76e-15 | amps | from RGA scan listing
AMU 53 | 1e-19 | amps | from RGA scan listing
AMU 55 | 2.23e-17 | amps | from RGA scan listing
AMU 57 | 1e-19 | amps | from RGA scan listing
Sum Flag H/C AMUs | 9.1425e-15 | amps
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
AMU 40 (w/leak open) | 2.2e-14 | amps
AMU 40 (background) | 3.3e-15 | amps
Calib leak contributes | 1.87e-14 | amps | = (w/leak open) - (background)
Leak_Rate_Cal_factor | 12620.3209 | Torr-L/s/A
Leak_Rate_Cal_factor_per_sqcm | 140.2258 | Torr-L/s/A/cm^2
Flag H/C Outgassing | 1.1538e-10 | torr l/s | = (Sum Flag H/C AMUs) x (Calib leak rate)/(Calib leak contrib.)
Test item surf area | 90 | cm2
Normalized outgassing | 1.282e-12 | torr l/s-cm2 |  | = Flag H/C Outgassing/Test item surf area
Full description: | 6 microscope glass slides 7.5 X 2.0 cm with approx. 0.010 inch thick coating of Masterbond EP30-2 Epoxy.
Epoxy was prepared with 10 parts A to 1 part B per manufactures instruction sheet and cured at room temperature for 24 Hrs. in the clean room.
The epoxy did fully harden within 24 Hrs. to a smooth translucent white surface.
CONFIRM DESCRIPTION & PRE-SCAN BAKE WITH BOB!!
Pre-scan bake: | low temperature bake (~40C) for 12 hrs. to drive off water
total for all AMUs | 3.2579e-11 | amps
approx. total outgas | 4.1116e-07 | torr-l/s
approx. total outgas | 4.5684e-09 | torr-l/s/cm^2
pumping speed | 60 | l/s
estimated total pressure | 6.8526e-09 | torr | (usually 3E-08 torr as measured by ion gauge or FC mode)