| AUTHOR: | CHECKED: | DATE | APPROVALS |
| --- | --- | --- | --- |
|  |  |  | DCN NO. | REV | DATE |
| Rodica Martin | Mike Smith | 09-8-10 |  | V1 |  |

Material and Polishing Specifications

| Material | Low absorption Fused Silica, Grade Corning 0A, Heraeus Suprasil 312 or equivalent. |
| --- | --- |
| Diameter | 2.00 in +/-0.01 in |
| Thickness | 0.50 in +/-0.01 in |
| Clear Aperture | > 80% of the optic diameter |
| Bubble and inclusion cross section within the clear aperture | Determined by the glass Grade Corning 0A, Heraeus Suprasil 312 or equivalent. No bubbles or inclusions within the clear aperture |
| Striae within the clear aperture | Grade A according to MIL-G-174 |
| Serial Number and Registration Marks | The optics should be serialized as TFP-w-56o-xx, where xx increments start at 01. Serial Numbers and Registration Marks shall be scribed or etched on the barrel of the optic. An arrow indicates Surface 1, the polarizer surface. |
| Bevel | Bevel for safety 0.03 in +/-0.01 in @ 45 deg |
| Wedge | 0.5 deg +/-0.1 deg |
| Surface quality | No scratches, sleeks and point defects of radius greater than 2 micrometers within the central 20 mm diameter 10–5 scratch–dig outside the central 20 mm diameter |
| Microroughness | <0.2 nm |

Coating Requirements:

| Surface 1 – Polarizer at 1064 nm |
| --- |
| Coating Area | 2-3 mm to bevel |
| Coating Deposition Method | Ion Beam Sputtering |
| Angle of Incidence (AOI) | 56 deg (Brewster’s angle for fused silica) |
| Suppression of s-polarization Ts/Tp | Better than 1:5000 |
| Transmission Efficiency | Tp > 99.5% for p-polarization |
| Absorption at 1064 nm | < 10 ppm |
| Transmitted Wavefront Error | 1/20th wave at 633 nm, over the clear aperture |
| Reflected Wavefront Error | 1/10th wave at 633 nm, over the clear aperture |
| Temperature Stability | The specified optical performance must be maintained over a temperature range 20 oC – 30 oC. |
| High Average Optical Power | > 300 kW/cm2 sustained |
| Surface 2 – Uncoated |

Measurement Matrix - Frequency and Method

| Specification | Test Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Material | Process Control Material Certification | 100% | Certification |
| Physical Dimensions | Measurement | 100% | Diameter, Thickness, Wedge Angle |
| Inclusions | Visual Inspection | 100% | Hand sketches indicating location, depth, and dimensions. |
| Striae | MIL-G-174, Section 4.4.5, method 1 or 2 | 100% | Certification |
| Surface Defect Analysis | Visual Inspection | 100% | Hand Sketches including defect dimensions and digital images at four points equally spaced along a circumference of a centered 10 mm diameter. |
| Microroughness | Interferometry | 100% | Certification |