LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T050030-00-W	LIGO	Feb. 23, 2005

Summary of H1 g factor measurements

R. Savage, M. Rakhmanov, K. Kawabe, J. Betzwieser

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW17-161 175 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

In February 2005, a series of measurements were made at the LIGO Hanford Observatory as part of ongoing efforts to understand heating effects in the input test masses of the H1 (4-km) interferometer.  The measurements summarized in this document recorded the temporal evolution of the arm cavity g factors resulting from heating with either the 1064 nm interferometer laser light or the 10 m light of the Thermal Compensation System (TCS).

From the g factor measurements, we infer the ITM radius of curvature by assuming the metrology value for the ETM radius of curvature.  g = g1g2; g1,2 = 1 – L/R1,2 where L is the length of the arm cavity.

The individual measurements and results are described in LHO elog entries on Feb. 9, Feb. 11, Feb. 18, and Feb. 20. Slides for the presentation on this topic given at the LIGO detector commissioning meeting on Monday, Feb. 14, 2005 can be found in the LIGO DCC.

The nominal test mass radii of curvature are from metrology performed by the fabricator.  They are:

ITMX – 14240 m

ITMY – 13600m

ETMX – 7260

ETMY - 7320

The reduced data from these measurements, which was used to generate the plots in this document, can be found at the end of this document.  Once this document is available in the DCC, one should be able to cut and paste the data from the .doc file for comparison with thermal models, etc.

Descriptions of the measurement technique can be found in:

LIGO- P030044-00-W, Characterization of the LIGO 4km Fabry-Perot Cavities Via Their High-Frequency Dynamic Responses to Length and Laser Frequency Variations, M. Rakhmanov, F. Bondu, O. Debieu, and R. L. Savage, Jr., Classical and Quantum Gravity, Vol. 21, No. 5, pp. S487-S492, 2004

# TCS annulus heating

## Xarm

This measurement began at approximately 11:56 AM PST on Feb. 9, 2005.  A nominal 1.0 W of TCS annulus heating was applied to ITMX.  A direct measurement of the TCS output power in this heating mode was not made, but C. Vorvick was confident in the calibration.

Figure 1 H1 ITMX; 1.0 W of annulus TCS heating.

## Yarm

This measurement began at approximately 5:15 PM PST on Feb. 9, 2005.  Unlike the Xarm measurement, immediately preceding this measurement the interferometer had been locked in the full “power up” mode (3 W) with the TCS compensation fully active.

Figure 2 H1 ITMY; 1.0 W of annulus TCS heating.

# TCS central heating

For these measurements, the interferometer had been out of lock for approximately six hours and the TCS system had been off.  Thus we are reasonably confident that the optics started in their “cold” state.

## Xarm

This measurement began at approximately 12:10 PM PST on Feb. 10, 2005.   The TCS heating was turned on at t = 0 and turned off at t = 25.5 minutes.

Figure 3 H1 ITMX; 37 mW of central TCS heating (nominal 0.2 W).

`

## Yarm

This measurement began at approximately 10:42 AM PST on Feb. 10, 2005.  The TCS heating was turned on at t = 0 and turned off at t = 24.0 minutes.

Figure 4 H1 ITMY; 45 mW of central TCS heating (nominal 0.2 W).

# Interferometer (1064 nm) heating

For these measurements, the full interferometer was locked at with 1.9 W of input power and the TCS heating turned off on both arms.  The ifo. was locked in this configuration for at least two hours before breaking the lock to begin the measurements.  The locks were broken at t = 0.

## Xarm

This measurement began at 12:05 PM PST on Feb. 18, 2005.  The interferometer was locked with 1.9 W of input power and no TCS heating immediately prior to this measurement.

Figure 5 H1 ITMX; after 2 hours full ifo. lock at 1.9 W input power with no TCS heating.

## Yarm

This measurement began at 2:40 PM PST on Feb. 19, 2005.  The interferometer was locked for 2.5 hours at 1.9 W of input power with no TCS heating immediately prior to this measurement.  Before that the ifo. was locked for about one hour at full power (3W) with full TCS heating active.

Figure 6 ITMXY; after 2 hours full ito. lock at 1.9 W input power with no TCS heating.

# Data for plots

format: time (minutes), g factor, radius of curvature (km)

ITMX annulus TCS heating

========================

0.00 	 0.324406 	 14.3375

2.00 	 0.323349 	 14.2176

3.00 	 0.323063 	 14.1855

4.00 	 0.322763 	 14.1520

5.00 	 0.322598 	 14.1336

6.00 	 0.322441 	 14.1162

8.00 	 0.322268 	 14.0970

9.00 	 0.322317 	 14.1024

10.00 	 0.322277 	 14.0980

11.00 	 0.322392 	 14.1108

12.00 	 0.322435 	 14.1155

13.00 	 0.322431 	 14.1151

14.00 	 0.322464 	 14.1187

15.00 	 0.322471 	 14.1195

16.00 	 0.322528 	 14.1258

17.00 	 0.322549 	 14.1282

18.00 	 0.322633 	 14.1375

19.00 	 0.322728 	 14.1481

20.00 	 0.322792 	 14.1552

21.00 	 0.322919 	 14.1694

22.00 	 0.322960 	 14.1740

23.00 	 0.322971 	 14.1752

ITMY annulus TCS heating

========================

0.00 	 0.320741 	 13.5946

2.00 	 0.319427 	 13.4621

3.00 	 0.319254 	 13.4448

4.00 	 0.319137 	 13.4332

5.00 	 0.319072 	 13.4267

6.00 	 0.319039 	 13.4235

7.00 	 0.318987 	 13.4183

9.00 	 0.319188 	 13.4383

11.00 	 0.319375 	 13.4569

13.00 	 0.319559 	 13.4753

15.00 	 0.319922 	 13.5117

17.00 	 0.320151 	 13.5348

19.00 	 0.320342 	 13.5541

21.00 	 0.320574 	 13.5776

23.00 	 0.320759 	 13.5964

28.00 	 0.321068 	 13.6280

33.00 	 0.321342 	 13.6561

40.00 	 0.321614 	 13.6841

ITMX central TCS heating

=========================

-5.00 	 0.323820 	 14.2708

-4.00 	 0.323834 	 14.2724

-3.00 	 0.323954 	 14.2860

-2.00 	 0.323848 	 14.2740

-1.00 	 0.323910 	 14.2810

0.00 	 0.323919 	 14.2820

0.50 	 0.324134 	 14.3065

1.00 	 0.324278 	 14.3229

2.00 	 0.324499 	 14.3482

3.00 	 0.324727 	 14.3744

4.00 	 0.324944 	 14.3994

5.00 	 0.325003 	 14.4062

7.00 	 0.325152 	 14.4234

8.00 	 0.325194 	 14.4283

9.00 	 0.325206 	 14.4297

10.00 	 0.325490 	 14.4627

12.00 	 0.325358 	 14.4473

15.00 	 0.325577 	 14.4728

17.00 	 0.325414 	 14.4538

19.00 	 0.325455 	 14.4586

22.00 	 0.325638 	 14.4799

25.00 	 0.325703 	 14.4875

25.50 	 0.325503 	 14.4642

26.00 	 0.325206 	 14.4297

27.00 	 0.324852 	 14.3888

28.00 	 0.324825 	 14.3856

29.00 	 0.324537 	 14.3525

30.00 	 0.324577 	 14.3571

31.00 	 0.324426 	 14.3398

32.00 	 0.324532 	 14.3520

33.00 	 0.324136 	 14.3067

34.00 	 0.324328 	 14.3286

35.00 	 0.324080 	 14.3003

37.00 	 0.324044 	 14.2962

40.00 	 0.323974 	 14.2883

45.00 	 0.323853 	 14.2745

ITMY central TCS heating

========================

-24.50 	 0.321685 	 13.6914

-22.50 	 0.321894 	 13.7131

-20.50 	 0.321984 	 13.7224

-18.50 	 0.321585 	 13.6811

-16.50 	 0.321770 	 13.7002

-14.50 	 0.321857 	 13.7092

-9.50 	 0.321915 	 13.7152

-4.50 	 0.322040 	 13.7282

-2.50 	 0.321933 	 13.7171

0.50 	 0.321993 	 13.7233

1.50 	 0.322250 	 13.7501

2.50 	 0.322553 	 13.7817

3.50 	 0.322753 	 13.8027

4.50 	 0.322848 	 13.8126

5.50 	 0.322958 	 13.8242

6.50 	 0.322899 	 13.8180

7.50 	 0.323154 	 13.8449

8.50 	 0.323149 	 13.8444

9.50 	 0.323231 	 13.8530

10.50 	 0.323109 	 13.8401

11.50 	 0.323274 	 13.8576

12.50 	 0.323356 	 13.8663

13.50 	 0.323275 	 13.8577

18.50 	 0.323631 	 13.8955

23.50 	 0.323540 	 13.8858

24.50 	 0.323586 	 13.8907

25.50 	 0.323328 	 13.8633

26.50 	 0.323147 	 13.8441

27.50 	 0.322982 	 13.8267

28.50 	 0.322790 	 13.8065

30.50 	 0.322673 	 13.7943

32.50 	 0.322613 	 13.7880

34.50 	 0.322363 	 13.7618

36.50 	 0.322422 	 13.7680

38.50 	 0.322316 	 13.7569

43.50 	 0.322266 	 13.7517

48.50 	 0.322191 	 13.7439

ITMX ifo. heating

==================

2.42 	 0.325428 	 14.4554

3.00 	 0.325253 	 14.4351

3.57 	 0.325100 	 14.4174

4.00 	 0.324952 	 14.4003

4.50 	 0.324817 	 14.3847

5.50 	 0.324662 	 14.3669

6.50 	 0.324521 	 14.3507

7.50 	 0.324405 	 14.3374

8.50 	 0.324279 	 14.3230

9.50 	 0.324161 	 14.3096

10.50 	 0.324135 	 14.3066

11.50 	 0.324085 	 14.3009

12.50 	 0.323990 	 14.2901

13.50 	 0.323956 	 14.2862

14.50 	 0.323838 	 14.2728

16.50 	 0.323760 	 14.2640

18.50 	 0.323624 	 14.2486

20.50 	 0.323593 	 14.2451

22.50 	 0.323584 	 14.2441

25.50 	 0.323476 	 14.2319

28.50 	 0.323557 	 14.2411

31.50 	 0.323504 	 14.2351

34.50 	 0.323531 	 14.2381

38.50 	 0.323439 	 14.2277

41.50 	 0.323436 	 14.2274

47.50 	 0.323415 	 14.2250

51.50 	 0.323300 	 14.2121

ITMY ifo. heating

=================

2.75 	 0.321965 	 13.7204

3.37 	 0.321900 	 13.7137

3.95 	 0.322163 	 13.7410

4.53 	 0.321775 	 13.7007

5.12 	 0.321743 	 13.6974

5.72 	 0.321936 	 13.7174

6.28 	 0.321634 	 13.6862

6.88 	 0.321582 	 13.6808

7.47 	 0.321606 	 13.6833

8.05 	 0.321724 	 13.6955

9.03 	 0.321482 	 13.6705

10.00 	 0.321256 	 13.6473

10.98 	 0.321410 	 13.6631

11.95 	 0.321219 	 13.6435

12.95 	 0.321255 	 13.6472

13.92 	 0.321096 	 13.6309

14.90 	 0.321066 	 13.6278

16.07 	 0.321102 	 13.6315

17.05 	 0.320991 	 13.6201

19.00 	 0.320958 	 13.6168

20.95 	 0.321011 	 13.6222

22.92 	 0.320988 	 13.6198

25.05 	 0.321005 	 13.6216

27.02 	 0.321007 	 13.6218

29.15 	 0.321017 	 13.6228

32.10 	 0.320923 	 13.6132

35.03 	 0.321017 	 13.6228

37.97 	 0.320994 	 13.6204

41.08 	 0.320940 	 13.6149

44.00 	 0.321029 	 13.6240

47.13 	 0.320986 	 13.6196

50.07 	 0.320975 	 13.6185

# Footnotes

http://blue.ligo-wa.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=02/09/2005&anchor_to_scroll_to=2005:02:09:18:20:59-rick

http://blue.ligo-wa.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=02/11/2005&anchor_to_scroll_to=2005:02:11:11:37:51-rick

http://blue.ligo-wa.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=02/18/2005&anchor_to_scroll_to=2005:02:18:18:27:46-rick

http://blue.ligo-wa.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=02/20/2005&anchor_to_scroll_to=2005:02:20:16:23:59-malik

LIGO- G050076-00-W in the LIGO DCC or http://apex.ligo-wa.caltech.edu/~rick/Talks/050213CommGfactor.pdf.

http://www.ligo.caltech.edu/docs/P/P030044-00.pdf