| AUTHOR(S) | DATE | Document Change Notice, Release or Approval |
| --- | --- | --- |
| Kiwamu Izumi | 23/9/2013 | see LIGO DCC record Status |

- Introduction

This document covers reflectivity (or equivalently transmissivity) measurement performed in the lab with a 95% beam splitter (E1000871-02) and a 1064 nm free space laser at two different polarizations.

- Background

In the original drawing of HAM1 optical layout (D1000313-v10) this beam splitter was labeled as ‘M14’ and placed so as to reflect a S-polarizing beam instead of P-polarizing beam. However, on the other hand, because this particular beam splitter was designed for P-polarizing beam we needed to measure the actual reflectivity of it with a S-polarizing beam at wavelength of 1064 nm.

3. Set up

Figure.1 Measurement setup

As for the laser source I used the existing Lightwave NPRO which had been setup for a reference cavity and SURF squeezer experiment at the Hanford OSB optics lab. A polarizing beam splitter (PBS), which was already setup, was used to generate both S- and P-polarizing beams as shown in figure 1.

Items used :

- 1064 nm laser source (which had been already setup)

- Polarizing cube beam splitter, 1064nm (which had been already setup)

- Half wave plate, 1064 nm (which had been already setup)

- ND01 neutral density filter

- Steering mirror, 1064 nm

- Ophir laser power meter, Vega , with filter removed

- Measurements

Below are the results of the measurements (S-pol was measured with the reflected light off of the PBS and P-pol with the transmitted light). Losses are assumed to be very small so that R+T =1.

R = 99.4 +/- 0.1 % for S-pol at 1064 nm, 45 deg

R = 94.8 +/- 0.9 % for P-pol at 1064 nm, 45 deg

- Links

- Beam splitter specification : E1000871-02-v1 https://dcc.ligo.org/DocDB/0029/E1000871/001/E1000871-v1.pdf

- HAM1 optical layout : D1000313-v10 https://dcc.ligo.org/DocDB/0009/D1000313/010/D1000313-v10_HAM1.pdf

- LHO alog 7817                                                                                            https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=7817