| AUTHOR: | DATE | CHECKED: |
| --- | --- | --- |
| P. Fritschel | 3-25-11 | Mike Smith |
|  |  | Mindy Jacobson |
|  |  | Rick Savage |

This specification is for anti-reflection coated, laser-quality, plano-plano windows. They are to be used as viewports on UHV chambers.

Requirements

Physical Configuration

| Diameter | 6.00( +/- 0.01( |
| --- | --- |
| Thickness | 0.75( +/- 0.01( |
| Wedge | < 5 arc minutes |
| Bevel | 45 deg. bevel on both edges, width approx. 0.03( or as appropriate to prevent chipping |
| Clear Aperture | > 5.20( |
| Markings | Part and serial number, as defined below, shall be marked on the barrel |

Substrate Material

Fused silica, Corning 7980 or equivalent, grade 1B or better

Part and Serial Number

The Part/Serial number shall be of the format:
E1100267-v1 TYPE YY ,  where

TYPE		is  I or II, depending on the coating type – see below
YY 		is incremental for each optic type, starting at 01

Surface Quality: Surfaces 1 and 2, within the clear aperture

Laser grade polish:10-5 scratch-dig surface quality

Transmitted wavefront error, within the clear aperture

(/10 or less, at 633 nm

Anti-Reflection Coatings

Applied to surfaces S1 and S2.

Durability per MIL-C-675C, Coating Adhesion and Durability, or current compatible standard, to be approved by LIGO.

Surface S1 will be used in an ultra-high vacuum environment — coating performance therefore applies to vacuum environment; if S1 and S2 designs are different, S1 must be indicated appropriately on the barrel.

Two types of AR coatings (each optic receives the same type on each side); specifications for each side are:

| Type I. Dual wavelength | 1064 nm: R < 0.2% for 0-15 deg. AOI, s- & p-polarization |
| --- | --- |
|  | 532 nm: R < 0.3% for 0-15 deg AOI, s- & p-polarization |
| Type II. Broadband near IR | R < 0.5% from 800-1080 nm, 0 deg AOI |

Testing and Documentation

| Specification | Test Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Physical Dimensions | Visual Inspection | 100% | Certification |
| Surface Quality | Visual Inspection | 100% | Certification |
| Transmitted Wavefront Error | Interferometry | 100% | Certification |
| AR Coatings | Spectrophotometer | Witness sample for each coating run | Spectral scans |