| Author: | Dani Atkinson / Daniel Sigg |
| --- | --- |
| Refer to: | LIGO-E1100429-v3 |
| Date: | April 29, 2011 |

# Common Mode Servo Board Test Procedure

Test Preparation

Enter Name, Date, Revision, Board Serial Number and board to be tested: common mode board (CM), mode cleaner board (MC) or an acquisition light system board (ALS).

| Test Engineer | Date | Pass |
| --- | --- | --- |
| Board | Board Serial Number |  |
| D040180 rev. E | CM or ALS or MC |  |

Required Test and Ancillary Equipment

- 1 - Common Mode Board D1003364 Tester

- 1 - Tektronix AFG 3101 Signal Generator or equivalent

- 1 - Tektronix TDS 210 Oscilloscope or equivalent

- 1 – Fluke Multimeter or equivalent

- 1 - HP 4395A Network analyzer (1Hz to 10MHz) or equivalent

- 1 - Stanford Research Systems Signal Analyzer Model SR785

- 1 - GPIB to Cat5 adapter

- 1 - Cat5 cable

- 1 – Laptop CPU using Windows operating system

- 1 – Folder containing Test File Scripts

- 2 - DC Power Supplies (Five Channels Required. Continuous Supply Voltages: +/- 24VDC, +/- 17VDC, and +5VDC)

- 1 - 17VDC Power Cable

- 1 - 24VDC Power Cable

- 1 – 5VDC Power Cable (Banana Plug to Banana Plug Cable and Jumper)

- 1 - custom cable adapting the DB9 Monitor port on the D0901781 front panel into three BNCs. ( Refer to Common Mode Board: DAQ, Number D040180 Rev E, Sheet 17 of 17 for DB9 pinout detail)

- 3 – BNC Female to Female Adapters (Barrels)

- 1 - BNC Tee Connector

- 3 - BNC Female to Double Stacking Banana Plugs

- 1 – BNC Male to Mini Grabber Test Leads Cable

- 2 – 50 ohm BNC terminations

- 4 – BNC Male to BNC Male Cables

Table of Contents

Test Preparation.............................................................................................................................	1

Required Test and Ancillary Equipment……………………………………………………….	1

Important Notes………………………………………………………………………………….	3

Tests Part 1……………………………………………………………………………………….	4

Power Board Voltage ……………………………………………………………..............	4

Power Supplies Test……………………………………………………………………….	4

Oscillations………………………………………………………………………………...	5

Adjust DC Bias……………………………………………………………………………	6

Signal Gain………………………………………………………………………………...	7

Crossbar Switches…………………………………………………………………………	9

Excitation A……………………………………………………………………………….	9

Split………………………………………………………………………………………	10

Latching………………………………………………………………………………….	10

Excitation B………………………………………………………………………………	11

Limiter……………………………………………………………………………………	11

Gain Slider C…………………………………………………………………………….	12

EPICS Readbacks………………………………………………………………………..	12

Limit Indicator…………………………………………………………………………...	13

Tests Part 2………………..…………………………………………………………………….	13

Important Notes…….…………………………………………………………………	13

Power Board Noise…….………………………………………………………………...	14

Monitor Channel Filtering…….…………………………………………………………	14

Adjustment Channel Filtering…….……………………………………………………...	15

Distortion……….……………………………………………………………………….	15

Noise Spectra…………………………………………………………………………….	15

Basic Transfer Functions…………………………………………………………………	16

Transfer Functions of Boost Gain Stages………….…………………………………….	17

Transfer Functions of DAQ Channels…………..……………………………………….	17

Tests Part 3…………………………………………………………………………………….	18

High Frequency Transfer Function………..……………………………………………18

IMPORTANT NOTES: 1. On the Common Mode Servo Tester (D1003364) front panel, all switches must be returned to default positions after each test and/or step, unless otherwise instructed. 2. The default position for most switches is UP, with the exception of switches D22, D25, D28, and D31, which are DOWN. The switch default positions are shown in Picture 1 below.

Picture 1

Front of D0901781 Common Mode Servo and D1003364 Common Mode Servo Tester in default configuration.

NOTE: Common Mode Servo ALS and MC Variants 1. Unless otherwise marked, nominal values listed are for all boards. Where the ALS and/or MC boards vary from the CM board, those values will be green for ALS or red for MC.

# Tests Part 1.

# Power Board Voltage (Low Noise Power Circuit Board Assembly D0901846)

Connect +/-17VDC and +/- 24VDC to the Common Mode Servo and +5VDC to the Common Mode Servo Tester.

Turn ON Power Supplies.

On the Low Noise Power Circuit Board Assembly, Connect the positive multimeter test lead to the following test points and Connect the negative multimeter test lead to GRD.

Record the observed voltages in the data boxes below.

Turn Off Power Supplies.

| TP1 | TP2 | TP3 | TP4 | TP5 | TP6 | TP7 | TP8 | TP9 | TP10 | TP11 | TP12 | TP13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +17V | -17V | GND | GND | +5V | -15V | +24V | GND | -24V | GND | +15V | +VREF | -VREF |

** Correct voltage indications are: TP14 ~3VDC and front panel OK light lit.

## Power Supplies

Turn OFF Power Supplies.

Connect 50 pin Control cables 1 and 2 to corresponding Control Mode Servo Tester and Common Mode Servo jacks.

Turn ON Power Supplies

.

Check current draw from the ±17V power supply is between 0.3A and 0.6A.

On the front panel of Power Supplies, Observe and Record the amperage displayed.

| Power supply | Current | Nominal |
| --- | --- | --- |
| +24V |  | 0.02 |
| -24V |  | 0.02 |
| +17V |  | .45 |
| -17V |  | .45 |

Oscillations

Connect oscilloscope and Set oscilloscope coupling to AC Coupling.

Connect oscilloscope probe to the following outputs. Ensure no oscillating wave forms are observed.

Place checkmark in corresponding box below each output.

| Outputs | OUT1 | OUT2 | SERVO | A:TST1 | A:TST2 | B:TST1 | B:TST2 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CheckBox |  |  |  |  |  |  |  |
| Outputs | D32 Input Mon | D33 Split Mon | D34 Fast Mon | D39 Slow FB Mon | D40 Output Mon |  |  |
| CheckBox |  |  |  |  |  |  |  |
| Outputs | IMON | FMON | SMON |  |  |  |  |
| CheckBox |  |  |  |  |  |  |  |

Adjust DC Bias

Set Oscilloscope coupling to DC Coupling.

Connect Input Mon (D32) and Offset Adj. (D36) to the oscilloscope.

Ground IN1 using a BNC 50 ohm termination.

Adjust DC bias (R54) for zero volts observed at Input Mon (D32) ensure D32 remains zero when D36 is removed.

Picture 2

Connect FB Mon (D39) and Offset Adj. (D37) to oscilloscope.

Adjust R137 to zero volts observed at FB Mon (D39) when D37 is removed.

Connect OUT1 to oscilloscope.

Turn ON D15 (switch down).

Adjust R54 for zero volts observed.

Return D15 to default position.

Turn ON D16 (switch down).

Adjust R54 for zero volts observed.

Turn ON D15 and D16.

Adjust R54 for zero volts observed at OUT1.

Return switches to default positions.

Record observations below.

| Zero D32 via R54. | VDC |
| --- | --- |
| Zero D39 via R137. | VDC |
| Zero OUT1 via R54 with D15 enabled. | VDC |
| Zero OUT1 via R54 with D16 enabled. | VDC |
| Zero OUT1 via R54 with both D15 and D16 enabled. | VDC |

Signal Gain

Gain slider A (Gain IN1):

Toggle switch D14 Down (IN1 position).

Connect OUT2 to the oscilloscope.

Connect Function Generator Output to Common Mode Servo IN1 jack.

Set Function Generator to frequency 100Hz, Sine wave and an Amplitude of 1 Vpp.

Inject a 100Hz / 1Vpp Sine wave signal.

Measure the voltage at 0dB (all switches in default position) and Record.

Individually, Toggle each switch down (GND) and Record observed voltage.  After each voltage observation, Return the switch to default position.

Continue to Toggle each switch, Record the observed voltage and Return each switch to default position.

Return D14 to the default position.

** Tolerance is + / - 1.059 V (+/-0.5dB).

| Binary input (Switch Setting) | Measured Vpp | Nominal Vpp |
| --- | --- | --- |
| —(0dB) |  | 1 |
| D0 (1dB) |  | 1.12 |
| D1 (2dB) |  | 1.26 |
| D2 (4dB) |  | 1.59 |
| D3 (8dB) |  | 2.51 |
| D4 (16dB) |  | 6.31 |
| D3 & D4 (24dB) |  | 15.9 |
| D5 (-32dB) |  | 0.025 |
| D5 & D3 (-24dB) |  | 0.063 |
| D5 & D4 (-16dB) |  | 0.159 |
| D5 & D3 & D4 (-8dB) |  | 0.398 |

Gain slider B (Gain IN2):

Toggle switch D13 down (D14 is in the default position IN2).

Connect OUT2 to an oscilloscope.

Set Function Generator to frequency 100Hz, Sine wave and an Amplitude of 1 Vpp.

Connect Function Generator Output to Common Mode Servo IN2 jack.

Inject a 100Hz / 1Vpp Sine wave signal into IN2.

Measure the voltage at 0dB (all switches in default position) and Record.

Toggle each switch individually Down (GND) and Record observed voltage.  Return the switch to default position.

Continue to Toggle each switch, Record the observed voltage and Return each switch to default position.

Return D13 to the default position.

** Tolerance is + / - 1.059 V (+/-0.5dB).

| Binary Input (slider gain) | Measured Vpp | Nominal Vpp |
| --- | --- | --- |
| — |  | 1 |
| D6 (1dB) |  | 1.12 |
| D7 (2dB) |  | 1.26 |
| D8 (4dB) |  | 1.59 |
| D9 (8dB) |  | 2.51 |
| D10 (16dB) |  | 6.31 |
| D9 & D10 (24dB) |  | 15.9 |
| D11 (-32dB) |  | 0.025 |
| D11 & D9 (-24dB) |  | 0.063 |
| D11 & D10 (-16dB) |  | 0.159 |
| D11 & D9 & D10 (-8dB) |  | 0.398 |

Crossbar switches

Inject a 100Hz/1Vpp Sine wave to IN1. Individually, Toggle each Crossbar switches Down. Using an oscilloscope, Record the voltage states at OUT1 and OUT2. Voltage states are either ON or OFF.

| Binary input | OUT1 | Nominal | OUT2 | Nominal |
| --- | --- | --- | --- | --- |
| Switches in Default Positions |  | On |  | Off |
| D12 (input 1 disabled) |  | Off |  | Off |
| D13 (input 2 enabled) |  | On |  | Off |
| D14 (output switch) |  | On |  | On |

Inject a 100Hz/1Vpp Sine wave to IN2. Record the voltage states at OUT1 and OUT2 while toggling the switches Down. Voltages states are either ON or OFF.

| Binary input | OUT1 | Nominal | OUT2 | Nominal |
| --- | --- | --- | --- | --- |
| Switches in Default Positions |  | Off |  | On |
| D12 (input 1 disabled) |  | Off |  | On |
| D13 (input 2 enabled) |  | On |  | On |
| D14 (output switch) |  | Off |  | Off |

Excitation A

Inject a 100Hz/1Vpp Sine wave to IN1. Measure and Record the voltage at A:TEST1 and A:TEST2 while toggling the switches Down.   ** Tolerance is +/-0.5dB.

| Binary input | A:TEST1 | Nominal Vpp | A:TEST2 | Nominal Vpp |
| --- | --- | --- | --- | --- |
| Switches in Default |  | 1.00 |  | -1.00 |

Inject a 100Hz/1Vpp Sine wave to A:EXC. Measure  and Record the voltage at A:TEST2 and OUT1 while toggling the switches Down.  ** Tolerance is +/-0.5dB. (Red = MC)

| Binary input | A:TEST2 | Nominal Vpp | OUT1 | Nominal Vpp |
| --- | --- | --- | --- | --- |
| Default |  | Off |  | Off |
| D18 (com exc enable) |  | -0.10 |  | 0.10 / 0.50 |
| D18 & D19 (com option) |  | -0.10 |  | Off |

Split

Inject a 100Hz/1Vpp Sine wave to IN1. Measure and Record the voltage at OUT1 and SERVO while toggling the switches Down.  ** Tolerance is +/-0.5dB.

| Binary input | OUT1 | Nominal Vpp | SERVO | Nominal Vpp |
| --- | --- | --- | --- | --- |
| — |  | -1.00 |  | -1.00 |
| Lift D22 (disable fast) |  | -1.00 |  | Off |
| D21 (common filter) |  | -1.00 |  | -1.00 |
| D23 (fast polarity) |  | -1.00 |  | +1.00 |
| D20 (slow polarity) |  | 1.00 |  | -1.00 |
| D24 (slow option) |  | Off |  | -1.00 |
| Lift D28 (slow comp) |  | 3.98 (phase offset) |  | -1.00 |
| D29 (slow boost) |  | 4.12 (phase offset) |  | -1.00 |
| D30 (slow filter) |  | -1.00 |  | -1.00 |
| D25 (slow bypass) |  | 1.00 |  | -1.00 |
| D27 (slow offset enable) |  | 1.00 (change offset with slow offset D38) |  | -1.00 |
| D27 and D26 (slow 5V offset) |  | 1.00 (5 V offset) |  | -1.00 |

Latching

Inject a 100Hz/1Vpp Sine wave to IN1. Toggle Down LE switch (P1/11 latch enable).

Measure and Record the voltage at SERVO.

Toggle D12 Down (IN1 1 enable) and make sure the signal at the output stays on all the time.

Return LE switch and D12 switch to default positions.

| SERVO | Vpp |
| --- | --- |
| Check |  |

Excitation B

Inject a 100Hz/1Vpp Sine wave to IN1. Measure the voltage at B:TEST1 and B:TEST2 while toggling the switches Down. Tolerance is +/-0.5dB.

| Binary input | B:TEST1 | Nominal Vpp | B:TEST2 | Nominal Vpp |
| --- | --- | --- | --- | --- |
| — |  | -1.00 |  | 1.00 |
| Lift D22 |  | Off |  | Off |
| D49 (fast/slow) and lift D22 |  | -1.00 |  | 1.00 |

Inject a 100Hz/1Vpp Sine wave to B:EXC. Measure the voltage at OUT1 and SERVO while toggling the switches Down. Tolerance is +/-0.5dB.

| Binary input | OUT1 | Nominal Vpp | SERVO | Nominal Vpp |
| --- | --- | --- | --- | --- |
| — |  | Off |  | Off |
| D47 (exc. enable) |  | Off |  | 0.10 |
| D47 & D48 (fast option) |  | Off |  | Off |
| D47 & D49 |  | 0.10 |  | Off |
| D47, D49, & D24 |  | Off |  | Off |

Limiter

Inject a 100Hz/10Vpp Sine wave to IN1. Measure the voltage at SERVO while toggling switch D31/35 UP (on the tester, red is on and green is off for this switch). The measured voltage should be within 25% of the nominal value.

| Binary input | Measured Vpp | Nominal Vpp |
| --- | --- | --- |
| - |  | 20.0 Vpp |
| D31 (fast limiter) |  | 6.6 Vpp |

Gain slider C

Inject a 100Hz/1Vpp Sine wave to IN1. Measure the voltage at SERVO while toggling the switches Down. Tolerance is +/-0.5dB.

| Binary input (slider gain) | Measured Vpp | Nominal Vpp |
| --- | --- | --- |
| — |  | 1 |
| D41 (1dB) |  | 1.12 |
| D42 (2dB) |  | 1.26 |
| D43 (4dB) |  | 1.59 |
| D44 (8dB) |  | 2.51 |
| D45 (16dB) |  | 6.31 |
| D44 & D45 (24dB) |  | 15.9 |
| D46 (-32dB) |  | 0.025 |
| D46 & D44 (-24dB) |  | 0.063 |
| D46 & D45 (-16dB) |  | 0.159 |
| D46 & D45 & D44 (-8dB) |  | 0.398 |

EPICS Readbacks

Inject a 1Hz/1Vpp Sine wave to IN1.  Observe analog outputs for a peak to peak value and Record the observed voltage.

Inject a 100Hz/1Vpp Sine wave to IN1 and Record the observed voltage.

**The voltage tolerance is 1 dB (6dB for D34) of the nominal value.

(Red =  MC) (Green =  ALS)

| EPICS readback | 1Hz | Nominal Vpp | 100Hz | Nominal Vpp |
| --- | --- | --- | --- | --- |
| D32 (input mon) |  | -1.00 |  | 0.080 |
| D33 (split mon) |  | -1.00 |  | 0.080 |
| D34 (fast mon) |  | -0.4 / 7.5 / 7.5 |  | 0.80 |
| D39 (slow FB mon) |  | 1.00 |  |  |
| D40 (output mon) |  | -1.00 |  |  |

Limit indicator

Inject a 0.1Hz/10Vpp Square wave to IN1.  Observe D35 Indicator Light (limit indicator) is ON and Record the observed voltage.  Compare with the nominal response; see Appendix A6.

| D35 Indicator Light Check |  |
| --- | --- |
| Voltage |  |

Inject a 100Hz Sine wave to IN1. Increase injected signal amplitude from 0.0V, in 0.1V steps, until D35 Indicator Light goes from high (ON) to low (OFF).  Record the observed voltage.

| Binary input | Measured [Vpp] | Nominal [Vpp] |
| --- | --- | --- |
| — |  | Approx. 6.0 Vpp |

## Tests Part 2: SR785 Signal Analyzer Tests

Important Notes: 1. Ensure all Common Mode Servo Tester switches are in the default position. 2. Closely Read and follow all On-Screen prompts.

On a Windows operating system laptop, Create and Save a file called TEST_DATA to C: drive. The path is C:\Test_DATA\.

Save Test Scripts in TEST_DATA.

Connect an SR785 Signal Analyzer to the laptop with a GPIB to Cat5 adapter.

From the DOS CMD window, Type cd..  , Enter,  Type cd..  ,Enter and Type cd TEST_DATA.

Type and Run 'setgpib.bat' and Enter the adapter's IP address (which should be labeled on the adapter).

Reset the SR785's settings with 'resetSR785.bat'. If the SR785 resets when the script is run, the SR785 is properly connected to the PC.

Power Board Noise (SR785PowerBoardNoise.bat)

One pair of probes (MiniGrabbers) are required to check the noise levels at 140Hz on the low noise power board.

In the DOS CMD window, Type SR785PowerBoardNoise.

Read and Follow the On-Screen prompts for proper test equipment configuration and procedure.

Record the collected On-Screen data in the boxes below.

** Test values must be less than the values indicated in the table below.

| TP12 | < [nV/√Hz] | TP13 | < [nV/√Hz] | TP11 | < [nV/√Hz] | TP6 | < [nV/√Hz] |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 20 |  | 30 |  | 30 |  | 30 |

# Monitor Channel Filtering (SR785MonitorTFs.bat)

In the DOS CMD window, Type SR785MonitorTFs

Read and Follow the On-Screen prompts for proper test equipment configuration and procedure.

Measure test transfer functions at 100Hz to 1Hz on IN1 to the indicated monitor channels on the tester and Record the data in the table below.

** Tolerances for Lowpass filtering are +/-1dB and +/-5deg from nominal.

| Boost # | @1Hz | Nominal | @10Hz | Nominal | @100Hz | Nominal |
| --- | --- | --- | --- | --- | --- | --- |
| Input Mon (D32) |  | -0.1dB 173deg |  | -4.1dB 129deg |  | -22dB 95deg |
| Split Mon (D33) |  | -0.1dB 173deg |  | -4.1dB 129deg |  | -22dB 95deg |
| Fast Mon (D34) (CM) |  | -8.8dB 150deg |  | 10.5dB 5deg |  | -2.5dB -79deg |
| D34 (MC/ALS) |  | 19.9dB -7deg |  | 15.9dB -51deg |  | -2.0dB -85deg |
| FB Mon (D39) |  | -0.1dB -7deg |  | -4.1dB -51deg |  | -22dB -85deg |
| Output Mon (D40) |  | -0.1dB 173deg |  | -4.1dB 129deg |  | -22dB 95deg |

Return all Common Mode Servo Tester switches to the default position.

Adjustment Channel Filtering (SR785AdjustmentTFs.bat)

Type SR785AdjustmentTFs

Test the transfer functions at 10kHz to 1Hz on the indicated adjustment channels on the tester to OUT1. Toggle Down D27 when testing D38. Verify filtering of at least -60dB at 100Hz for each channel and Record levels below in the boxes below.

Return switch D27 to default position.

| Offset Adj.(D36) |  | Offset Adj. (D37) |  | Output Adj. (D38) |  |
| --- | --- | --- | --- | --- | --- |

Distortion (SR785DistortionMeasurement.bat)

Type SR785DistortionMeasurement

Inject a 1kHz/1Vrms sine wave to IN1. Use a spectrum analyzer (SR785) to measure the harmonic components at SERVO; see Appendix A7. On the SR785, Press Marker to display THD level.  Repeat the measurement for IN2 (D13 on).  Record the measurements in the boxes below.

Return D13 to the default position.

|  | IN1 | SERVO | IN2 | SERVO |
| --- | --- | --- | --- | --- |
| Total Harmonic Distortion (THD) |  | <-70dB |  | <-70dB |

Noise Spectra (SR785NoiseMeasurements.bat)

Type resetSR785 and Allow the SR785 to reset.  Type SR785NoiseMeasurements

Terminate IN1 and IN2 using 50 ohm terminations. Measure the noise density at OUT1, OUT2 and SERVO.  Record the values at 100Hz, 1kHz, 10kHz and 100kHz in the table below. See Appendix A1 for typical examples.

| Frequency | OUT1 | < [nV/√Hz] | OUT2 | < [nV/√Hz] | SERVO | < [nV/√Hz] |
| --- | --- | --- | --- | --- | --- | --- |
| 10Hz |  |  |  |  |  |  |
| 100Hz |  | 40 |  | 30 |  | 50 |
| 1kHz |  | 30 |  | 30 |  | 40 |
| 10kHz |  | 30 |  | 30 |  | 40 |
| 100kHz |  | 30 |  | 30 |  | 40 |

Basic Transfer Functions (SR785BasicTFs.bat)

Type SR785BasicTFs

Sweep the frequency from 100kHz down to 1Hz with 100mV source amplitude and Measure the transfer function from IN1 to OUT1, from IN1 to SERVO and from IN2 to OUT2. Record the values at 10Hz, 100Hz, 1kHz, 10kHz and 100kHz in the table below.  See Appendix A2 for typical examples.

** Tolerances must be within 1dB and 5deg of nominal.  See Appendix A2 for typical examples.

| OUT1/IN1 | dB | Nom (CM MC ALS) | deg | Nom (CM MC ALS) |
| --- | --- | --- | --- | --- |
| 1Hz |  | 0.0dB 14.0dB 0.0dB |  | 180deg 180deg 180deg |
| 10Hz |  | 0.0dB 14.0dB 0.0dB |  | 180deg 180deg 179deg |
| 100Hz |  | 0.0dB 14.0dB -0.2dB |  | 180deg 177deg 169deg |
| 1kHz |  | 0.0dB 13.0dB -7.0dB |  | 180deg 153deg 117deg |
| 10kHz |  | 0.0dB -0.2dB -26.0dB |  | 175deg 102deg 94deg |
| 100kHz |  | -3.0dB -20.0dB -46.0dB |  | 130deg 86deg 85deg |

| SERVO/IN1 | dB | Nom (CM MC ALS) | deg | Nom (CM MC ALS) |
| --- | --- | --- | --- | --- |
| 1Hz |  | -28.3dB 0.0dB 0.0dB |  | -23deg -180deg 180deg |
| 10Hz |  | -1.9dB 0.0dB 0.0dB |  | -127deg -180deg 180deg |
| 100Hz |  | 0.0dB 0.0dB 0.0dB |  | -174deg -180deg 180deg |
| 1kHz |  | 0.0dB 0.0dB 0.0dB |  | -180deg -180deg 180deg |
| 10kHz |  | 0.0dB 0.1dB 0.0dB |  | 89deg -177deg 89deg |
| 100kHz |  | 0.0dB 3.0dB 0.0dB |  | 81deg -170deg 81deg |

| OUT2/IN2 | dB | Nom (CM MC ALS) | deg | Nom (CM MC ALS) |
| --- | --- | --- | --- | --- |
| 1Hz |  | 0.0dB 0.0dB 0.0dB |  | 180deg 180deg 180deg |
| 10Hz |  | 0.0dB 0.0dB 0.0dB |  | 180deg 180deg 180deg |
| 100Hz |  | 0.0dB 0.0dB 0.0dB |  | 180deg 180deg 180deg |
| 1kHz |  | 0.0dB 0.0dB 0.0dB |  | 180deg 180deg 180deg |
| 10kHz |  | 0.0dB 0.0dB 0.0dB |  | 180deg 180deg 180deg |
| 100kHz |  | 0.0dB 0.0dB 0.0dB |  | 177deg 177deg 177deg |

Transfer Functions of Boost Gain Stages (SR785BoostGainTFs.bat)

Type SR785BoostGainTFs

Note: 1. Switch D5 must be Down (low) for all measurements. 2. All other switches are in default unless prompted otherwise

It is also possible to measure these boost stages by using TP3, TP8, TP9, TP10 and TP11A. See Appendix A4 for typical examples.

** Tolerances must be within 1dB and 5deg of nominal.

| Boost # | @10Hz | Nom | @100Hz | Nom | @1kHz | Nom |
| --- | --- | --- | --- | --- | --- | --- |
| Common Comp. (D17) |  | 39.7dB -14deg |  | 31.4dB -67deg |  | 12.3dB -74deg |
| 1. (D15) |  | 26.3dB -1deg |  | 26.3dB -5deg |  | 23.4dB -42deg |
| 2. (D16) |  | 26.3dB -1deg |  | 26.3dB -5deg |  | 23.4dB -42deg |
| 3. (D15+D16) |  | 23.5dB -2deg |  | 23.1dB -17deg |  | 12.9dB -61deg |
| Lift D28 (slow comp) |  | 31.3dB -68deg |  | 11.9dB -88deg |  | -8.1dB -90deg |
| D29 (slow boost) |  | 31.4dB -67deg |  | 12.3dB -74deg |  | 0.6dB -22deg |

Transfer Functions of DAQ Channels (SR785DAQTFs.bat)

Type SR785DAQTFs

Measure the transfer function from SR785 CH1 A to D0901781 Monitor jack (DAQ channels). Sweep the frequency from 10kHz down to 1Hz at 1mV source amplitude. Record the values at 1Hz and 10kHz in the table below. See Appendix A5 for typical examples.

** Tolerances must be within 1dB and 5deg of nominal.

| Frequency | 1Hz | Nominal | 10kHz | Nominal |
| --- | --- | --- | --- | --- |
| IMON |  | 26dB, 0deg |  | 26dB, 0deg |
| FMON |  | –dB, –deg |  | 46dB, -180deg |
| SMON |  | -26dB, 0deg |  | 26dB, -12deg |

## Tests Part 3: 4395A Network/Spectrum Analyzer

Connect the 4395A in a similar fashion to the SR785, with a GPIB to Cat5 adapter.

High Frequency Transfer Function (AG4395AHighFreqTF.bat)

Type AG4395AHighFreqTF

Use a network analyzer to measure the transfer function from IN1 to SERVO. Sweep the frequency from 10MHz down to 10kHz with –20dBm source. To remove cable delays first measure the transfer function against a BNC barrel and use as a reference.  Record the displayed values at 100kHz, 300kHz and 1MHz in the table below.  Nominal values are given for CM. See Appendix A3 for typical examples.

** Tolerances are within 1dB and 5deg of nominal.

| Frequency | SERVO/IN1 [dB] | Nominal | SERVO/IN1 [deg] | Nominal |
| --- | --- | --- | --- | --- |
| 100kHz |  | 0dB |  | 170deg |
| 300kHz |  | 0dB |  | 150deg |
| 1MHz |  | 2dB |  | 75deg |