LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100848 	                 LIGO                   Oct 3, 2013

aLIGO BSC-ISI, Pre-integration Testing report,

Phase II (before cartridge install)

E1100848 – V6

Hugo Paris, Jim Warner

for the SEI team

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

1.   Phase II-a	4

1.   Hardware changes	4

1.   CPS – E1100369	4

2.   GS13 – E1100740	4

3.   L4C – E1100740	4

4.   T240 – E1100740	4

5.   Cables – E1100822	4

6.   Misc	4

2.   Electronic Inventory	4

3.   Models Modifications	4

4.   Mass distribution	5

5.   Basic functionalities just after installing the BSC-ISI on the teststand	6

1.    Pressure sensors	6

i.   Spectra	7

ii.   Actuators-cables resistance	10

iii.   Offsets CPS Unlocked vs locked	10

iv.   Offset local drive	11

v.   Offset Cartesian drive	11

vi.   Range of motion	12

vii.   Linearity test	13

6.   Transfer functions and Comparison with measurements done without SUS	15

1.   On the LVEA test stand – With SUS	15

i.   Recent ETMX data for comparison – ST1 CPS	19

ii.   Comparisons with measurements prior to SUS install 	20

7.   Conclusion Phase II-a	24

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC3 was moved from the Staging building to the VEA test stand May 2013.

This document presents results of tests (Phase II) performed on the ISI-BSC3 (ITMX) before installation in the chamber. Testing performed in May is available in version 1 of this report. Since then suspensions were installed and a new round of testing was performed, between 9/30/13 and 10/4/13. Related test results are presented here.

Tests were performed, and report was written, on the model of H1:ISI-ETMX (BSC9) – E1100849 v1

This unit was tester on the H1:ISI_TST LVEA test stand model. All results are currently, refered to, and saved in:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/TST/

Once this report is validaed, results will be posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/ITMX/

The following type of document can be found in the SVN:

- Excell spreadsheet (.xls)

- Data location

- Figures location

- Masses distribution scheme (ppt)

# Phase II-a

## Hardware changes

### CPS – E1100369

CPS have not been replaced since phase I testing in the staging building.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

T240 have not been replaced since phase I testing in the staging building.

### Cables – E1100822

The cables installed on ISI-BSC3 are reported in the table below. Further information can be found in E1100822.

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

## Models Modifications

The model was updated and recompiled since MEDM screens were modified

Test result:                                              Passed:     X              Failed:         .   W Waived:   __

## Mass distribution

Here is the payload distribution of the ISI. SUS-QUAD is installed. It will be collected prior to the cartridge install.

| Stage 1 |
| --- |
| Location | Weight (lb) |
| Corner 1 | 2 Vib Abs, 1x D0902613, 1x D0902616-4 |
| Corner 2 | 2 Vib Abs, 1x D0902613, 1x D0902616-4, 1x -1, 1x -3, 2x -5, 1x -2 |
| Corner 3 | 2 Vib Abs, 1x D0902616 -3, 2x -1 |
| Total |  |

Table 1 - Payload Stage 1

Stage 1 payload is unchanged since first installation in the staging building.

| Stage 2 - 2612lb - 5% = 2481lb |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Mass Budget - Mass type | Quantity | Weight | Unit | Weight (lb) |
| D1003136 | 2 | 50 | lb | 100 |
| D1003161-21.6 | 6 | 47.5 | lb | 285 |
| D1003161-20.6 | 10 | 45.3 |  | 453 |
| 600 lb leg Element | 1 | 600 | lb | 600 |
| type 0 | 1 | .6 | lb | .6 |
| type 1 | 3 | 1.1 | lb | 3.3 |
| type 2 | 1 | 2.2 | lb | 2.2 |
| type 3 | 1 | 4.5 | lb | 4.5 |
| type 4 | 2 | 7.9 | lb | 15.8 |
| type 5 | 3 | 15.6 | lb | 46.8 |
| type 6 | 7 | 27.2 | lb | 190.4 |
|  |  |  |  | 1701.6 |

Table 2 - Payload Stage 2

Test result:                                              Passed:    X             Failed:         .   W Waived:  __

## Basic functionalities just after installing the BSC-ISI on the teststand

### Pressure sensors

Data can be found in the SVN at:

/seismic/BSC-ISI/H1/ITMX/Data/Static_Tests/

-H1_ISI_ITMX_Pressure_Sensors_Check_Raw_20131003.mat

Script(s) can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

-Pressure_Sensor_Check_Raw_BSC_ISI.m

|  | Pressure (cts) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | -25 | 233 | -1095 |
| ST1-L4C-D | 24796 | 24691 | 24449 |
| ST1-GS13-P | 1245 | 594 | -1107 |
| ST1-GS13-D | 24791 | 25005 | 25115 |
| ST1-T240-P | 13885 | 13504 | 13628 |

Table 3 - Geophones Pressure sensors

Test result:                                              Passed:     X              Failed:         .   W Waived:   __

### Spectra

Spectra are consistent between corner. No response drop in low frequency.

Spectra can be found in the SVN at:

seismic/BSC-ISI/H1/ITMX/Data/Spectra/Undamped/

-H1_ISI_ITMX_on_TST_ASD_m_LOC_CPS_T240_L4C_GS13_2020_10_03.mat

-H1_ISI_TST_on_TST_ASD_m_L4C_GS13_Stage_Tilted_2020_10_03.mat

Script(s) used for the regular spectra test can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

- ASD_Measurements_Local_BSC_ISI.m
- Plot_ASD_Local_BSC_ISI.m

Script(s) used for the tilted spectra test can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

- ASD_Measurements_Stages_Tilted_BSC_ISI.m
- Plot_ASD_Tilted_Stage_BSC_ISI.m

Figures can be found in the SVN at :

/ligo/svncommon/SeiSVN/seismic/BSC-ISI/H1/ITMX/Data/Figures/Spectra/Undamped
-H1_ISI_ITMX_on_TST_ASD_CT_LOC_CPS_T240_L4C_GS13_2020_10_03.fig

-H1_ISI_ITMX_on_TST_ASD_m_LOC_CPS_T240_L4C_GS13_2020_10_03_3.fig

-H1_ISI_TST_on_TST_Tilted_ASD_m_LOC_ST1_L4C_2020_10_03.fig

-H1_ISI_TST_on_TST_Tilted_ASD_m_LOC_ST2_GS13_2020_10_03.fig
-H1_ISI_TST_on_TST_Tilted_ASD_CT_LOC_ST1_L4C_2020_10_03.fig

-H1_ISI_TST_on_TST_Tilted_ASD_CT_LOC_ST1_L4C_2020_10_03.pdf

Figure 1 - Spectra inboard instruments - ISI Unlocked

Figure 2 - Spectra inboard instruments - ISI Tilted -ST2

Figure 3 - Spectra inboard instruments - ISI Tilted -ST1

All spectra show a peak around 1.7Hz. With the mass on V1 actuator pocket, the ST1 L4Cs do not have this feature anymore. Same remark applies to the spectra measured in

Test result:                                              Passed:    X               Failed:         .   W Waived:   __

### Actuators-cables resistance

Will be collected 10/04/13.

Test result:                                              Passed:                 Failed:         .   W Waived:    X

### Offsets CPS Unlocked vs locked

Script(s) used for this test can be found in the SVN:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

- Offset_STD_CPS_BSC_ISI.m

Table 4 - Locked vs Unlocked Position

Test result:                                              Passed:   X              Failed:         .   W Waived:

### Offset local drive

Data for this test was accidentaly not recorded in the SVN. The test was however performed on this ISI in the LVEA earlier this year. The only difference being that that the suspensions were not installed then. Results should not be affected.

Script(s) used for this test can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

- Static_Test_Local_Basis_BSC_ISI.m

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3670 | 1493 | 1466 | -12.5 | 12.7 | -16.5 |
|  | ST1 - H2 | 1503 | 3767 | 1489 | -6.5 | 11.6 | -.02 |
|  | ST1 - H3 | 1445 | 1458 | 3581 | -.97 | 3.2 | -11.8 |
|  | ST1 - V1 | 14 | -143.7 | 68.9 | 2906 | -477 | -483 |
|  | ST1 - V2 | 72.9 | 27.9 | -134 | -450.4 | 2828 | -475 |
|  | ST1 - V3 | -145 | 75 | 20.6 | -489 | -47.5 | 2441 |

Table 5 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2144 | 333 | 319 | 1.3 | -15.9 | 6.2 |
|  | ST2 - H2 | 340 | 2174 | 336 | 5 | -.5 | -43 |
|  | ST2 - H3 | 321 | 324 | 2133 | -2 | .4 | 15.6 |
|  | ST2 - V1 | 67.4 | 107 | -163 | 2507 | 347 | -39.7 |
|  | ST2 - V2 | -170 | 62.4 | 101 | 40 | 2537 | 352 |
|  | ST2 - V3 | 97.5 | -171 | 82 | 290 | 47.5 | 2440 |

Table 6 - Static Tests – Local to Local - Stage 2

Test result:                                              Passed:   X               Failed:          .   W Waived:   __

### Offset Cartesian drive

These tests are redundant with the local test. Tests in the Cartesian basis were not performed.

Test result:                                              Passed:                  Failed:         .   W Waived:   _X_

### Range of motion

The range of motion measurement is performed by pushing on the table with the actuators, in each direction collinear to the CPS.

Data can be found in the SVN at:

/seismic/BSC-ISI/H1/ITMX/Data/Static_Tests/

-H1_ISI_ITMX_on_TST_Range_Of_Motion_20131003.mat

Script(s) can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

-Range_Motion_BSC_ISI.m

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude of the Range of Motion (counts) | Amplitude of the Range of Motion (mil) |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15767.232 | -23 | 15365.748 | 31133 | 37 |
| ST1 - H2 | -15530.388 | -458 | 15503.408 | 31034 | 37 |
| ST1 - H3 | -15304.828 | 195 | 15714.632 | 31019 | 37 |
| ST1 - V1 | -11096.702 | 1287 | 13686.218 | 24783 | 30 |
| ST1 - V2 | -13168.77 | -1100 | 10995.104 | 24164 | 29 |
| ST1 - V3 | -13302.58 | -899 | 11513.224 | 24816 | 30 |
| ST2 - H1 | -9614.1804 | -419 | 8768.3934 | 18383 | 5 |
| ST2 - H2 | -13225.366 | -3974 | 5267.3394 | 18493 | 6 |
| ST2 - H3 | -5024.0496 | 4078 | 13162.96 | 18187 | 5 |
| ST2 - V1 | -11697.22 | -939 | 9784.8366 | 21482 | 6 |
| ST2 - V2 | -8794.4858 | 2040 | 12862.9 | 21657 | 6 |
| ST2 - V3 | -10705.258 | -62 | 10601.33 | 21307 | 6 |

Table 7 - Range of motion - Actuator drive in the LVEA

| Sensors | Push in positive direction | Push in negative direction |
| --- | --- | --- |
| ST1 - H1 | 17750 | -15000 |
| ST1 - H2 | 18000 | -16000 |
| ST1 - H3 | 17500 | -15000 |
| ST1 - V1 | 15000 | -17000 |
| ST1 - V2 | 15000 | -19000 |
| ST1 - V3 | 15000 | -19000 |

Table 8 – Manual Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:    X              Failed:         .   W Waived:

### Linearity test

Data can be found in the SVN at:

/seismic/BSC-ISI/H1/ITMX/Data/Linearity_Test/

-H1_ISI_ITMX_on_TST_Linearity_test_20131003.mat

Script(s) can be found in the SVN at:
/seismic/BSC-ISI/Common/Testing_Functions_BSC_ISI/Version_1/

-Linearity_Test_Awgstream_BSC_ISI.m

Figure 4 - Linearity test – ITMY– LVEA w/Quad

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | .525 | -18 | .524 | .2 |
|  | ST1 - H2 | .531 | -450 |  | 1.3 |
|  | ST1 - H3 | .517 | -203 |  | -1.5 |
|  | ST1 - V1 | .413 | 1287 | .410 | .82 |
|  | ST1 - V2 | .402 | -1095 |  | -1.77 |
|  | ST1 - V3 | .413 | -907 |  | .95 |
| Stage 2 | ST2 - H1 | .307 | -429 | .307 | .17 |
|  | ST2 - H2 | .309 | -3969 |  | .74 |
|  | ST2 - H3 | .304 | 4055 |  | -.92 |
|  | ST2 - V1 | .359 | -940 | .359 | .20 |
|  | ST2 - V2 | .361 | 2019 |  | .65 |
|  | ST2 - V3 | .356 | -59 |  | -.85 |

Table 9 - Slope – Offset Linearity test

Test result:                                              Passed:    X              Failed:         .   W Waived:

## Transfer functions and Comparison with measurements done without SUS

### On the LVEA test stand – With SUS

A little bit  saturation occurred on the inertial sensors in the 700mHz-10Hz range.

Measurement data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/H1/ITMX/Data/Transfer_Functions/Measurements/Undamped/ -H1_ISI_TST_Data_L2L_500Hz_1000Hz_ST1_ST2_20131002-202032.mat

-H1_ISI_TST_Data_L2L_700mHz_10Hz_ST1_ST2_20131002-160148.mat

-H1_ISI_TST_Data_L2L_100mHz_700mHz_ST1_ST2_20131003-020707.mat

-H1_ISI_TST_Data_L2L_10Hz_100Hz_ST1_ST2_20131002-232958.mat

-H1_ISI_TST_Data_L2L_100Hz_500Hz_ST1_ST2_20131002-214502.mat

Concatenated TF can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/H1/ITMX/Data/Transfer_Functions/Measurements/Undamped/

- H1_ISI_TST_TF_L2L_Raw_2013_10_03.mat

Figures can be found in the SVN at:

seismic/BSC-ISI/H1/ETMX/Data/Figures/Transfer_Functions/Measurements/Undamped/

-H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST2_CPS_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST2_GS13_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST1_L4C_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST1_T240_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2013_10_03.fig

-H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2013_10_03.fig

Figure 5 - TF ST1 ACT to ST1 CPS

A few suspicious resonances can be seen above 20Hz. They can also be seen on stage 2. Similar resonances were seen on ETMX as well. Relevant ETMX transfer functions are shown in the next section.

Figure 6 - TF ST1 ACT to ST1 L4C

Figure 7 - TF ST1 ACT to ST1 T240

Figure 8 - TF ST2 ACT to ST2 CPS

Figure 9 - TF ST2 ACT to ST2 GS13

### Recent ETMX data for comparison – ST1 CPS

Figure 10 - Recent CPS TF ETMX

### Comparisons with measurements prior to SUS install

The following figures shows the comparison between the transfer functions before SUS install (05/03/2013) and after (10/03/2013).

Script(s) can be found in the SVN at:

/ligo/svncommon/SeiSVN/seismic/BSC-ISI/H1/ITMX/Scripts/Misc/

-Comparison_TF_L2L_H1_ISI_ITMX.m

Figures can be found in the SVN at:

-LHO_ISI_ITMX_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20130503_vs_20131003.fig  -LHO_ISI_ITMX_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20130503_vs_20131003.fig

-LHO_ISI_ITMX_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20130503_vs_20131003.fig  -LHO_ISI_ITMX_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20130503_vs_20131003.fig

-LHO_ISI_ITMX_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20130503_vs_20131003.fig  -LHO_ISI_ITMX_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20130503_vs_20131003.fig

-LHO_ISI_ITMX_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20130503_vs_20131003.fig  -LHO_ISI_ITMX_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20130503_vs_20131003.fig

Figure 11 - Transfer functions comparison - ST1 ACT H to ST1 CPS H

Figure 12 - Transfer functions comparison - ST1 ACT V to ST1 CPS V

Figure 13 - Transfer functions comparison - ST1 ACT H to ST1 L4C H

Figure 14 - Transfer functions comparison - ST1 L4C V to ST1 L4C V

Figure 15 - Transfer functions comparison - ST2 ACT H to ST2 CPS H

Figure 16 - Transfer functions comparison - ST2 ACT V to ST2 CPS V

Figure 17 - Transfer functions comparison – ST2 ACT H to ST2 GS13 H

Figure 18 - Transfer functions comparison – ST2 ACT V to ST2 GS13 V

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

## Conclusion Phase II-a

Tests were validated with the SEI team on the Friday October 4th telecom meeting.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

Phase IIb Tests after Cartridge Install

Basic Functionality after Cartridge Installation

    - Pressure Sensors

All pressure sensors are working.

seismic/BSC-ISI/H1/ITMX/Data/Static_Tests/

Table 10 - Geophones Pressure sensors

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Spectra

Spectra of the seismometers can be found in the svn at:

seismic/BSC-ISI/H1/ITMX/Data/Spectra/Undamped/

seismic/BSC-ISI/H1/ITMX/Data/Figures/Spectra/Undamped/

Figure 19 - St1 CPS Power Spectra

Figure 20 - St1 L4C Power Spectra

Figure 21 - St1 T240 Power Spectra

Figure 22 - St2 CPS Power Spectra

Figure 23 - St2 GS13 Power Spectra

After many trials, this unit passes.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Lock Unlock Shifts

Table 11 - Lock/Unlock Shifts.

Locker shifts are within spec.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Range of motion

Table 12 - Range of motion under actuator drive

Range of motion is within spec.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - CPS Linearity Test

Figure 24 St1 Horizontal Sensor

Figure 25 St1 Vertical

Figure 26 Stage 2 Horizontal

Figure 27 Stage 2 Vertical

CPS linearity is within spec.

Test result:                                    Passed:   X              Failed:         .   W Waived:

Table 13 - Slope – Offset Linearity test

Test result:                                    Passed:   X              Failed:         .   W Waived:

  - Transfer Functions after Cartridge Install

Transfer Functions are located in the SVN at:

/SeiSVN/seismic/BSC-ISI/H1/ITMX/Data/Transfer_Functions//Undamped/

H1_ISI_ITMX_Data_L2L_100mHz_700mHz_ST1_ST2_20131122-010237.mat

H1_ISI_ITMX_Data_L2L_700mHz_10Hz_ST1_ST2_20131121-173602.mat

H1_ISI_ITMX_Data_L2L_700mHz_10Hz_ST1_ST2_20131121-173602.mat

H1_ISI_ITMX_Data_L2L_10Hz_100Hz_ST1_ST2_20131121-215531.mat

L2L concatenated

/SeiSVN/seismic/BSC-ISI/H1/ITMX/Data/Transfer_Functions/Simulations/Undamped/

-H1_ISI_ITMX_TF_L2L_Raw_2013_11_21.mat

Figure 28 St1 Act to CPS

Figure 29 St1 Act to T240

Figure 30 St1 Act to L4C

Figure 31 St2 Act to GS13

Figure 32 St2 Act to St2 CPS

Transfer functions look healthy, except for an L4C that seems to have been poorly plugged on the in-air side of the chamber. Will need to keep an eye on this unit during operation.

Test result:                                    Passed:   X              Failed:         .   W Waived:

## Conclusion Phase II-b

Test result:                                              Passed:   X              Failed:         .   W Waived:   _