LIGO Laboratory / LIGO Scientific Collaboration

E1100867-v17	        LIGO                           24 September 2013

Advanced LIGO Arm Cavity Baffle ITM

Assembly Procedure

Stephany Foley, Lisa Austin, Mike Smith, Nichole Washington, Heidy Kelman

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

CHANGE LOG

| Date, version | Summary of Changes |
| --- | --- |
| 2012-05-03 V8 | Baffle Suspension Assembly updated with notes from LHO assembly on 12/11 and CIT assembly on 5/12. |
| 2013-03-04 V9 | Updated Box Assembly instructions for 1 hole vs. older version 2 hole design |
| 2013-03-05 V10 | Added Change Log Shown PD Reference information |
| 2013-03-19 V11 | Updates and corrections to Suspension Assembly instructions. |
| 2013-03-19 V12 | Updates and corrections to Suspension Assembly instructions. Corrected Figure numbers and references. Added cable ties and cable testing information. |
| 2013-03-29 V13 | Added assembly serial number entry locations. |
| 2013-04-01 V14 | Updated document with assembly information for acceptance review. Added cable test result requirement. |
| 2013-9-11 V15 | Changed Upper Hinge Plate and added captive screws Various other corrections to distinguish ITM build |
| 2013-09-23 V16 | Defined serial numbering Corrected errors Added visuals |
| 2013-09-24 V17 | Corrected typo |

# Introduction

This document details the assembly of the AOS SLC Arm Cavity Baffle ITM and its subassemblies:

I.	Introduction	3

A.	Clean room standards	3

B.	ITM Arm Cavity Baffle Assembly	4

Table of Figures	7

II.	ARM CAVITY BAFFLE BOX ASSEMBLY	8

A.	Arm Cavity Baffle Hinge Assembly – D1002173	8

B.	Arm Cavity Baffle Shelf Assembly – D1100391	9

C.	Arm Cavity Baffle Box Assembly –	11

D.	Photodetector Assembly (D1003013) & Lower Cable (D1003111)	18

E.	Counter Weight Assemblies	22

F.	Baffle Box Final Assembly	23

III.	ARM CAVITY BAFFLE SUSPENSION ASSEMBLY	24

A.	D1002582 - SLC BAFFLE TUBE UP ASSEMBLY	25

B.	D1001007 - SLC ACB LO TUBE ASSEMBLY	26

C.	D1002564 - SLC EDDY CURRENT DAMPING 8 DIA TUBE ASSEMBLY	28

D.	D1001005 - ARM CAVITY BAFFLE BLADE ASSEMBLY	30

E.	ACB SUSPENSION ASSEMBLY for BSC1 & BSC3 –	34

## Clean room standards

For a clean assembly all LIGO standards should be followed, as presented in the latest version of the LIGO Contamination Control Plan (E0900047). Clean room garb including UHV gloves should be worn when working with parts.

All tools that come in contact with assembly should be cleaned to class B standards.

Assembly will be done under a portable clean room. Any time a part of the assembly is not covered by the portable clean room or not being actively worked on it should be covered with appropriate clean covers. (C3 polyester or equivalent).

All parts that will be included in the final assembly must be cleaned to LIGO standards, Class A. The list of parts to be Class A-cleaned includes screws, washers, inserts, and assorted other hardware. All tooling and other parts that are not included in the final assembly, but that contact Class A parts during assembly must be cleaned to LIGO standards, Class B.

## ITM Arm Cavity Baffle Assembly

### Overview of the ITM Arm Cavity Baffle Assembly

The ITM Arm Cavity Baffle is comprised of the following sub-assemblies and fixtures.

| Drawing Number | Description | Image |
| --- | --- | --- |
| Baffle Box Assembly (D1200314–x and D1200657–y) |
| Bottom Hinge Assembly | D1002173 |  |
| Shelf Assembly | D1100391 |  |
| Extender Shelf Assembly | D1100359 |  |
| Counter Weight Assembly (Left and Right) | D1200610 and D1200665 |  |
| Suspension Assembly (D1200275) |
| Blade Assembly | X-Arm: D1001005-3 Y-Arm: D1001005-2 |  |
| Lower Tube Assembly | D1001007 |  |
| Upper Tube Assembly | D1002582 |  |
| Eddy Current Damping 8” Diameter Tube Assembly | D1002564 |  |
| Anti-Rotation Assembly | D1201228 |  |

Notes for all Subassembly:

- Assembly requires two people

- Watch for foil scratched surface

Table of Figures

Figure 1: Lower Hinge Assembly	8

Figure 2: Shelf Assembly	9

Figure 3: Baffle Box Assembly	11

Figure 4: "W" Order of Assembly	13

Figure 5: Middle Reinforcing Plate	13

Figure 6: Bottom Skin	14

Figure 7: PD Housing mating	14

Figure 8: Lifting Brace attachment	15

Figure 9: Lower Cable Routing	19

Figure 10: PD Assembly	19

Figure 11: Peek Insulator installation	19

Figure 12: PD locations for D1200657 (BSC1)	20

Figure 13: PD locations for D1200314 (BSC3)	20

Figure 14: Hat Section	21

Figure 15: Suspension Assembly	24

Figure 16: Baffle Tube Up Assembly	25

Figure 17: Lo Tube Assembly	26

Figure 18: Lo Tube magnet placement	27

Figure 19: Eddy Current Damping Tube Assembly	28

Figure 20: Copper Plate location	29

Figure 21: Spring Blade Loadin	31

Figure 22: Blade Mounting Bracket position	33

Figure 23: Suspension Rod	33

Figure 24: Suspension Rod Hardware	36

Figure 25: Suspension Rod and Tube Up Assembly	37

# ARM CAVITY BAFFLE BOX ASSEMBLY

## Arm Cavity Baffle Hinge Assembly – D1002173

Figure 1: Lower Hinge Assembly

### List of parts and tools needed:

CLASS A –

D1001622 – 1					SN:____________

1185-4EN250 – 4

CLASS B

Tool to insert Heli-Coils

### Install Heli-Coil inserts into “Arm Cavity Baffle Lower Mounting Hinge” D1001622

### D1002173 serial number – SAME AS D1001622.

## Arm Cavity Baffle Shelf Assembly – D1100391

Figure 2: Shelf Assembly

### List of parts and tools needed -

CLASS A PARTS

D1001026 – 1				SN:____________

D1001027 – 1				SN:____________

D1100340 – 1				SN:____________

D1100327 – 1				SN:____________

D1100347 – 3				SN:____________

D1100352 – 3				SN:____________

FA-605-NA – 18

CLASS B TOOLS

5/64” Hex Allen Tool

### Attach “SLC ACB Lo Captured Plate” D1100352 (Short) to center of “SLC Arm Cavity Baffle Lower Leaf” D1001027

  - Install hardware 3 places

    - FA-605-NA, Flat Head Screw, #6-32 X .312, SSTL - Silver Plated,  Qty 3

### Attach “SLC ACB Right Side Panel” D1100340 to “SLC Arm Cavity Baffle Lower Leaf” D1001027 with “SLC ACB Lo Captured Plate” D1100352 (Short)

- Install hardware 3 places

- FA-605-NA, Flat Head Screw, #6-32 X .312, SSTL - Silver Plated,  Qty 3

### Repeat Step 3 on opposite side with “SLC ACB Left Side Panel” D1100327

### Attach “SLC ACB Upper Captured Plate” D1100347 (Long) to center of “SLC Arm Cavity Baffle Upper Leaf” D1001026

- Install hardware 3 places

- FA-605-NA, Flat Head Screw, #6-32 X .312, SSTL - Silver Plated,  Qty 3

### Attach “SLC Arm Cavity Baffle Upper Leaf” D1001026 to “SLC ACB Right Side Panel” D1100340 with “SLC ACB Upper Captured Plate” D1100347 (Long)

- Install hardware 3 places

- FA-605-NA, Flat Head Screw, #6-32 X .312, SSTL - Silver Plated, Qty 3

### Repeat Step 6 on opposite side with “SLC ACB Left Side Panel” D1100327

### D1100391 serial number – SAME AS D1001026

## Arm Cavity Baffle Box Assembly –

Figure 3: Baffle Box Assembly

### List of assemblies, parts, hardware and tools needed -

CLASS A ASSEMBLIES

1 - D1002173		SN:____________

4 - D1003013		SN:____________

1 - D1100391		SN:____________

1 - D1100359 (will be assembled in this section)		SN: SAME AS D1100288

90945A740 – 16

90945A760 – 24

92200A242 – 8

92200A537 – 12

D1100288 – 1		SN:____________

D1100353 – 1		SN:____________

D1100354 – 1		SN:____________

D1100458 – 1		SN:____________

D1101234 – 2		SN:____________

N-1024-A – 8

N-2520-A - 12

CLASS A PARTS

1 - D1000974						SN:____________

1 - D1000975						SN:____________

1 - D1000976						SN:____________

2 - D1001365						SN:____________

4 - D1003025						SN:____________

1 - D1003111						SN:____________

1 - D1101427						SN:____________

1 - D1101796						SN:____________

1 - D1200296 (for D1200657 only - BSC1)		SN:____________

1 - D1200313 (for D1200314 only – BSC3)		SN:____________

Balance Weights as described on LIGO-Document E1300070

D1200835 	SN:____________

D1001826 	SN:____________

D1200779 	SN:____________

D1201230 	SN:____________

CLASS A HARDWARE

2 – 92196A114

65 – 92200A537

4 – 92200A540

2 – 92200A541

18 – 98019A330

156 – 98019A355 (interchangeable with WF-25)

2 – BU-2012-N

12 – C-2008-NA

73 – C-2012-N (interchangeable with 92200A537)

81 – C-2014-N (interchangeable with 92200A537)

4 – C-2024-N

12 – FA-2412-N (must be oxidized)

12 – FA-606-NA

18 – N-1024-A

75 – N-2520-A

156 – WF-25 (interchangeable with 98019A355)

CLASS B TOOLS

### Place “SLC Arm Cavity Baffle Skin” D1200296 (or D1200313) “W” upside down on build surface.

Figure 4: "W" Order of Assembly

NOTE: Order of Assembly:  start at the middle “V” and work outwards.

### Loosely attach “SLC Arm Cavity Baffle Middle Reinforcing Plate” D1001365 to “SLC Arm Cavity Baffle Skin”  (“W”)

#### Install hardware 6 places

#### Slotted Holes (screw, SSTL, washer, nut)

    - FA-2412-N, Flat Head Screw #10-24 X 3/4”L, SSTL - oxidized,  Qty 6

    - 98019A330, Flat #10 Washer, SSTL, Qty 6

    - N-1024-A, Hex Nut #10-24, SSTL - Silver Plated,  Qty 6

Figure 5: Middle Reinforcing Plate

### Repeat Step 3 on other side

### Loosely attach “SLC Arm Cavity Baffle Skin” (“W”) to “SLC Arm Cavity Baffle Bottom Skin” D1000975

#### Install hardware 14 places

#### Eight Slotted and six Round Holes (screw, washer, SSTL, washer, nut)

- 92200A537, Socket Head Cap Screw 1/4”-20 X 1/2”L, SSTL Qty 14

- 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 28

- N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 14

Figure 6: Bottom Skin

### Attach four “Photo Detectors Housings” D1003025 to “SLC Arm Cavity Baffle Skin” (“W”)

  - Install hardware 12 places

    - FA-606-NA, Flat Head Screw #6-32 X 3/8”L, SSTL - Silver Plated, Qty 12

Figure 7: PD Housing mating

### Lay Assembly down

### Attach “SLC Arm Cavity Baffle Center Skin” D1000976 and “Arm Cavity Baffle Center, Mirror Skin” D1101796 to “SLC Arm Cavity Baffle Bottom Skin” D1000975 and “Brace, Lifting, Arm Cavity Baffle” D1101427. Refer to D1200657 or D1200314 assembly drawing Detail B on Sheet 2 for clarification.

  - Install hardware 6 places

  - Round Holes (screw, SSTL, Brace, washer, nut)

    - FA-2412-NA, Flat Head Screw #10-24 X 3/4”L, SSTL - oxidized,  Qty 6

    - 98019A330, Flat #10 Washer, SSTL, Qty 6

    - N-1024-A, Hex Nut #10-24, SSTL - Silver Plated,  Qty 6

Figure 8: Lifting Brace attachment

### Place assembly upright.

### Loosely attach “SLC Arm Cavity Baffle Top Skin” D1000974 to “SLC Arm Cavity Baffle Skin” (“W”), “SLC Arm Cavity Baffle Center Skin” D1000976, and “Arm Cavity Baffle Center, Mirror Skin” D1101796

#### Install hardware 19 places

#### Eight Slotted and eleven Round Holes (screw, washer, SSTL, washer, nut)

- 92200A537, Socket Head Cap Screw 1/4”-20 X 1/2”L, SSTL Qty 19

- 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 38

- N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 19

### Tighten all screws.

### Attach completed “ACB Shelf Assembly” D1100391 to “Arm Cavity Baffle Box Assembly” and tighten all locations.

  - Install hardware 20 places

  - Round Holes (screw, washer, SSTL (x2), washer, nut)

    - 92200A537, Socket Head Cap Screw 1/4”-20 X 1/2”L, SSTL Qty 20 – longer screws are available, if needed.

    - 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 40

    - N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 20

### Lay Baffle Box assembly flat with front facing down.

**** Assembly and attachment of D1100359 – Adjustable Extender Shelf Assembly ****

### Loosely attach “SLC ACB Bottom Extender” D1100458 to bottom of “ACB Shelf Assembly” D1100391 into “SLC ACB Lo Captured Plate” D1100352

#### Install center hardware 2 places – Do not tighten

#### Slotted Holes (screw, washer, SSTL (x2), capture plate)

    - C-2008-NA, Socket Head Cap Screw, ¼-20 X 1/2”L, SSTL Silver Plated Qty 2

    - 98019A355, Flat ¼” Washer .63” OD, SSTL  Qty 2

#### Install side hardware 4 places – Do not tighten

#### Slotted Holes (screw, washer, SSTL (x2), capture plate)

- C-2008-NA, Socket Head Cap Screw, ¼-20 X 1/2”L, SSTL Silver Plated Qty 4

- 98019A355, Flat ¼” Washer .63” OD, SSTL  Qty 4

### Loosely attach “SLC ACB Top Extender” D1100288 to top of “ACB Shelf Assembly” D1100391 into “SLC ACB Upper Captured Plate” D1100347

#### Install center hardware  2 places – Do not tighten

#### Slotted Holes (screw, washer, SSTL (x2), capture plate)

- C-2008-NA, Socket Head Cap Screw, ¼-20 X 1/2”L, SSTL Silver Plated Qty 2

- 98019A355, Flat ¼” Washer .63” OD, SSTL  Qty 2

#### Install side hardware 4 places – Do not tighten

#### Slotted Holes (screw, washer, SSTL (x2), capture plate)

- C-2008-NA, Socket Head Cap Screw, ¼-20 X 1/2”L, SSTL Silver Plated  Qty 4

- 98019A355, Flat ¼” Washer .63” OD, SSTL  Qty 4

### Loosely attach “SLC ACB Left Side Extender” D1100353 to “SLC ACB Bottom Extender” D1100458 and “SLC ACB Top Extender” D1100288

#### Install hardware 6 places – Do not tighten

#### Round Holes (screw, washer, SSTL (x2), washer, nut)

- 92200A537, Socket Head Cap Screw 1/4”-20 X 1/2”L, SSTL Qty 6

- 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 12

- N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 6

### Repeat Step 16 on opposite side with “SLC ACB Right Side Extender” D1100354

### Attach “ACB Stiffener Side Extender” D1101234 to “SLC ACB Left Side Extender” D1100353

#### Install hardware 4 places

#### Round Holes (screw, washer, Aluminum, SSTL, washer, nut)

- 92200A242, Socket Head Cap Screw #10-24 X 1/2”L, SSTL Qty 4

- 98019A330, Flat #10 Washer, SSTL Qty 8

- N-1024-A, Hex Nut #10-24, SSTL - Silver Plated,  Qty 4

### Repeat Step 18 on opposite side with “SLC ACB Right Side Extender” D1100354

### Set alignment of Extender parts and placement as defined in assembly drawing D1200657 for Y-Arm on Sheet 5 or D1200314 for X-Arm on Sheet 4. Tighten screws enough to hold measurement.

### Tighten all screws. Verify placement measurements are accurate. Loosen and re-tighten, if needed, to meet requirement.

### Attach “ACB Hinge Assembly” D1002173 to top of “ACB Baffle Box”

#### Install hardware 2 places

#### Round Holes (screw, Hinge, SSTL, washer, nut)

    - BU-2012-N, Button Head Screw, ¼-20 X 3/4”L, SSTL Qty 2

    - 98019A355, Flat ¼” Washer, SSTL Qty 2

    - N-2520-A, Hex Nut ¼”-20, SSTL Silver Plated, Qty 2

#### Install hardware 4 places

#### Round Holes (screw, Hinge, SSTL, washer, nut)

- 92200A540, Socket Head Cap Screw, ¼-20 X 3/4”L, SSTL Qty 4

- 98019A355, Flat ¼” Washer, SSTL Qty 4

- N-2520-A, Hex Nut ¼”-20, SSTL Silver Plated, Qty 4

### Attach “SLC ACB Balancing Weight” to top of “ACB Baffle Box” as defined from selection of Suspension Assembly Spring Blade serial number. See Document E1300070 for balance weights.

- Install hardware 4 places

- Round Holes (screw, washer, Balance Weight, SSTL, washer, nut)

- C-2024-N, Socket Head Cap Screw, ¼-20 X 1.5”L, SSTL, Qty 4

- 98019A355, Flat ¼” Washer, SSTL Qty 8

- N-2520-A, Hex Nut ¼”-20, SSTL Silver Plated, Qty 4

## Photodetector Assembly (D1003013) & Lower Cable (D1003111)

### List of parts, subassemblies and tools needed –

CLASS A PARTS

D1001346-1 – 1						SN:____________

D1001363 – 2							SN:____________

D1003013 - 4 (will be assembled in this section)	Same SN as PD Housing (D1003025)

D1003024 – 4						SN:____________

D1101713 – 4						SN:____________

FA-410-A – 12

CLASS A HARDWARE

92196A114 – 2

92200A537 - 20

92200A541 – 2

98019A355 – 44

F442-003C – as required

F402-003C – as required

N-2520-A – 22

WF-25 – 44 (interchangeable with 98019A355)

### Attach “aLIGO Cable Bracket, Ham Table Assembly” D1001346-1 to top of “ACB Baffle Box”

- Install hardware 2 places

- Round Holes (screw, washer, Bracket, SSTL, washer, nut)

- 92200A541, Socket Head Cap Screw, ¼-20 X .88”L, SSTL, Qty 2

- 98019A355, Flat ¼” Washer, SSTL Qty 4

- N-2520-A, Hex Nut ¼”-20, SSTL Silver Plated, Qty 2

### Attach the two “SLC Photodetector Cable Lower Assembly” D1003111 to “aLIGO Cable Bracket, Ham Table Assembly” D1001346-1

- Install hardware 2 places

- Round Holes (screw)

- 92196A114, 4-40 X 7/8”L, SSTL Qty 2

### Route cables through slots in side panels of “ACB Baffle Box”, 4 places.

Figure 9: Lower Cable Routing

### Place “PD PCB Support” D1003028 on top of “Arm Cavity Baffle ALS PD PCB” D1003239 at the ends of the “SLC Photodetector Cable Lower Assembly” D1003111, 4 places. Carefully insert “Photodiode, ACB” D1101713 into sockets in “Arm Cavity Baffle ALS PD PCB” D1003239, 4 places. Watch for bent pins and correct orientation.

Figure 10: PD Assembly

### Insert “PD Front Peek Insulator” D1003028 into “Photo Detectors Housings” D1003025 aligning holes for attachment, 4 places.

Figure 11: Peek Insulator installation

Figure 12: PD locations for D1200657 (BSC1)

Figure 13: PD locations for D1200314 (BSC3)

### Attach ends of cable to “Photo Detectors Housings” D1003025 at locations marked on “PD Back Peek” D1003014, 4 places.

- Install hardware 12 places

- Round Holes (screw, peek, PCB, kapton, PD, peek, SSTL)

- FA-410-A, 4-40 X 5/8”L, vented, silver plated SSTL Qty 12

### Attach “ACB Side Reinforcing Hat Section” D1001363 to each side of the “SLC Arm Cavity Baffle Skin” (“W”) while routing cables using Clips F442-003C and F402-003C, as needed.

#### Install hardware 10 places each side

#### Round Holes (screw, washer, SSTL (x2), washer, nut)

- 92200A537, Socket Head Cap Screw 1/4”-20 X 1/2”L, SSTL Qty 20

- 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 40

- N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 20

Figure 14: Hat Section

### Route cables securely to “ACB Baffle Box” using cable ties F442-003C and F402-003C, as needed.

## Counter Weight Assemblies

### List of parts and tools needed –

CLASS A PARTS

D1200644 – 2						SN:____________

D1200686 – 2						SN:____________

CLASS A HARDWARE

C-2028-NA - 4

WF-25 – 4

C-2014-N – 8

98019A355 – 16

N-2520-A - 8

### Attach right side “Counter Weight Slider” D1200644 to “Counter Weight Base” D1200686 as shown in drawing D1200665.

- Install hardware 2 places

- Slotted Holes (screw, washer, Slider, Base)

- C-2028-NA, Socket Head Cap Screw, ¼-20 X 1.75”L, Ag-plated SSTL, Qty 2

- WF-25, Flat ¼” Washer, SSTL Qty 2

### Attach left side “Counter Weight Slider” D1200644 to “Counter Weight Base” D1200686 as shown in drawing D1200610.

- Install hardware 2 places

- Slotted Holes (screw, washer, Slider, Base)

- C-2028-NA, Socket Head Cap Screw, ¼-20 X 1.75”L, Ag-plated SSTL, Qty 2

- WF-25, Flat ¼” Washer, SSTL Qty 2

### Attach “Counter Weight Assembly – Right Side” D1200665 to top of “ACB Box Assembly” as shown in drawing D1200657 for Y-Arm or D1200314 for X-Arm.

- Install hardware 4 places

- Round Holes (screw, washer, Counter Weight Assembly, SSTL, washer, nut)

    - C-2014-N, Socket Head Cap Screw 1/4”-20 X 7/8”L, SSTL Qty 4

    - 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 8

    - N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 4

### Attach “Counter Weight Assembly – Left Side” D1200610 to top of “ACB Box Assembly” as shown in drawing D1200657 for Y-Arm or D1200314 for X-Arm.

- Install hardware 4 places

- Round Holes (screw, washer, Counter Weight Assembly, SSTL, washer, nut)

- C-2014-N, Socket Head Cap Screw 1/4”-20 X 7/8”L, SSTL Qty 4

- 98019A355, Flat ¼” Washer, .63” OD, SSTL Qty 8

- N-2520-A, Hex Nut  ¼”-20, SSTL - Silver Plated,  Qty 4

## Baffle Box Final Assembly

### Test cables per Procedure T1100637.		PASS		FAIL	(circle one)

Comments: ________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

### Torque all screws, as required.

### Record serial number of applicable Box Assembly:

D1200657 (BSC1)		SN: SAME AS D1000974

D1200314 (BSC3)		SN: SAME AS D1000974

### Weigh assembly and record: _______________ lbs. (kg/grams)

# ARM CAVITY BAFFLE SUSPENSION ASSEMBLY

Figure 15: Suspension Assembly

## D1002582 - SLC BAFFLE TUBE UP ASSEMBLY

LIGO Parts:

(1) D1002612 - SLC UPPER TUBE					SN:____________

(1) D1002610 - SLC TUBE UP CONNECTOR PLATE		SN:____________

(1) D1002581 - SLC SUSPENSION ROD SUPPORT		SN:____________

Hardware:

(4) 92200A537 (1/4-20 x 1/2 SHCS)

(4) 92200A242 (10-24 x 1/2 SHCS)

Tools:

3/16” Hex L-Key tool (¼-20 SHCS)

5/32” Hex L-Key tool (10-24 SHCS)

### Insert D1002581 – SLC SUSPENSION ROD SUPPORT into D1002612 – SLC UPPER TUBE on end closest to view holes, see Figure 16. Align with fastener holes. Parts may not slide together with ease.

  - Install hardware 4 places

    - 92200A537 (1/4-20 x 1/2 SHCS)

- Insert D1002610 – SLC TUBE UP CONNECTOR PLATE into D1002612 – SLC UPPER TUBE on end furthest from view holes, see Figure 16: Baffle Tube Up Assembly. Align with fastener holes.

  - Install hardware 4 places

    - 92200A242 (10-24 x 1/2 SHCS)

### D1002582 serial number – Same as D1002612.

Figure 16: Baffle Tube Up Assembly

## D1001007 - SLC ACB LO TUBE ASSEMBLY

LIGO Parts:

(1) D1001009 - ARM CAVITY BAFFLE LO TUBE			SN:____________

(1) D1000684 - SLC TUBE LOWER MOUNTING PLATE		SN:____________

(1) D1002618 - SLC TUBE LOWER CONNECTOR PLATE		SN:____________

(1) D1000930 - SLC MAGNET HOLDER STEEL PLATE			SN:____________

Other Parts:

(4) N35P1000500HT - BUNTING MAGNETIC-NEODYMIUM (1.00D x .50H)

Hardware:

(4) 92200A537 (1/4-20 x ½ SHCS)

(2) 94518A510 (1/4-20 x ½ FHCS)

(4) 92200A242 (10-24 x ½ SHCS)

Tooling/Fixtures:

D1201014 – Magnet Spacer Tool

Small test magnet for Section II.B.4

D1200635 – Magnet Placement Tool

Tools:

3/16” Hex L-Key tool (¼-20 SHCS)

5/32” Hex L-Key tool (10-24 SHCS)

Philips Screwdriver

Figure 17: Lo Tube Assembly

- Insert D1000684 - SLC TUBE LOWER MOUNTING PLATE into D1001009 - ARM CAVITY BAFFLE LO TUBE, see Figure 17. Align with larger fastener holes.

  - Install hardware 4 places

    - 92200A537 (1/4-20 x ½ SHCS)

- Insert D1002618 - SLC TUBE LOWER CONNECTOR PLATE into D1001009 - ARM CAVITY BAFFLE LO TUBE, see Figure 17. Align with smaller fastener holes.

  - Install hardware 4 places

    - 92200A242 (10-24 x ½ SHCS)

- Attach D1000930 - SLC MAGNET HOLDER STEEL PLATE to D1002618 - SLC TUBE LOWER CONNECTOR PLATE, see Figure 18.

  - Install hardware 2 places

    - 94518A510 (1/4-20 x ½ FHCS)

- Attach N35P1000500HT - BUNTING MAGNETIC-NEODYMIUM Magnets in 4 places.

  - Use Teflon tool D1200635 to install magnets.

  - Alternate polarity of magnets (need small test magnet to test polarity)

  - If magnets are wrong, use tool to extract the wrong magnets; then, reverse the magnets as required.

- D1001007 serial number – Same as D1001009.

Figure 18: Lo Tube magnet placement

## D1002564 - SLC EDDY CURRENT DAMPING 8 DIA TUBE ASSEMBLY

LIGO Parts:

(1) D1002560 - SLC DAMPING TUBE TOP PLATE		SN:____________

(1) D1002561 - SLC DAMPING 8 INCHES DIA TUBE		SN:____________

(1) D1002617 - SLC DAMPING TUBE LOWER PLATE		SN:____________

(1) D1000929 - SLC COPPER SUPPORT PLATE			SN:____________

(1) D1000909 - SLC COPPER PLATE				SN:____________

Hardware:

(12) 92200A242 (10-24 x ½ SHCS)

(2) 90233A815 (10-24 x ½ FHCS TITANIUM)

(2) 93286A044 (#25 FLAT WASHER ALUMINUM)

(2) 95435A755 (1/4-20 x 3/4 SHCS TITANIUM)

Tools:

3/16” Hex L-Key tool (¼-20 SHCS)

5/32” Hex L-Key tool (10-24 SHCS)

Phillips screw driver

Figure 19: Eddy Current Damping Tube Assembly

- Insert D1002560 - SLC DAMPING TUBE TOP PLATE into D1002561 - SLC DAMPING 8 INCHES DIA TUBE, see Figure 19. Align with fastener holes and cut-outs.

  - Install hardware 6 places

    - 92200A242 (10-24 x ½ SHCS)

- Insert D1002617 - SLC DAMPING TUBE LOWER PLATE into D1002561 - SLC DAMPING 8 INCHES DIA TUBE, see Figure 19. Top and lower plate edges should be aligned with fastener holes.

  - Install hardware 6 places

    - 92200A242 (10-24 x 1/2 SHCS)

- Attach D1000909 – SLC COPPER PLATE to D1000929 – SLC COPPER SUPPORT PLATE.

  - Install hardware 2 places

    - 90233A815 (10-24 x 1/2 FHSC)

- Loosely attach D1000929 - SLC COPPER SUPPORT PLATE to D1002617 - SLC DAMPING TUBE LOWER PLATE in orientation shown in Figure 20. Fasteners will be tightened in Section II.E.12.

- Install hardware 2 places

- 93286A044 (#25 FLAT WASHERS ALUM)

- 95435A755 (1/4-20 x 3/4 SHCS TITANIUM)

- D1002564 serial number - Same as D1002561.

Figure 20: Copper Plate location

## D1001005 - ARM CAVITY BAFFLE BLADE ASSEMBLY

Record assembly TYPE: -2 for BSC1   Record assembly serial number – Same as D1002608

Record assembly TYPE: -3 for BSC3   Record assembly serial number – Same as D1002608

LIGO Parts:

(1) D1002609 - SLC BLADE MOUNTING BRACKET			SN:____________

(1) D1200903 - SLC ACB INTERFACE MOUNTING PLATE (BSC1)	SN:____________

(1) D1200920 - SLC ACB INTERFACE MOUNTING PLATE (BSC3) 	SN:____________

(1) D1002608 - SLC ACB SUSPENSION BLADE				SN:____________

(1) D1002844 - SLC ACB BLADE CLAMP					SN:____________

(1) D1200781 - SLC ACB SUSPENSION ROD				SN:____________

Hardware:

(7) C-1820-NA (5/16-18 x 1 1/4 SHCS Ag)

(7) 98019A385 (5/16 FLAT WASHER)

Tooling/Fixtures:

(1) D1101099 - ARM CAVITY BAFFLE BEND FIXTURE

(1) D1101298 - BLADE SPRING BENDING HOOK ASSEMBLY

(1) D1101184 - SLC BEND FIXTURE PLATE

(4) D1001700 - SLC INTERFACE MOUNTING CLAMP

VARIOUS 1lb-20lb WEIGHTS

Tooling/Fixture Hardware:

(4) C-2008-NA (1/4-20 x ½ SHCS Ag)

(4) 98019A355 (1/4  FLAT WASHER)

(2) C-2016-NA (1/4-20 x 1 SHCS Ag)

(2) C-2012-NA (1/4-20 x ¾ SHCS Ag)

(2) 92200A246 (10-24 x 7/8 SHCS)

Tools:

1/4” Hex L-Key tool (5/16-18 SHCS)

3/8” Wrench (#10 Hex Nut)

3/16” Hex L-Key tool (¼-20 SHCS)

5/32” Hex L-Key tool (10-24 SHCS)

SAFETY GLASSES MUST BE WORN DURING THE FOLLOWING ASSEMBLY STEPS

- Secure D1200903/ D1200920 - INTERFACE MOUNTING PLATE to table with (4) D1001700 - SLC INTERFACE MOUNTING CLAMP spread to distribute weight. Placement of clamps should allow for placement of mounting bracket in next step.

  - Install hardware 4 places

##### C-2032-NA (1/4-20 x 2 SHCS Ag)

##### 98019A355 (1/4  FLAT WASHER)

- Attach D1002609 - SLC BLADE MOUNTING BRACKET in a backwards position to top of D1200903/ D1200920 - SLC ACB INTERFACE MOUNTING PLATE using D1101290 to guide placement of bracket to measure ½” from edge of MOUNTING PLATE.

  - Install hardware 4 places

    - C-1820-N (5/16-18 x 1 ¼ SHCS)

    - 98019A385 (5/16 FLAT WASHER)

- Attach D1002844 - SLC ACB BLADE CLAMP and D1002608 - SLC ACB SUSPENSION BLADE to BLADE MOUNTING BRACKET. Blade clamp lip should line up with suspension blade.

  - Install hardware 3 places

    - C-1820-NA (5/16-18 x  SHCS)

    - 98019A385 (5/16 FLAT WASHER)

Note: Guide washers out of clearance for ¼” holes.

- Attach D1101099 - aLIGO ACB BEND FIXTURE to D1002844 – SLC ACB BLADE CLAMP. Clearance holes in BEND FIXTURE fit over 3 socket head cap screws attaching BLADE CLAMP, SUSPENSION BLADE to BLADE MOUNTING BRACKET.

  - Install hardware 4 places

    - C-2008-NA (1/4-20 x ½ SHCS Ag)

- Turnover INTERFACE MOUNTING PLATE and place on table with bracket clearing edge of table, see Figure 21.

Figure 21: Spring Blade Loadin

- Secure D1200903/ D1200920 - INTERFACE MOUNTING PLATE with (4) D1001700 - SLC INTERFACE MOUNTING CLAMP spread to distribute weight. to table so that it can sustain 180 lbs. suspended from bracket.

- Attach D1101298 - BLADE SPRING BENDING HOOK ASSEMBLY to tip of D1002608 - SLC ACB SUSPENSION BLADE.

- Attach approx. 160bs of weights onto hook end of D1101298 - BLADE SPRING BENDING HOOK ASSEMBLY bending SUSPENSION BLADE to align with curve on BEND FIXTURE. Weight opening placement should be staggered.

- Record weight applied to bend blade - ________________ lbs/grams/kg

- Attach D1101184 - aLIGO ACB BEND FIXTURE PLATE to D1101099 - ARM CAVITY BAFFLE BEND FIXTURE with D1002608 - SLC ACB SUSPENSION BLADE SUSPENSION BLADE sandwiched between both parts. This should be done when spring is fully bent.

  - Install hardware in 2 places

    - 92200A246 (10-24 x 7/8 SHCS)

Note: SUSPENSION BLADE should be securely attached to BEND FIXTURE before continuing to next step.

- Slowly remove weights and D1101298 - BLADE SPRING BENDING HOOK ASSEMBLY.

- Carefully release D1200903/ D1200920 - SLC ACB INTERFACE MOUNTING PLATE from table.

- Flip assembly over and place securely on table.

- Temporarily attach (1) D1001700 – SLC INTERFACE MOUNTING CLAMP to D1200903/ D1200920 - SLC ACB INTERFACE MOUNTING PLATE - closer to the BLADE MOUNTING BRACKET to ensure clearance for removal. This will secure MOUNTING PLATE to table during assembly.

- Carefully detach the assembled D1002609 - SLC BLADE MOUNTING BRACKET from the D1200903/ D1200920 - SLC ACB INTERFACE MOUNTING PLATE, rotate 180˚ and re-attach to table.

  - Removing and installing hardware in 4 places

    - C-1820-N (5/16-18 x 1 ¼ SHCS)

    - 98019A385 (5/16 FLAT WASHER)

Note: Do not install a screw to the center hole of BLADE MOUNTING BRACKET.

- Remove D1101099 - aLIGO ACB BEND FIXTURE from D1002844 – SLC ACB BLADE CLAMP.

  - Remove hardware installed in 6 places.

    - (4) C-2008-NA (1/4-20 x ½ SHCS Ag)

    - (2) 92200A246 (10-24 x 7/8 SHCS)

Figure 22: Blade Mounting Bracket position

- Attach D1200781 - SLC ACB SUSPENSION ROD to tip of SUSPENSION BLADE. Short threaded end into blade tip. Orient the flat of rod so that a wrench can hold the flat from the blade tip direction, see Figure 23.

Figure 23: Suspension Rod

Note: Suspension Rod orientation is important. Make sure Suspension Rod is lined up with the center hole of the Interface Mounting Plate.

## ACB SUSPENSION ASSEMBLY for BSC1 & BSC3 –

Record assembly serial number - Same as D1002564

LIGO Parts and Assemblies:

(1) D1002582 - SLC BAFFLE TUBE UP ASSEMBLY			SN:____________

(1) D1001005-2 – ACB BLADE ASSEMBLY (for BSC1)			SN:____________

(1) D1001005-3 – ACB BLADE ASSEMBLY (for BSC3)			SN:____________

(1) D1002564 - EDDY CURRENT DAMPING 8 DIA TUBE ASSY	SN:____________

(2) D1201361 - EARTHQUAKE STOP SLEEVE				SN:____________

(2) D1201235 – LOCKED THREADED STUD				SN:____________

(2) D1001120 - SLC EARTHQUAKE STOP RING 			SN:____________

(1) D1001007 - ACB TUBE LO ASSEMBLY				SN:____________

(1) D1300531 - ACB UPPER MOUNTING HINGE				SN:____________

(4) D1300526 – UPPER HINGE CAPTIVE SCREWS

(1) D1001186 - 3/4-10 x 4 HEX HEAD SCREW, MODIFIED		SN:____________

(2) D1102316 - ¾-10 x 3/8 NI-CU HEX NUT, MODIFIED			SN:____________

(1) D1101285 – TRANSPORT, LOCKING BRACKET			SN:____________

LIGO Assemblies:

- D1201228 – ANTI-ROTATION ASSEMBLY				SN – Same as D1201198

(1) D1201199– ANTI-ROTATION CLAMP			SN:____________

(1) D1201198– ANTI-ROTATION BLOCK			SN:____________

(6) 1185-5EN938 (HELICOIL, 5/16-18 X 0.938 LONG)

(4) 1185-4EN250 (HELI-COIL INSERT, 1/4-20 X 1/4 LG)

(6) C-1832-N (5/16-18 x 2 SHCS)

(6) WF-31 (5/16 FLAT WASHER)

Hardware:

(1) N-1024-A (10-24 HEX NUTS Ag)

(5) 98019A330 (#10 FLAT WASHER)

(4) C-2016-NA (1/4-20 x 1 SHCS Ag)

(12) 98019A355 (¼ FLAT WASHER)

(2) 98017A200 (3/8 FLAT WASHER)

(4) 92200A242 (10-24 x 1/2 SHCS)

(8) C-2016-N (1/4-20 x 1 SHCS)

(3) 98017A220 (3/4 FLAT WASHER)

(4) 98019A509 (FLAT WASHER .50 Sz)

(2) 91845A031 (3/8-16 HEX NUT)

(6) WF-25 (¼ FLAT WASHER)

(6) C-2012-N (1/4-20 x .75 SHCS)

(6) 1185-4EN250 (HELI-COIL INSERT, 1/4-20 X 1/4 LG)

(2) V1156 2-116 (2-116 FKM VITON O-RINGS 75D BLACK PARKER V747-75)

Tooling/Fixtures:

(1) D1101288 - ALIGNMENT FIXTURE, TUBE UP ASSEMBLY

(1) D1201251 - VARIABLE HEIGHT BRACKET, 9 DEGREE

(1) D1201250 – SUSPENSION ROD GAUGE

(1) D1201014 – MAGNET SPACING TOOL

Tooling/Fixture Hardware:

(4) C-2016-NA (1/4-20 x 1 SHCS Ag)

(19) 98019A355 (¼ FLAT WASHERS)

(13) 92200A540 (¼-20 x ¾ SHCS)

(1) N-2520-A (¼-20 HEX NUT Ag)

Tools:

1/4” Hex L-Key tool (5/16-18 SHCS)

1/8” Hex L-Key tool (10-24 FHCS)

3/8” Wrench (#10 Hex Nut)

3/16” Hex L-Key tool (¼-20 SHCS)

5/32” Hex L-Key tool (10-24 SHCS)

9/16” Wrench (3/8-16 Stud/Nut)

1-1/8” Wrench (¾-10 Nut)

(2) 1 1/8” Hex L-Key tool (10-32 SHCS)

### Install Heli-Coil inserts into D1201198 – ANTI-ROTATION BLOCK.

#### Install hardware in 10 places.

##### (4) 1185-4EN250 (HELI-COIL INSERT, 1/4-20 X 1/4 LG)

##### (6) 1185-5EN938 (HELICOIL, 5/16-18 X 0.938 LONG)

### Install Heli-Coil inserts into D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE.

#### Install hardware in 6 places.

##### (6) 1185-4EN250 (HELI-COIL INSERT, 1/4-20 X 1/4 LG)

### Add V1156 2-116 - O-Rings to D1000909 – SLC COPPER PLATE securely, 2 places.

THE FOLLOWING ASSEMBLY PROCEDURE IS PERFORMED UPSIDE DOWN WITH THE INTERFACE MOUNTING PLATE ON A TABLE

### Attach hardware to D1200781 - SLC ACB SUSPENSION ROD. Located on D1001005-2/-3 - ARM CAVITY BAFFLE BLADE ASSEMBLY.

#### Install hardware in 1 place

    - N-1024-A (10-24 HEX NUT)

    - 98019A330 (#10 FLAT WASHER)

### Using the D1201250 – SUSPENSION ROD GAUGE, install hardware on long tread of rod. Once measurement is taken loosen nut 1 ½ turns before placing tube. Be very careful not to bend the suspension rod.

Figure 24: Suspension Rod Hardware

### Temporarily place D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY to align earthquake holes to D1002582 - SLC BAFFLE TUBE UP ASSEMBLY. Once aligned remove D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY.

### Attach D1002582 - SLC BAFFLE TUBE UP ASSEMBLY to D1200781 - SLC ACB SUSPENSION ROD located on D1001005 – ARM CAVITY BAFFLE BLADE ASSEMBLY. Screw the TUBE UP ASSEMBLY into the rod until it touches the D1201250 – SUSPENSION ROD GAUGE.

### Attach D1101288 - ALIGNMENT FIXTURE, TUBE UP ASSEMBLY to D1200903/ D1200920 - SLC ACB INTERFACE MOUNTING PLATE on D1001005 – ARM CAVITY BAFFLE BLADE ASSEMBLY. Bottom slotted part must be loose to slide to the interface plate.

  - Install hardware in 2 places

    - C-2016-NA (1/4-20 x 1 SHCS Ag)

    - 98019A355 (¼ FLAT WASHER)

### Align and level D1002582 - SLC BAFFLE TUBE UP ASSEMBLY. Have measuring fixture in place while twisting TUBE UP.

### Attach D1101288 - ALIGNMENT FIXTURE, TUBE UP ASSEMBLY to the bottom of D1002582 - SLC BAFFLE TUBE UP ASSEMBLY.

  - Install hardware in 2 places

- C-2016-N (1/4-20 x 1 SHCS)

- 98019A355 (¼ FLAT WASHER)

### Recheck measurement by holding suspension rod in place with wrench. With an addition wrench tighten the nut at correct measurement. Be sure not to let rod twist.

Figure 25: Suspension Rod and Tube Up Assembly

### Remove D1101288 - ALIGNMENT FIXTURE, TUBE UP ASSEMBLY.

  - Remove hardware in 4 places

    - C-2016-NA (1/4-20 x 1 SHCS Ag)

    - 98019A355 (¼ FLAT WASHER)

### Place D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY on D1001005 – ARM CAVITY BAFFLE BLADE ASSEMBLY. With one person holding base of UP TUBE ASSEMBLY, then reaching inside of 8” TUBE and grabbing loose end of UP TUBE ASSEMBLY.

### Insert (2) D1201235 – LOCKED THREADED STUD through D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY and D1002582 - SLC BAFFLE TUBE UP ASSEMBLY. Mate STUD to (2) D1201361 - EARTHQUAKE STOP SLEEVE. TUBE UP ASSEMBLY needs to be held while threaded studs are getting placed.

  - Install hardware in 2 places (stud, .5 washer, tube, .5 washer, sleeve, 3/8 washer, nut).

##### 98017A200 (3/8 FLAT WASHER)

##### 98019A509 (FLAT WASHER .50 Sz)

##### 91845A031 (3/8-16 HEX NUT)

### Attach D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY to D1001005 – ARM CAVITY BAFFLE BLADE ASSEMBLY. Measurement from end of mounting plate to base of 8” TUBE should be ½”. TUBE UP ASSEMBLY still being held in place until EQ RING in Step 10 is in place.

  - Install hardware in 4 places

    - C-2016-NA (1/4-20 x 1 SHCS Ag)

    - 98019A355 (¼ FLAT WASHER

### Attach (2) D1001120 - SLC EARTHQUAKE STOP RING to D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY.

  - Install hardware in 4 places

    - 92200A242 (10-24 x ½ SHCS)

    - 98019A330 (#10 FLAT WASHER)

### Attach D1001007 - ACB TUBE LO ASSEMBLY to D1002582 - SLC BAFFLE TUBE UP ASSEMBLY. Loosely, do not tighten.

  - Install hardware in 4 places

    - C-2016-N (1/4-20 x 1 SHCS)

##### 98019A355 (¼ FLAT WASHER)

### Adjust distance between the magnets and COPPER PLATE to be 0.090” using D1201014 – MAGNET SPACING TOOL. Tighten screws from Section II.C.4.

### Attach D1101285 - TRANSPORT, LOCKING, ACB to D1002564 - SLC EDDY CURRENT DAMPING 8 DIAMETER TUBE ASSEMBLY and D1001007 - ACB TUBE LO ASSEMBLY. Place on side next to BLADE base. Align flush to 8” TUBE edge and tighten.

  - Install hardware in 4 places

##### C-2016-N (1/4-20 x 1 SHCS)

##### 98019A355 (¼ FLAT WASHER)

### Tighten TUBE LO ASSEMBLY to TUBE UP ASSEMBLY and loosen D1101285 - TRANSPORT, LOCKING, ACB.

### Assemble D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE to D1001186 (SCREW HEX HD 3/4-10 x 4, MODIFIED)

  - Install hardware listed (screw, washer, hinge, washer, nut, nut, washer, tube)

    - (2) D1102316 (3/4-10 x 3/8 HEX NUT NI-CU, MODIFIED)

    - (3) 98017A220 (¾ FLAT WASHER)

### Attach D1201198 – ANTI-ROTATION BLOCK to D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE.

  - Install hardware in 6 places

    - C-2012-N (1/4-20 x ¾ SHCS)

    - WF-25 (1/4 FLAT WASHER)

### Insert D1001186 (SCREW HEX HD 3/4-10 x 4, MODIFIED) into D1001007 - ACB TUBE LO ASSEMBLY. Screw in approximately 1.5 inches. Final rest should have the hinge of the D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE positioned on the same side as the D1000909 – SLC COPPER PLATE.

### Attach D1201251 VARIABLE HEIGHT BRACKET, 9 DEGREE to 8” TUBE. Place on side opposite to BLADE base. Edge of BRACKET should be flush with 8” TUBE base. Tighten to 8” TUBE base.

  - Install hardware in 2 places.

##### 92200A540 (1/4-20 x 3/4 SHCS)

##### 98019A355 (¼ FLAT WASHER)

### Adjust distance between D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE and D1001007 - ACB TUBE LO ASSEMBLY by Bolt and Nuts. The scribe line on D1201198 – ANTI-ROTATION BLOCK must line up with the scribe line on D1201251 VARIABLE HEIGHT BRACKET, 9 DEGREE.

### Attach D1201251 VARIABLE HEIGHT BRACKET, 9 DEGREE to D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE and D1201198 – ANTI-ROTATION BLOCK.

- Install hardware in 4 places

##### 92200A540 (¼-20 x ¾ SHCS)

##### 98019A355 (¼ FLAT WASHER)

### Tighten D1001186 - SCREW HEX HD 3/4-10 x 4, MODIFIED against D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE and D1001007 - ACB TUBE LO ASSEMBLY.

### Attach D1201199 – ANTI-ROTATION CLAMP to the D1201198 – ANTI-ROTATION BLOCK.

- Install hardware in 6 places

##### C-1832-N (5/16-18 x 2 SHCS)

##### WF-31 (5/16 FLAT WASHER)

### Set spacing between MAGNETS and COPPER PLATE to approximately 0.090 using D1201014 – MAGNET SPACING TOOL. Tighten the COPPER PLATE to the 8” TUBE.

### Insert (4) D1300526 - ACB MOUNTING HINGE CAPTIVE SCREWS to D1300531 - ARM CAVITY BAFFLE UPPER MOUNTING HINGE

### Torque all screws, as required.

### Weigh assembly and record: _______________ lbs. (kg/grams)

# Text boxes

D1002610 –

SLC TUBE UP CONNECTOR PLATE

D1002581 –

SLC SUSPENSION ROD SUPPORT

D1002612 –

SLC UPPER TUBE