# Cable Visual Inspection Summary

- CONNECTORS – match drawings.

If not ⟹ reject

Right

Metalization and ears match drawing

Ears

Metalized

Plain PEEK

Metalized

Wrong

Connector metalized where it shouldn’t be

DB25

D# – version – length* – s/n

length*:	D1000225

D1000234

D1002522

microD25

D# – version –length* – s/n

length*:	D1002522

DB9

D# – version – position* – s/n

length*:	D1000227

microD9

Leg (A, B, C, D) – first 3 digits of s/n

D3

D# – version – s/n

Mighty Mouse

Leg (A, B, C, D)

SCRIBED – on every connector. Scribe unmarked connectors on the side without a vent hole.

Length must be scribed only if a drawing has multiple lengths.

If not ⟹ scribe

- CLAMPING BAND – has double band-it, and

does not slide off easily during twist and pull test.

If single band-it does not slide off ⟹ pass

If band-it (single or double) slides off ⟹ replace

- SHAGGY SNOUT – copper does not extend past

clamping bands.

If shaggy ⟹ trim

- BROKEN COPPER BRAID – copper does not

protrude through black PEEK overbraid .

If broken ⟹ trim

Right

Double band-it

Not Ideal

Single band-it

- SCREWS – are in place, and facing in the right

direction.

If fillister head screws are missing ⟹ replace

If fillister head screws are backwards ⟹ reposition

If captive screws are missing ⟹ reject if pre-bake,

replace if post-bake

Wrong

Captive screws backwards

bac

Right

Captive crews face right way

bac

Right	               Wrong

Has screws       Lost screws

bac

Wrong

Shaggy snout

Wrong

Broken copper sticking out

- HELICOIL HOLES – are sized correctly. Use gage to verify.

If not ⟹ re-tap hole