*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T1200616-v2 *ADVANCED LIGO* 23 January 2012

Buoyancy correction for suspensions

Mark Barton, Brett Shapiro

Distribution of this document:

DCC

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology**  **LIGO Project – MS 18-34**  **1200 E. California Blvd.**  **Pasadena, CA 91125**  Phone (626) 395-2129  Fax (626) 304-9834  E-mail: [info@ligo.caltech.edu](mailto:info@ligo.caltech.edu) | **Massachusetts Institute of Technology**  **LIGO Project – NW22-295**  **185 Albany St**  **Cambridge, MA 02139**  Phone (617) 253-4824  Fax (617) 253-7014  E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory**  **P.O. Box 1970**  **Mail Stop S9-02**  **Richland WA 99352**  Phone 509-372-8106  Fax 509-372-8137 | **LIGO Livingston Observatory**  **P.O. Box 940**  **Livingston, LA 70754**  Phone 225-686-3100  Fax 225-686-7189 |

http://www.ligo.caltech.edu/

**Table of Contents**

1 Introduction 3

1.1 Purpose and Scope 3

1.2 References 3

1.3 Version history 3

2 Background 3

3 Theory 4

4 Results 5

4.1 Test mass chain 5

4.2 Compensation plate chain 5

4.3 End reaction mass chain 5

4.4 BS/FM metal build (as for one-arm test dummy installation) 6

4.5 BS/FM glass build 6

5 Installation procedure 6

# Introduction

## Purpose and Scope

When a chamber containing a suspension is pumped down, the loss of buoyancy from the air causes a slight sag of the hanging masses, with implications for OSEM and EQ stop adjustment. The initial version of this document estimates the magnitude for quad TM, CP and ERM chains, updating the calculation first made by Brett for the noise prototype at LASTI. -v2 adds the BS/FM. Later versions will add estimates for other suspension types.

## References

LASTI ilog: <http://www.ligo.mit.edu/ilog/pub/ilog.cgi>

Brett’s LASTI ilog entries of [5/13/09](http://emvogil-3.mit.edu/ilog/pub/ilog.cgi?group=lasti&task=view&date_to_view=05/12/2009&anchor_to_scroll_to=2009:05:12:20:04:30-brett) and [6/29/10](http://www.ligo.mit.edu/ilog/pub/ilog.cgi?group=lasti&task=view&date_to_view=06/29/2010&anchor_to_scroll_to=2010:06:29:13:40:18-brett)

T1100455: [Analysis of position/voltage data on a BOSEM](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=69627)

SUS SVN (browser view): <https://redoubt.ligo-wa.caltech.edu/websvn/listing.php?repname=sus&>

SUS SVN (client view): <https://redoubt.ligo-wa.caltech.edu/svn/sus/>

M1100256: [Suspension Earthquake Stops to be Set at 0.75 mm](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=72511)

D0901346: [Advanced LIGO Quadruple Suspension](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=3413)

E1101037: [aLIGO BSC ISI/Quad Installation Procedure](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=73780)

D1000392: [aLIGO BS/FM MAIN ASSEMBLY](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=9458)

## Version history

1/12/18: -v1. Initial version with analysis for TM, CP and ERM chains.

1/23/12: -v2. Added analysis for glass and metal build BS/FM. Removed counts values where there was no associated OSEM. Added references, especially M1100256.

# Background

Suspended masses (and other objects) in the open vacuum chambers will experience an upward buoyancy force in proportion to their volume. When a chamber is pumped down, the force due to the air is removed. This turns out to produce a significant sag for the quad suspension, which is vertically soft due to multiple layers of blade springs.

This has the following effects for the quad (the BS/FM is very similar but with only one chain):

1. The LF and RT BOSEMs for the top mass in each chain are attached to the structure and vertically oriented, and will see the drop of the mass directly, as a move away from the working position set in initial adjustment.
2. The F1, F2, F2, and SD BOSEMs for the top mass are attached to the structure and horizontally oriented, and will be de-centered slightly. This is unlikely to be a big problem but they could change their gain slightly, or, if they are poorly centered to begin with, they could be brought to the point of fouling.
3. The BOSEMs and AOSEMs on the upper-intermediate and penultimate masses will be also decentered slightly if the sag is different between the chains (which it turns out it is).
4. All of the masses will move vertically relative to the earthquake stops. This is particularly significant for the lower masses.

# Theory

The following theory is an update of the calculation used in Brett’s LASTI ilog entries of [5/13/09](http://emvogil-3.mit.edu/ilog/pub/ilog.cgi?group=lasti&task=view&date_to_view=05/12/2009&anchor_to_scroll_to=2009:05:12:20:04:30-brett) and [6/29/10](http://www.ligo.mit.edu/ilog/pub/ilog.cgi?group=lasti&task=view&date_to_view=06/29/2010&anchor_to_scroll_to=2010:06:29:13:40:18-brett).

The LIGO LVEA is kept at around 70F (John Worden). The density of air at this temperature is  1.2 kg/m3.

The force on a single mass depends on the volume, which in turn depends on the mass and average density:



where .

The glass masses are fairly homogeneous and for the density we take the bulk value of 2200 kg/m3. The metal masses are a composite of different materials, so we take as representative the density of stainless steel, 7900 kg/m3.

A disappearing buoyancy force at one mass creates sag at all four masses, as the force propagates via the springs. This is conveniently summed up in a compliance matrix, and the net effect at each mass can be calculated by multiplying the vector of forces at the four masses by the compliance matrix to get a vector of sag displacements:



Mathematica models of the three quad chain types of interest (TM, CP, ERM) are in the following directory of the SUS SVN

^/trunk/Common/MathematicaModels/QuadLite2Lateral/mark.barton

The cases of interest are 20111121TMproductionTM, 20111121TMproductionCP and 20111121TMproductionERM.

Because the blade springs are common to all chains, there is a single quad compliance matrix (units in mm/100 g):

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Top | UIM | PM | TM |
| Top | 0.344389 | 0.344389 | 0.344389 | 0.344389 |
| UIM | 0.344389 | 0.643029 | 0.643029 | 0.643029 |
| PUM | 0.344389 | 0.643029 | 0.850089 | 0.850089 |
| TM | 0.344389 | 0.643029 | 0.850089 | 0.864969 |

The derivation of this is in the SUS SVN at

…/20111121TMproductionTM/sagcalc/ASUS4XLLateralModelCalcSag.nb.

The compliance matrix for the BS/FM is

|  |  |  |  |
| --- | --- | --- | --- |
|  | Top | IM | BS/FM |
| Top | 0.344389 | 0.344389 | 0.344389 |
| IM | 0.344389 | 0.643029 | 0.643029 |
| BS/FM | 0.344389 | 0.643029 | 0.864969 |

The derivation is at

^/trunk/Common/MathematicaModels/TripleLite2/mark.barton/20120120bsNW/sagcalc/ASUS3L2ModelCalcSag.nb. The model used corresponds to Jeff Kissel’s production BS/FM model in Matlab at

^/trunk/Common/MatlabTools/TripleModel\_Production/bsfmopt\_metal.m (revision 2005 of 1/19/12)

For the BOSEMs, the sensitivity to flag position as a fraction of open light is 1.22 OL/mm (average of three values from T1100455-v2), and BOSEMs will be operated with digital normalization to 30000 counts, for a net sensitivity of 36600 counts/mm.

# Results

## Test mass chain

parameter set: mark.barton/20111121TMproductionTM

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mass (kg) | Density | Volume (cc) | Buoyancy (g) | Sag (mm) | Sag (counts) |
| Top | 21.999 | 7900 | 2.78 | 3.34 | 0.17 | 4221 |
| UIM | 21.526 | 7900 | 2.72 | 3.27 | 0.31 |  |
| PUM | 39.633 | 2200 | 18.02 | 21.62 | 0.40 |  |
| TM | 39.631 | 2200 | 18.01 | 21.62 | 0.40 |  |

## Compensation plate chain

parameter set: mark.barton/20111121TMproductionCP

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mass (kg) | Density | Volume (cc) | Buoyancy (g) | Sag (mm) | Sag (counts) |
| Top | 22.002 | 7900 | 2.79 | 3.34 | 0.09 | 2248 |
| UIM | 21.532 | 7900 | 2.73 | 3.27 | 0.16 |  |
| PRM | 59.229 | 7900 | 7.50 | 9.00 | 0.20 |  |
| CP | 20.039 | 2200 | 9.11 | 10.93 | 0.20 |  |

## End reaction mass chain

parameter set: mark.barton/20111121TMproductionERM

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mass (kg) | Density | Volume (cc) | Buoyancy (g) | Sag (mm) | Sag (counts) |
| Top | 22.002 | 7900 | 2.79 | 3.34 | 0.10 | 2412 |
| UIM | 21.532 | 7900 | 2.73 | 3.27 | 0.17 |  |
| PRM | 53.536 | 7900 | 6.78 | 8.13 | 0.22 |  |
| ERM | 25.19 | 2200 | 11.45 | 13.74 | 0.22 |  |

## BS/FM metal build (as for one-arm test dummy installation)

We assume densities corresponding to stainless, aluminum and aluminum.

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mass (kg) | Density | Volume (cc) | Buoyancy (g) | Sag (mm) | Sag (counts) |
| Top | 12.63 | 7900 | 1.60 | 1.92 | 0.05 | 1208 |
| IM | 13.575 | 2700 | 5.03 | 6.03 | 0.09 |  |
| BS/FM | 14.21 | 2700 | 5.26 | 6.32 | 0.10 |  |

## BS/FM glass build

We assume densities corresponding to stainless, aluminium and fused silica.

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mass (kg) | Density | Volume (cc) | Buoyancy (g) | Sag (mm) | Sag (counts) |
| Top | 12.63 | 7900 | 1.60 | 1.92 | 0.05 | 1330 |
| IM | 13.575 | 2700 | 5.03 | 6.03 | 0.10 |  |
| BS/FM | 14.21 | 2200 | 6.46 | 7.75 | 0.11 |  |

# Installation procedure

1. Check at the output of the OSEMINF MEDM screen and make sure the BOSEMs are normalized to ±15000 counts (with the standard sign convention where negative counts mean entered and positive counts mean retracted).
2. Set the LF and RT BOSEMs physically so that they're negative by about the number of counts in green.
3. Check that the V (vertical) value is positive by about the number of counts in green (i.e., that much less than “1/2 full” on the OSEM\_ALIGN speedo display in MEDM).
4. Set the EQ stops so that they will give the 0.75 mm gap called for in M1100256 after the masses have dropped by about the distance in orange. The allowed tolerance is 0.25 mm, so feel free to ignore sag values smaller than that. The priorities are the PUM and TM of the quad main chain.