LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1200107	                 LIGO                   September 19th, 2012

aLIGO HAM-ISI, Installation Test Report, Phase II,

LLO HAM 6

E1200107-V1

Céline Ramet, Michael Vargas

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	3

I.	SIDE CHAMBER TESTING	4

	Step 1: GS13	4

o	Step 1.1 – Horizontal GS-13s	4

o	Step 1.2 – Vertical GS-13s	6

	Step 2: Inventory (E1000052)	9

	Step 3: Check level of Stage 1 Optical Table	10

	Step 4: Shim thickness	10

	Step 5: Blade spring profile	10

	Step 6: Gap checks on actuators-after installation on Stage 1	11

	Step 7: Mass budget	11

	Step 8: Lockers adjustment	13

	Step 9- Electronics Inventory	13

	Step 10– Cables inventory – E1100822	13

	Step 11 - Set up sensors gap	14

	Step 12- Measure the Sensor gap	14

	Step 13 - Check Sensor gaps after the platform release	15

	Step 14– Performance of the limiter	16

	Step 14.1 - Test Nº1 - Push “in the general coordinates”	16

	Step 15- Position Sensors unlocked/locked  Amplitude Spectral Densities	17

	Step 16 - GS13 ASD -tabled tilted	19

	Step 17- GS13 pressure readout	20

	Step 18- Actuators Sign and range of motion (Local drive)	20

	Step 19 - Static Testing (Tests in the local basis)	21

	Step 20- Linearity test	22

	Step 21 Cartesian Basis Static Testing	23

	Step 22- Frequency response	25

	Step 22.1 - Local to local measurements	25

Conclusion of Side Chamber testing	29

Introduction

This document presents the tests performed to characterize and validate the modifications made to the eLIGO HAM-ISI, previously installed in HAM 6 at LIGO Livingston. This unit was removed from the vacuum system in November 2010, kept in a shipping container with weekly dry Nitrogen purge until June 2012.

This was the 4th unit pulled out from storage for installation at Livingston, to be installed in HAM 6 chamber. All seismometers, Capacitive Positions Sensors, and in-vacuum cables were swapped for aLIGO ones. Also, the Capacitive Position Sensors cables were shielded.

There are 4 distinctive parts:

- Part I: side chamber testing results

- Part II: post insertion testing results

- Part III: Transfer functions taken during optics installation 
	3.1 transfer functions after first triple install 
	3.xx ....

- Part IV: Tests in final configuration

# SIDE CHAMBER TESTING

Side chamber testing was conducted in the LLO LVEA with actual electronics and field cables from June 22nd until July 3rd 2012. Temporary extensions to those cables were used to reach the test stand located further away from the electronics than the HAM chamber is.

Steps 1 and 2 capture data from testing done previously, whereas all the following steps were done (or waived) during that side-chamber testing period.

## Step 1: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at:

seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 1.1 – Horizontal GS-13s

Huddle testing

Figure  - Huddle testing of Horiz GS-13 859, 824, and ??? after complete aLIGO modifications

## Step 1.2 – Vertical GS-13s

Huddle testing

Figure  - Huddle testing of Vert GS-13 679,683 and 699 after aLIGO modifications

Driven testing

Figure  - Driven Transfer Function of Vert GS-13 679,683 and 699 after aLIGO modifications

Issues/difficulties/comments regarding this test:

Not all GS-13s have exactly the same resonant frequency (cf huddle test results). This parameter was adjusted as well as manageable after modifications of the instruments. Instruments were passed as long as their resonant frequencies were within 0.95 and 1.2 Hz.

Some instruments appear out of phase compared to the model. This only means the instrument was placed at 180 deg compared to the expected position when the model was created . This was ignored in order to approve the instrument for use.

Acceptance Criteria

- All instruments match their expected response.

Test result:	Passed:   X   	Failed:         .

## Step 2: Inventory (E1000052)

Data shown in red indicate changes made  side-chamber, ** indicate parts older than 2009.

| DCC/Vendor number | Part name | Configuration | S/N | S/N | S/N |
| --- | --- | --- | --- | --- | --- |
| D071001 | Stage 0 base |  | ** |  |  |
| D071051 | Stage 1 base |  | ** |  |  |
| D071050 | Optical table |  | ** |  |  |
| D071002 | Spring Post |  | ** | ** | ** |
| D071100 | Spring |  | ** | ** | ** |
| D071102 | Flexure |  | ** | ** | ** |
| ADE | Position sensor | Horizontal | 12044 | 11984 | 12072 |
|  |  | Vertical | 12074 | 12065 | 12063 |
| D047812 | GS-13 pod | Horizontal | 20 | 37 | 84 |
|  |  | Vertical | ?? | 36 | 40 |
| D047823 | L4C pod | Horizontal | N/A | N/A | N/A |
|  |  | Vertical | N/A | N/A | N/A |
| D0902749 | Actuator | Horizontal | 19** | 09** | 14** |
|  |  | Vertical | 20** | 02** | 10** |

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 3: Check level of Stage 1 Optical Table

Due to past experience showing we had to adjust the shims both side chamber and in chamber we decided not to run this test.

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

Test result:						Passed:      	Failed:         . Waived:   X_

## Step 4: Shim thickness

Issues/difficulties/comments regarding this test: We adjusted shims under 2 lockers in order to improve the level of the table.

| Lockers | Shim thickness (mil) |
| --- | --- |
| A |  |
| B |  |
| C |  |
| D |  |

Table – Shims Thickness

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 5: Blade spring profile

Because we know we’ll likely have to adjust the shims in chamber, we did not take this measurement.

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:   _        Failed:          Waived:   X   .

## Step 6: Gap checks on actuators-after installation on Stage 1

Gaps were inspected on all vertical actuators and adjusted to be within requirements. No records were taken. Horizontal actuators will be inspected/adjusted in chambers.

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:   X   	Failed:         .

## Step 7: Mass budget

Figure – Keel Masses and Wall masses location

|  |  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.1 | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 | 2 |  |  | 2 |  |  | 1 | 1 | 47.2 | 21.41 |
| w1 |  |  | 1 |  |  |  | 1 | 1 | 43.9 | 19.91 |
| w2 | 3 |  |  |  | 1 |  | 1 | 1 | 47.3 | 21.45 |
| w3 |  |  |  | 1 |  | 0 | 1 | 1 | 45 | 20.41 |
| w4 |  |  | 1 | 1 | 1 |  | 1 | 1 | 50.6 | 22.95 |
| w5 |  |  | 1 |  |  |  | 1 | 1 | 43.9 | 19.91 |
| w6 | 3 |  | 1 |  |  |  | 1 | 1 | 43.9 | 19.91 |
| w7 | 4 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w8 |  |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| Side Masses Total | 12 | 0 | 4 | 4 | 2 | 0 | 9 | 9 | 407.4 | 184.79 |

Table – Wall masses distribution

|  |  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.1 | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  |  |  |  |  | 0.00 | 0.00 |
| k2 | 1 |  |  |  |  |  | 1 | 1 | 42.80 | 21.23 |
| k3 |  |  |  |  |  |  |  |  | 0.00 | 0.00 |
| k4 | 1 |  |  |  |  |  | 1 | 1 | 42.80 | 21.23 |
| k5 |  |  |  |  |  |  |  |  | 0.00 | 0.00 |
| k6 | 1 |  |  |  |  |  | 1 | 1 | 42.80 | 21.23 |
|  | 3 | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 128.40 | 63.69 |

Table – Keel masses distribution

|  | D972213 | D972214 | D972215 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 2.5 kg | 5 kg | 10 kg |  |  |
|  | 610 | 375 | 230 | 5.5 | 11 | 22 | lbs | kgs |
| A |  | 2 |  |  |  | 2 | 794 | 360.15 |
| B |  |  |  |  |  |  | 0 | 0.00 |
| C |  |  |  |  |  |  | 0 | 0.00 |
| D |  |  |  |  |  |  | 0 | 0.00 |
| E-1 |  |  |  |  | 1 | 2 | 55 | 24.95 |
| E-2 |  |  |  |  |  | 2 | 44 | 19.96 |
| E-3 |  |  |  |  |  | 2 | 44 | 19.96 |
| Top Masses | 0 | 2 | 0 | 0 | 1 | 8 | 937 | 425.02 |

Table - Optical Table Masses distribution

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 184.79 | 63.69 | 425.02 | 673.50 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427) +/-25Kg (5%)

Issues/difficulties/comments regarding this test:

This load is inexplicably high (fails by 20%) in comparison with the aLIGO ISIs. A difference between the springs because those were manufactured from a different batch several years before is possible, but still unlikely to explain such a disparity.

Test result:	Passed:     	Failed:   X      .

## Step 8: Lockers adjustment

We used dial indicators for initial balancing, but fine balancing and locker adjustment was done using the CPSs. (See Step 12).

Test result:	Passed:            Failed:         . Waived:   X__

## Step 9- Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1103331 |
|  |  | S1103318 |
| Anti Image filter | D1100202 | S1200954 |
| Anti aliasing filter | D1000269 | S1200950 |
|  |  | S1200949 |
| Interface chassis | D1000067 | S1200667 |
|  |  | S1107437 |

Table - Inventory electronics

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 10– Cables inventory – E1100822

The location of all cables must be reported in the spreadsheet E1100822.

| Cable Connects | Cable S/N |
| --- | --- |
| Part Name | Configuration | Corner 1 | Corner 2 | Corner 3 |
| GS13 | Horizontal | S1104642-S1104609 | S1104647-S1104596 | S1104641-S1104470 |
| GS13 | Vertical |  |  |  |
| Actuator | Horizontal | S1104752 | S1104475 | S1107757 |
|  | Vertical | S1104473 | S1104716 | S1104750 |

Table – Cables inventory

Acceptance Criteria

- Cable inventory completed

- E110082 spreadsheet updated

Test result:	Passed:   X   	Failed:         .

## Step 11 - Set up sensors gap

|  | No mass |
| --- | --- |
| Table locked | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation |
| H1 | 125.72 | 26.371 |
| H2 | -331.76 | 13.788 |
| H3 | 54.683 | 18.929 |
| V1 | -219.05 | 19.088 |
| V2 | -229.87 | 17.625 |
| V3 | -85.35 | 20.183 |

Table – Capacitive position sensor readout after gap set-up

Issues/difficulties/comments regarding this test: HAM-ISI – LLO HAM 6 uses synchronized boxes, with power boards installed on their back.

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 12- Measure the Sensor gap

This test was not done any more due to risk of damage to sensor targets.

Test result:	Passed:   _        Failed:          Waived:   X

## Step 13 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | 125.72 | 26.371 | -660.85 | -786.57 |
| H2 | -331.76 | 13.788 | -586.56 | -254.8 |
| H3 | 54.683 | 18.929 | -1327.3 | -1382 |
| V1 | -219.05 | 19.088 | -253.35 | -34.3 |
| V2 | -229.87 | 17.625 | -734.46 | 504.58 |
| `V3 | -85.35 | 20.183 | -1216.4 | -1131.1 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 14– Performance of the limiter

## Step 14.1 - Test Nº1 - Push “in the general coordinates”

Issues/difficulties/comments regarding this test:

It is to be noted in this test and in all the following ones involving the actuators, than despite having everything supposedly similar to all the aLIGO ISIs, the actuators from this ISI were always driving in the opposite direction as the others. This will be solved after installation by a modification on the field cables.

|  | CPS read out | Calculated after calibration |
| --- | --- | --- |
| Sensors | UP (Counts) | Down (Counts) | UP (mil) | Down (mil) |
| V1 | 20300 | -20000 |  |  |
| V2 | 20400 | -19500 |  |  |
| V3 | 20300 | -20600 |  |  |
|  | CPS read out | Calculated after calibration |
| Sensors | CW(-RZ) | CCW (+RZ) | CW (mil) | CCW (mil) |
| H1 | 20100 | -22000 |  |  |
| H2 | 21900 | -22700 |  |  |
| H3 | 22800 | -19600 |  |  |

Table - Optic table range of motion

Acceptance criteria:

- The vertical sensor readout be positive when the optic table is pushed in the +Z direction

- The horizontal sensor readout be negative when the optic table is pushed in the +RZ direction

Step 15.1

- Absolutes value of all estimated motions must be higher than 16000counts (~0.020”)

Test result:	Passed:   X   	Failed:         .

## Step 15- Position Sensors unlocked/locked  Amplitude Spectral Densities

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Spectra/Undamped/

- LLO_ISI_HAM6_ASD_m_CPS_T240_L4C_GS13_Locked_vs_Unlocked_2012_07_03.pdf

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Unlocked_Locked_HAM_ISI.m

- Plot_ASD_Unlocked_Locked_Group_HAM_ISI.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM6/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM6_ASD_m_CPS_Requirements_Locked_vs_Unlocked_2012_07_03.pdf

- LLO_ISI_HAM6_ASD_m_GS13_Requirements_Locked_vs_Unlocked_2012_07_03.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13s

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

Issues/difficulties/comments regarding this test:

Test result:	Passed:   X   	Failed:          .

## Step 16 - GS13 ASD -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – ASD Calibrated GS13 with mass at corner

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Spectra/Undamped/

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2012-06-27.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Tilted_Stage_HAM_ISI.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM6/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM6_ASD_CT_GS13_Tilted_2012_06_27.pdf

Acceptance criteria:

- With table unlocked and tilted, magnitudes of power spectra must be fully included within:

Issues/difficulties/comments regarding this test:

One can notice that V1 seismometer, even though displaying similar behavior than the other vertical seismometers, has an amplitude about half as those. This was due to a connection issue between the field cable and the HAM-ISI interface. Those tests were not taken again after the find.

Test result:	Passed:   X   	Failed:          .

## Step 17- GS13 pressure readout

As of 07/02/12, readout was as follow:

|  | Corner 1 | Corner 2 | Corner 3 |
| --- | --- | --- | --- |
| GS-13 Direct Pressure | 99 | 100 | 78 |
| GS-13 Diff Pressure | 0 | 0 | 0 |

Acceptance criteria:

- The pressure on direct channels must be 102KPa +/-8 KPa

- The pressure on differential channels must be 0KPa +/-8 KPa

Issues/difficulties/comments regarding this test:

The channels in orange do not pass because of a gain difference in the GS-13 channels of the HAM-ISI interfaces. Those out of vacuum electronics will be corrected later on.

Test result:	Passed:   X   	Failed:       .

## Step 18- Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | 24356.63 | -23581.5 |
| H2 readout (count) | 23436.76 | -23902.2 |
| H3 readout (count) | 25256.95 | -24624.4 |
| V1 readout (count) | 19121.5 | -19818.9 |
| V2 readout (count) | 24979.59 | -24208.1 |
| V3 readout (count) | 21160.23 | -22650.5 |

Table - Range of motion - Local drive

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Static_Tests/

- LLO_ISI_HAM6_Range_Of_Motion_20120703.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 19 - Static Testing (Tests in the local basis)

|  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- |
| H1 | -1761.71 | -1103.13 | -1100.84 | -15.9095 | 12.5823 | 2.6116 |
| H2 | -1094.45 | -1743.49 | -1097.05 | 17.8259 | 9.667 | -14.0226 |
| H3 | -1119.06 | -1125.8 | -1788.9 | 12.42717 | 12.1842 | -15.1616 |
| V1 | -165.306 | -170.033 | 332.568 | -1278.77 | 37.36716 | 559.6561 |
| V2 | 317.665 | -188.887 | -171.338 | 550.6327 | -1246.62 | 19.856 |
| V3 | -168.966 | 313.483 | -168.426 | 30.41616 | 551.4492 | -1256.83 |

Table - Main and cross coupling

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Static_Tests/

- LLO_ISI_HAM6_Offset_Local_Drive_20120703.mat

Scripts files for taking data in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Static_Test_Local_Basis_HAM_ISI.m

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:   X    	Failed:         .

## Step 20- Linearity test

|  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- |
| H1 | -1.7604 | -1076.71 | -1.76438 | -0.22533 |  |
| H2 | -1.74814 | -913.349 |  | -0.92022 |  |
| H3 | -1.78459 | -1632.13 |  | 1.145553 |  |
| V1 | -1.28621 | 84.82623 | -1.26453 | 1.71472 |  |
| V2 | -1.2655 | -758.596 |  | 0.076528 |  |
| V3 | -1.24188 | -1326.95 |  | -1.79125 |  |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at:

/seismic/HAM-ISI/Common/ Common/Testing_Functions_HAM_ISI

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:

seismic/HAM-ISI/L1/HAM6/Data/Linearity_Test/

- LLO_ISI_HAM6_Linearity_test_20120703.mat

Figures in SVN at:

seismic/HAM-ISI/L1/HAM6/Data/Figures/Linearity_Test/

- LLO_ISI_HAM6_Linearity_test_20120703.pdf

Issues/difficulties/comments regarding this test:

A few actuators display higher slopes difference than expected (i.e. 1% to average). But it was deemed close enough to be acceptable here (field and in-vacuum cables have different lengths).

It took us numerous attempts before passing this test. Indeed we would regularly have one or the other actuator which would display a much lower slope. This was blamed on the unprotected end of the actuator cables (actuator side) contacting any other part of the ISI, therefore reducing the current flowing through the actuator. This can be measured at the air-side of the feedthru (the resistance between the 2 pins would be much higher than intended). This issue was found hard to detect and/or improve, and at this point it is not yet clear why this issue has been repeated so many times on this ISI and not on the others.

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:  X    	Failed:

## Step 21 Cartesian Basis Static Testing

| 1000 counts Drive | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- |
| X Drive | -204.493 | -204.728 | 461.3768 | -36.2418 | 1.80454 | 7.4758 |
| Y Drive | 380.6837 | -372.8 | -3.459 | -7.18695 | 3.55744 | 15.118 |
| Z Drive | -10.2574 | -4.40514 | 0.6114 | -247.434 | -217.85 | -196.366 |
| Rx Drive | 393.2914 | -409.009 | 3.7404 | 383.0866 | -1452.24 | 1043.733 |
| Ry Drive | 231.4814 | 219.8392 | -467.026 | 1474.289 | -335.09 | -1059.94 |
| Rz Drive | 1709.158 | 1704.476 | 1721.099 | 3.860908 | 8.25052 | 44.9336 |

Table - Tests in the local coordinate basis

| 1000 counts Drive | X | Y | Z | RX | RY | RZ |
| --- | --- | --- | --- | --- | --- | --- |
| X Drive | -443.579 | -5.88621 | -15.1806 | 2.052502 | -10.5316 | -17.7086 |
| Y Drive | -7.77595 | -441.642 | -18.2114 | 26.79738 | -24.171 | -6.9564 |
| Z Drive | 0.66663 | -4.74997 | -228.798 | 12.56017 | -11.6906 | 13.0496 |
| Rx Drive | -12.7533 | -13.9779 | -11.7751 | -2236.12 | -19.912 | 11.1894 |
| Ry Drive | -0.97062 | -9.49364 | -16.3234 | -3.48221 | -2284.08 | 11.3914 |
| Rz Drive | -2.26277 | 2166.048 | -19593.2 | 4545.38 | -35835.9 | -20711.3 |

Table - Tests in the general coordinate basis

Acceptance criteria:

- For a positive drive in the Cartesian basis Local sensor readout must have the same sign that in the following table:

| 1000 counts Drive | H1 | H2 | H3 | V1 | V2 | V3 | Direction read out |
| --- | --- | --- | --- | --- | --- | --- | --- |
| X Drive | + | + | - |  |  |  | + |
| Y Drive | - | + | 0 |  |  |  | + |
| Z Drive |  |  |  | + | + | + | + |
| Rx Drive |  |  |  | - | + | - | + |
| Ry Drive |  |  |  | - | + | + | + |
| Rz Drive | - | - | - |  |  |  | + |

Table – Reference table

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Issues/difficulties/comments regarding this test:

Apart from the known sign difference (see introduction), the local coordinate table abides by the reference perfectly. The general coordinate basis one seems more surprising, but at worse, this would be a matrix error, which would not interfere with the decision to install the HAM-ISI.

Test result:	Passed:   X   	Failed:         .

## Step 22- Frequency response

## Step 22.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

seismic/HAM-ISI/L1/HAM6/Data/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM6_Data_TF_L2L_50mHz_500mHz_20120705-174022.mat

- LLO_ISI_HAM6_Data_TF_L2L_500mHz_5Hz_20120705-204555.mat

- LLO_ISI_HAM6_Data_TF_L2L_5Hz_200Hz_20120705-085134.mat

- LLO_ISI_HAM6_Data_TF_L2L_200Hz_800Hz_20120703-172903.mat

Data collection script files:

seismic/HAM-ISI/L1/HAM6/Scripts/Data_Collection/

- Run_Exc_Batch_L1ISIHAM6.m

Scripts files for processing and plotting in SVN at:

seismic/HAM-ISI/L1/HAM6/Scripts/Control_Scripts/release/

- Step_1_TF_Loc_2_Loc_L1_ISI_HAM6.m

(note that here release was soft linked to Version_1)

Figures in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM6_TF_L2L_Raw_from_ACT_to_CPS_2012_07_05.fig

- LLO_ISI_HAM6_TF_L2L_Raw_from_ACT_to_GS13_2012_07_05.fig

Storage of measured transfer functions in the SVN at:

seismic/ HAM-ISI/L1/HAM6/Data/Transfer_Functions/Simulations/Undamped

- L1_ISI_HAM6_TF_L2L_Raw_2012_07_05.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements –Capacitive Position Sensors

Figure - Local to Local Measurements – Inertial sensors

Issues/difficulties/comments regarding this test:

We can see that horizontal sensors in corner 1 and 3 seem to see a slightly different behavior around 1 Hz. We tried to assess whether this could be due to the GS-13 characteristics themselves by extracting the sensor response from the transfer function and comparing this to the huddle test results.

Data collection script files:

seismic/HAM-ISI/L1/HAM6/Scripts/Data_Collection/

- L1_HAM6_GS13_Resp_Extraction_Fitting.m

Figures in SVN at:

/seismic/HAM-ISI/L1/HAM6/Data/Figures /Instrument_Responses/GS13/

- L1_ISI_HAM6_GS13_H1_Pod_20_Extracted_Response_VS_Huddle.fig

- L1_ISI_HAM6_GS13_H2_Pod_37_Extracted_Response_VS_Huddle.fig

- L1_ISI_HAM6_GS13_H3_Pod_84_Extracted_Response_VS_Huddle.fig

- L1_ISI_HAM6_GS13_V1_Pod_80_Extracted_Response_VS_Huddle.fig

- L1_ISI_HAM6_GS13_V2_Pod_64_Extracted_Response_VS_Huddle.fig

- L1_ISI_HAM6_GS13_V3_Pod_15_Extracted_Response_VS_Huddle.fig

Here are the plots for the 3 horizontal sensors:

Figure – Comparison between the responses of the horizontal GS-13s installed on HAM-ISI 6 and their respective huddle test results.

From those last 3  plots, we can notice that H1 and H3 resonant frequencies appear to be slightly under and over 1 Hz respectively, which may partly explain the slight difference between the horizontal corners. However, those slight differences were not visible on huddle test data. This may be due to a lower resolution on huddle test data

Conclusion of Side Chamber testing

When the actuator cables were properly isolated from the actuators (cf Step 20 linearity test comment), only a few minor issues were found during testing of this ISI:

- Slightly different actuator slope on H3 (in comparison to the other horizontal actuators)

- Slightly different response of the H1 and H3 GS-13s (in comparison to H2) between 0.8 and 1.1 Hz. This appears to be explained by the different sensors responses.

- The ISI mass load is inexplicably high compared to the other ISIs. We haven’t been able to find any mistakes in the weights recorded. (cf Step 7-Mass budget)

A number of mechanical tests (and adjustments) were waived at the time of side-chamber testing, leaving this for in-chamber initial testing:

- Level of the optical table

- Spring flatness measurements

- Horizontal actuator gaps