## Sheet1
Calculation of the DC path transimpedance based on design gain and measured output noise
Noise at 100Hz seen at differential DC output (dBVrms/rtHz) | 100Hz Noise translated into nV/rtHz | DC Path Calculated Shot Noise Limited Sensitivity assuming design whitening gain at 100Hz (mA)
-142 | 79.4328 | 3.4232
Calculation of the test path transconductance and the test switch isolation
Magnitude with switch on (dB) | Magnitude with switch off (dB | RF Transimpedance (ohms) | Test Path Transconductance (mA/V) | Test Switch Isolation (dB)
11.6 | -19.7 | 247 | 15.3923 | 31.3
This is Mike Z's method for shot noise limited sensitivity as outlined in T010121 section 3.3.1, it takes into account the instrument noise
Observed (dBm/rtHz) on RF Spectrum Analyzer with no light (Sdark) | RF Preamp Gain (dB) | Actual (dBm/rtHz) | Vrms/rtHz | V^2/Hz (Sdark)
-122 | 20.4 | -142.4 | 1.6962e-08 | 2.8772e-16
Observed (dBm/rtHz) on RF Spectrum Analyzer with 6mA to 10mA (Slamp) | RF Preamp Gain (dB) | Actual (dBm/rtHz) | Vrms/rtHz | V^2/Hz (Slamp) | Photocurrent (mA)
-114 | 20.4 | -134.4 | 4.2607e-08 | 1.8154e-15 | 10.03
Observed (dBm/rtHz) on RF Spectrum Analyzer with LNA terminated in 50 ohms (Sinst) | RF Preamp Gain (dB) | Actual (dBm/rtHz) | Vrms/rtHz | V^2/Hz (Sinst)
-143.2 | 20.4 | -163.6 | 1.4774e-09 | 2.1826e-18
V^2/Hz (S'dark) | V^2/Hz (S'lamp) | Shotnoise Limited Sensitivity (mA) | RF Transimpedance (ohms)
2.8554e-16 | 1.5277e-15 | 1.8747 | 689.9051