## Sequencers
Snumber | Location
S1202578 | L1-ISC-C3
S1202562 | L1-ISC-C4
S1202564 | L1-ISC-R1
S1202569 | L1-ISC-R2
S1202561 | L1-ISC-R3
S1202544 | L1-ISC-R4
S1202556 | L1-ISC-R5
S1202552 | L1-ISC-YR1
S1202547 | L1-ISC-YC1
S1202548 | L1-ISC-XC1
S1202568 | L1-ISC-XR1
S1202541 | LHO
S1202542 | LHO
S1202543 | LHO
S1202546 | LHO
S1202550 | LHO
S1202551 | LHO
S1202553 | LHO
S1202554 | LHO
S1202555 | LHO
S1202557 | LHO
S1202558 | LHO
S1202559 | LHO
S1202560 | LHO
S1202563 | LHO
S1202565 | LHO
S1202566 | LHO
S1202567 | LHO
S1202570 | LHO
S1202571 | LHO
S1202572 | LHO
S1202573 | LHO
S1202574 | LHO
S1202575 | LHO
S1202577 | LHO
S1202579 | LHO
S1202580 | LHO
S1202545 | LLO Spare - EE Lab
S1202549 | L1-TCS-R1
S1202576 | L1-TCS-R2
S1400672 | LLO Spare - CER
S1400673 | LLO Spare - CER
S1400674 | LLO Spare - CER
S1400675 | LLO Spare - CER
S1400676 | LHO Spare
S1400677 | LHO Spare
S1400678 | LHO Spare
S1400679 | LHO Spare
S1400680 | LHO Spare
S1400681 | LHO Spare

## Test Data
T1200379
D1200757 Power-Strip-Level DC Supply Sequencer Test Data
S# | Switch worked |  | LEDS lighted |  |  | Relays worked |  | Relay Turn-on +24VDC | Relay Turn-on -24VDC | ∆ +18 volts (Between Ports 1 and 2 on Relay 1) |  |  | ∆ -18 volts (Between Ports 1 and 2 on Relay 2)
|  | 24V In | 24V Out | 18V In | 18V Out | +18V | -18V |  |  | Volts | Amps | ∆ V | Volts | Amps | ∆ V
S1202541 | OK | ok | ok | ok | ok | ok | ok | 20.14 | 20.7 | 18 | 9.33 | 0.04 | 18 | 9.39 | 0.039
S1202542 | OK | ok | ok | ok | ok | ok | ok | 20.44 | 21.0 | 18 | 9.32 | 0.04 | 18 | 9.41 | 0.042
S1202543 | OK | ok | ok | ok | ok | ok | ok | 20.10 | 20.21 | 18 | 9.36 | 0.041 | 18 | 9.4 | 0.041
S1202544 | OK | ok | ok | ok | ok | ok | ok | 20.34 | 19.7 | 18 | 9.35 | 0.04 | 18 | 9.39 | 0.04
S1202545 | OK | ok | ok | ok | ok | ok | ok | 20.42 | 19.62 | 18 | 9.31 | 0.041 | 18 | 9.43 | 0.041
S1202546 | OK | ok | ok | ok | ok | ok | ok | 19.45 | 21.32 | 18 | 9.3 | 0.041 | 18 | 9.42 | 0.042
S1202547 | OK | ok | ok | ok | ok | ok | ok | 20.26 | 19.99 | 18 | 9.23 | 0.042 | 18 | 9.48 | 0.042
S1202548 | OK | ok | ok | ok | ok | ok | ok | 20.25 | 20.13 | 18 | 9.22 | 0.041 | 18 | 9.49 | 0.042
S1202549 | OK | ok | ok | ok | ok | ok | ok | 19.12 | 19.44 | 18 | 9.19 | 0.048 | 18 | 9.48 | 0.041
S1202550 | OK | ok | ok | ok | ok | ok | ok | 20.89 | 20.4 | 18 | 9.27 | 0.042 | 18 | 9.46 | 0.04
S1202551 | OK | ok | ok | ok | ok | ok | ok | 19.74 | 20.97 | 18 | 9.26 | 0.041 | 18 | 9.45 | 0.042
S1202552 | OK | ok | ok | ok | ok | ok | ok | 20.00 | 19.31 | 18 | 9.27 | 0.041 | 18 | 9.46 | 0.04
S1202553 | OK | ok | ok | ok | ok | ok | ok | 21.40 | 20.18 | 18 | 9.33 | 0.041 | 18 | 9.45 | 0.04
S1202554 | OK | ok | ok | ok | ok | ok | ok | 20.19 | 19.83 | 18 | 9.24 | 0.041 | 18 | 9.39 | 0.042
S1202555 | OK | ok | ok | ok | ok | ok | ok | 20.06 | 19.31 | 18 | 9.24 | 0.042 | 18 | 9.42 | 0.044
S1202556 | OK | ok | ok | ok | ok | ok | ok | 20.07 | 20.93 | 18 | 9.24 | 0.04 | 18 | 9.45 | 0.042
S1202557 | OK | ok | ok | ok | ok | ok | ok | 20.78 | 20.52 | 18 | 9.33 | 0.041 | 18 | 9.43 | 0.041
S1202558 | OK | ok | ok | ok | ok | ok | ok | 19.91 | 20.47 | 18 | 9.34 | 0.042 | 18 | 9.41 | 0.041
S1202559 | OK | ok | ok | ok | ok | ok | ok | 20.90 | 19.96 | 18 | 9.3 | 0.04 | 18 | 9.4 | 0.041
S1202560 | OK | ok | ok | ok | ok | ok | ok | 19.85 | 21.05 | 18 | 9.24 | 0.042 | 18 | 9.45 | 0.05
S1202561 | OK | ok | ok | ok | ok | ok | ok | 19.62 | 20.82 | 18 | 9.21 | 0.042 | 18 | 9.45 | 0.04
S1202562 | OK | ok | ok | ok | ok | ok | ok | 20.10 | 19.8 | 18 | 9.33 | 0.041 | 18 | 9.47 | 0.04
S1202563 | OK | ok | ok | ok | ok | ok | ok | 20.40 | 19.24 | 18 | 9.28 | 0.041 | 18 | 9.4 | 0.042
S1202564 | OK | ok | ok | ok | ok | ok | ok | 20.19 | 20.51 | 18 | 9.35 | 0.042 | 18 | 9.46 | 0.041
S1202565 | OK | ok | ok | ok | ok | ok | ok | 20.41 | 19.80 | 18 | 9.24 | 0.041 | 18 | 9.44 | 0.04
S1202566 | OK | ok | ok | ok | ok | ok | ok | 20.68 | 20.33 | 18 | 9.34 | 0.039 | 18 | 9.42 | 0.041
S1202567 | OK | ok | ok | ok | ok | ok | ok | 19.61 | 19.51 | 18 | 9.3 | 0.058 | 18 | 9.48 | 0.047
S1202568 | OK | ok | ok | ok | ok | ok | ok | 20.23 | 19.71 | 18 | 9.25 | 0.042 | 18 | 9.48 | 0.042
S1202569 | OK | ok | ok | ok | ok | ok | ok | 20.50 | 19.5 | 18 | 9.37 | 0.04 | 18 | 9.39 | 0.041
S1202570 | OK | ok | ok | ok | ok | ok | ok | 20.65 | 20.77 | 18 | 9.19 | 0.043 | 18 | 9.49 | 0.041
S1202571 | OK | ok | ok | ok | ok | ok | ok | 20.36 | 20.22 | 18 | 9.2 | 0.041 | 18 | 9.49 | 0.04
S1202572 | OK | ok | ok | ok | ok | ok | ok | 20.00 | 20.25 | 18 | 9.3 | 0.041 | 18 | 9.43 | 0.042
S1202573 | OK | ok | ok | ok | ok | ok | ok | 19.52 | 21.20 | 18 | 9.26 | 0.046 | 18 | 9.44 | 0.042
S1202574 | OK | ok | ok | ok | ok | ok | ok | 20.00 | 20.24 | 18 | 9.3 | 0.042 | 18 | 9.4 | 0.045
S1202575 | OK | ok | ok | ok | ok | ok | ok | 19.04 | 20.14 | 18 | 9.23 | 0.044 | 18 | 9.45 | 0.044
S1202576 | OK | ok | ok | ok | ok | ok | ok | 19.60 | 20.0 | 18 | 9.37 | 0.042 | 18 | 9.4 | 0.041
S1202577 | OK | ok | ok | ok | ok | ok | ok | 20.21 | 20.56 | 18 | 9.27 | 0.039 | 18 | 9.44 | 0.041
S1202578 | OK | ok | ok | ok | ok | ok | ok | 19.7 | 20.3 | 18 | 9.33 | 0.042 | 18 | 9.46 | 0.042
S1202579 | OK | ok | ok | ok | ok | ok | ok | 19.99 | 20.46 | 18 | 9.2 | 0.041 | 18 | 9.5 | 0.041
S1202580 | OK | ok | ok | ok | ok | ok | ok | 21.10 | 21.38 | 18 | 9.35 | 0.041 | 18 | 9.4 | 0.041
S1400672 | OK | ok | ok | ok | ok | ok | ok | 20.7 | 20.4 | 18 | 9.14 | 0.039 | 18 | 9.06 | 0.039
S1400673 | OK | ok | ok | ok | ok | ok | ok | 20.1 | 20.8 | 18 | 9.14 | 0.039 | 18 | 9.08 | 0.039
S1400674 | OK | ok | ok | ok | ok | ok | ok | 20.5 | 19.8 | 18 | 9.13 | 0.039 | 18 | 9.12 | 0.039
S1400675 | OK | ok | ok | ok | ok | ok | ok | 20.1 | 20.3 | 18 | 9.13 | 0.04 | 18 | 9.11 | 0.04
S1400676 | OK | ok | ok | ok | ok | ok | ok | 20.8 | 19.8 | 18 | 9.15 | 0.039 | 18 | 9.2 | 0.04
S1400677 | OK | ok | ok | ok | ok | ok | ok | 20.5 | 20.0 | 18 | 9.1 | 0.038 | 18 | 9.1 | 0.04
S1400678 | OK | ok | ok | ok | ok | ok | ok | 20.3 | 20.4 | 18 | 9.1 | 0.039 | 18 | 9.1 | 0.04
S1400679 | OK | ok | ok | ok | ok | ok | ok | 19.6 | 20.7 | 18 | 9.15 | 0.04 | 18 | 9.1 | 0.04
S1400680 | OK | ok | ok | ok | ok | ok | ok | 21.2 | 20.4 | 18 | 9.09 | 0.04 | 18 | 9.1 | 0.04
S1400681 | OK | ok | ok | ok | ok | ok | ok | 20.8 | 19.5 | 18 | 9.13 | 0.039 | 18 | 9.1 | 0.039

## D1200757 BOM
2014-02-19
|  | D1200757 Power-Strip_level DC Power Sequencer BOM | [7 empty] | Estimated | Quantity | 10 | Estimated
LIGO Part Number | Item | Part Type | Description | Designator | Manufacturer | Supplier | Supplier Part Number | Quantity | Unit | Unit Cost | Extended | Total count | Total cost
D1200757 | 1 | Sequencer Box | Sequencer Box |  | Hamilton Metalcraft, Inc. | Hamilton | D1200757 | 1 | @ | 78 | 78 | 10 | 780
D1200757 | 2 | Relay | Solid State Relays .005-100A 3.5-32VDC |  | Crydom | Mouser Electronics | 558-D06D100 | 2 | @ | 121.87 | 243.74 | 20 | 2437.4
D1200757 | 3 | Switch | Rocker Switches & Paddle Switches 3PST ON-NONE-OFF 30A |  | NKK Switches | Mouser Electronics | 633-SW3831-RO | 2 | @ | 33.86 | 67.72 | 20 | 677.2
D1200757 | 4 | Diode | DIODE ZENER 18V 2W DO-41 |  | Microsemi Commercial Components | Digi-Key | 2EZ18D5MSCT-ND | 2 | @ | 2.86 | 5.72 | 20 | 57.2
D1200757 | 5 | Diode | Current Regulator Diode 4.7mA |  | Central Semiconductor | Mouser Electronics | 1N5314 | 4 | @ | 4.79 | 19.16 | 40 | 191.6
D1200757 | 6 | LED | Green LED, Leaded, Panel mount |  | Chicago Miniature Lighting | Digi-Key | L10005-ND | 4 | @ | 0.69 | 2.76 | 40 | 27.6
D1200757 | 7 | Connector Shell | 3W3 Female Combo Shell - No Power |  | Conec | Conec | 3003W3SXX99A30X | 1 | @ | 2.44 | 2.44 | 10 | 24.4
D1200757 | 8 | Connector Shell | 3W3 Male Combo Shell - No Power |  | Conec | Conec | 3003W3PXX99A30X | 1 | @ | 2.44 | 2.44 | 10 | 24.4
D1200757 | 9 | Connector Shell | 3W3C Female Combo Shell - No Power |  | Conec | Conec | 303W3CSXX99A30X | 1 | @ | 2.48 | 2.48 | 10 | 24.8
D1200757 | 10 | Connector Shell | 3W3C Male Combo Shell - No Power |  | Conec | Conec | 303W3CPXX99A30X | 1 | @ | 2.48 | 2.48 | 10 | 24.8
D1200757 | 11 | Connector Contact | 20 AMP Male Power Contact |  | Conec | Conec | 131C11029X | 6 | @ | 1 | 6 | 60 | 60
D1200757 | 12 | Connector Contact | 20 AMP Female Power Contact |  | Conec | Conec | 132C11029X | 6 | @ | 1 | 6 | 60 | 60
D1200757 | 13 | Jack Socket | Jack Socket Kit |  | 3M | Digi-Key | MDVS44-ND | 4 | Kit | 2.4 | 9.6 | 40 | 96
D1200757 | 14 | Screw | 18-8 Stainless Steel Flat Undercut Head Phillips Machine Screw, 8-32 Thread, 5/8" Length |  |  | McMaster-Carr | 91099A266 | 4 | @ | 0.0772 | 0.3088 | 40 | 3.088
D1200757 | 15 | Nut | 18-8 Stainless Steel Nylon-Insert Hex Locknut, 8-32 Thread Size, 11/32" Width, 15/64" Height |  |  | McMaster-Carr | 91831A009 | 4 | @ | 0.0592 | 0.2368 | 40 | 2.368
D1200757 | 16 | Spade Terminal | Crimp-on Spade Terminal Block, Vinyl Insulated, 22-18 AWG, #10 Stud |  |  | McMaster-Carr | 69145K69 | 4 | @ | 0.1178 | 0.4712 | 40 | 4.712
D1200757 | 17 | Spade Terminal | Crimp-on Spade Terminal Block, Vinyl Insulated, 12-10 AWG, #10 Stud |  |  | McMaster-Carr | 69145K78 | 8 | @ | 0.0803 | 0.6424 | 80 | 6.424
D1200757 | 18 | Wire Connector | Standard Twist-on Wire Connector, 2(18 AWG) min, 4(14 AWG) & 1(18 AWG) Maximum, Yellow |  |  | McMaster-Carr | 7108K6 | 2 | @ | 0.1022 | 0.2044 | 20 | 2.044
D1200757 | 19 | Butt Splice | Standard Twist-on Wire Connector, 2(18 AWG) min, 4(14 AWG) & 1(18 AWG) Maximum, Yellow |  |  | McMaster-Carr | 7227K12 | 1 | @ | 0.1228 | 0.1228 | 10 | 1.228
D1200757 |  | Total Material | [5 empty] | 1 | Kit |  | 450.5264 | 10 | 4505.264
D1200757 |  | Labor | [5 empty] | 1 | Asmb | 90 | 90 | 10 | 900
D1200757 |  | Total Cost | [5 empty] | 1 | Cost |  | 540.5264 | 10 | 5405.264

## D1200757 Buy - 20140219
2014-02-19
|  | D1200757 Power-Strip-Level DC Power Sequencer BOM | [7 empty] | Estimated | Quantity | 10 | Estimated | Ordered?
LIGO Part Number | Item | Part Type | Description | Designator | Manufacturer | Supplier | Supplier Part Number | Quantity | Unit | Unit Cost | Extended | Total count | Total cost
D1200757 | 1 | Sequencer Box | Sequencer Box |  | Hamilton Metalcraft, Inc. | Hamilton | D1200757 | 1 | @ | 125 | 125 | 10 | 1250 | 2014-02-26
D1200757 | 2 | Relay | Solid State Relays .005-100A 3.5-32VDC |  | Crydom | Mouser Electronics | 558-D06D100 | 2 | @ | 121.87 | 243.74 | 20 | 2437.4 | 2014-02-25
D1200757 | 3 | Switch | Rocker Switches & Paddle Switches 3PST ON-NONE-OFF 30A |  | NKK Switches | Mouser Electronics | 633-SW3831-RO | 2 | @ | 33.86 | 67.72 | 20 | 677.2 | On Hand
D1200757 | 4 | Diode | DIODE ZENER 18V 2W DO-41 |  | Microsemi Commercial Components | Digi-Key | 2EZ18D5MSCT-ND | 2 | @ | 2.86 | 5.72 | 20 | 57.2 | 2014-02-25
D1200757 | 5 | Diode | Current Regulator Diode 4.7mA |  | Central Semiconductor | Mouser Electronics | 1N5314 | 4 | @ | 4.79 | 19.16 | 40 | 191.6 | On Hand
D1200757 | 6 | LED | Green LED, Leaded, Panel mount |  | Chicago Miniature Lighting | Digi-Key | L10005-ND | 4 | @ | 0.69 | 2.76 | 40 | 27.6 | On Hand
D1200757 | 7 | Connector Shell | 3W3 Female Combo Shell - No Power |  | Conec | Conec | 3003W3SXX99A30X | 1 | @ | 2.21 | 2.21 | 10 | 22.1 | On Hand
D1200757 | 8 | Connector Shell | 3W3 Male Combo Shell - No Power |  | Conec | Conec | 3003W3PXX99A30X | 1 | @ | 2.19 | 2.19 | 10 | 21.9 | On Hand
D1200757 | 9 | Connector Shell | 3W3C Female Combo Shell - No Power |  | Conec | Conec | 303W3CSXX99A30X | 1 | @ | 2.13 | 2.13 | 10 | 21.3 | On Hand
D1200757 | 10 | Connector Shell | 3W3C Male Combo Shell - No Power |  | Conec | Conec | 303W3CPXX99A30X | 1 | @ | 2.2 | 2.2 | 10 | 22 | On Hand
D1200757 | 11 | Connector Contact | 20 AMP Male Power Contact |  | Conec | Conec | 131C11029X | 6 | @ | 1.35 | 8.1 | 60 | 81 | On Hand
D1200757 | 12 | Connector Contact | 20 AMP Female Power Contact |  | Conec | Conec | 132C11029X | 6 | @ | 1.69 | 10.14 | 60 | 101.4 | On Hand
D1200757 | 13 | Jack Socket | Jack Socket Kit |  | 3M | Digi-Key | MDVS44-ND | 4 | Kit | 2.4 | 9.6 | 40 | 96 | On Hand
D1200757 | 14 | Screw | 18-8 Stainless Steel Flat Undercut Head Phillips Machine Screw, 8-32 Thread, 5/8" Length |  |  | McMaster-Carr | 91099A266 | 4 | @ | 0.0772 | 0.3088 | 40 | 3.088 | 2014-02-26
D1200757 | 15 | Nut | 18-8 Stainless Steel Nylon-Insert Hex Locknut, 8-32 Thread Size, 11/32" Width, 15/64" Height |  |  | McMaster-Carr | 91831A009 | 4 | @ | 0.0592 | 0.2368 | 40 | 2.368 | 2014-02-26
D1200757 | 16 | Spade Terminal | Crimp-on Spade Terminal Block, Vinyl Insulated, 22-18 AWG, #10 Stud |  |  | McMaster-Carr | 69145K69 | 4 | @ | 0.1178 | 0.4712 | 40 | 4.712 | On Hand
D1200757 | 17 | Spade Terminal | Crimp-on Spade Terminal Block, Vinyl Insulated, 12-10 AWG, #10 Stud |  |  | McMaster-Carr | 69145K78 | 8 | @ | 0.0803 | 0.6424 | 80 | 6.424 | On Hand
D1200757 | 18 | Wire Connector | Standard Twist-on Wire Connector, 2(18 AWG) min, 4(14 AWG) & 1(18 AWG) Maximum, Yellow |  |  | McMaster-Carr | 7108K6 | 2 | @ | 0.1022 | 0.2044 | 20 | 2.044 | On Hand
D1200757 | 19 | Butt Splice | Standard Twist-on Wire Connector, 2(18 AWG) min, 4(14 AWG) & 1(18 AWG) Maximum, Yellow |  |  | McMaster-Carr | 7227K12 | 1 | @ | 0.1228 | 0.1228 | 10 | 1.228 | On Hand
D1200757 |  | Total Material | [5 empty] | 1 | Kit |  | 502.6564 | 10 | 5026.564
D1200757 |  | Labor | [5 empty] | 1 | Asmb | 90 | 90 | 10 | 900
D1200757 |  | Total Cost | [5 empty] | 1 | Cost |  | 592.6564 | 10 | 5926.564