## End Remote
Ifo | Chassis | Rail | # | Terminal | Channel | # | Name | PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len
H1 | End Remote | L | 2 | EL3202 | RTD Inputs Channel | 1 | Value | PLC1 | Standard | Inputs | WORD | PemEndTempCerRack1DegcIn
H1 | End Remote | L | 2 | EL3202 | RTD Inputs Channel | 2 | Value | PLC1 | Standard | Inputs | WORD | PemEndTempSupRack1DegcIn
H1 | End Remote | R | 1 | EL6692 | IO Inputs | — | PLC1toEndY | PLC2 | Standard | Inputs | STRUCT | RecieveFromCornerPLC1 | —
H1 | End Remote | R | 1 | EL6692 | IO Inputs | — | PLC2toEndY | PLC2 | Standard | Inputs | STRUCT | RecieveFromCornerPLC2 | —
H1 | End Remote | R | 1 | EL6692 | IO Outputs | — | EndYtoPLC2 | PLC2 | Standard | Inputs | STRUCT | SendToCornerPLC2 | —

## ISC Common
Ifo | Chassis | Rail | # | Terminal | Channel | # | Name | PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len
H1 | ISC Common | L | 1 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndReflServoIn | SumMon
H1 | ISC Common | L | 1 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndReflServoIn | SplitMon
H1 | ISC Common | L | 2 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndReflServoIn | FastMon
H1 | ISC Common | L | 2 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndReflServoIn | SlowfbMon
H1 | ISC Common | L | 3 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndReflServoIn | SlowMon
H1 | ISC Common | L | 4 | El4132 | Channel | 1 | Output | PLC2 | Standard | Outputs | WORD | AlsEndReflServoOut | ComOfs
H1 | ISC Common | L | 4 | El4132 | Channel | 2 | Output | PLC2 | Standard | Outputs | WORD | AlsEndReflServoOut | SlowOfs
H1 | ISC Common | L | 5 | El4132 | Channel | 1 | Output | PLC2 | Standard | Outputs | WORD | AlsEndReflServoOut | SlowoutOfs
H1 | ISC Common | L | 7 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndReflServoIn | FastLimit
H1 | ISC Common | L | 7 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndReflServoIn | Ok
H1 | ISC Common | L | 8 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 0 | 0 | 1
H1 | ISC Common | L | 8 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 1 | 0 | 1
H1 | ISC Common | L | 8 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 2 | 0 | 1
H1 | ISC Common | L | 8 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 3 | 0 | 1
H1 | ISC Common | L | 9 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 4 | 0 | 1
H1 | ISC Common | L | 9 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastGain | 5 | 0 | 1
H1 | ISC Common | L | 9 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastExcEn
H1 | ISC Common | L | 9 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastOpt
H1 | ISC Common | L | 10 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | LatchEn
H1 | ISC Common | L | 10 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowExcEn
H1 | ISC Common | L | 10 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Pol
H1 | ISC Common | L | 10 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Pol
H1 | ISC Common | L | 11 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 0 | 0 | 1
H1 | ISC Common | L | 11 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 1 | 0 | 1
H1 | ISC Common | L | 11 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 2 | 0 | 1
H1 | ISC Common | L | 11 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 3 | 0 | 1
H1 | ISC Common | L | 12 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 4 | 0 | 1
H1 | ISC Common | L | 12 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1Gain | 5 | 0 | 1
H1 | ISC Common | L | 12 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 0 | 0 | 1
H1 | ISC Common | L | 12 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 1 | 0 | 1
H1 | ISC Common | L | 13 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 2 | 0 | 1
H1 | ISC Common | L | 13 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 3 | 0 | 1
H1 | ISC Common | L | 13 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 4 | 0 | 1
H1 | ISC Common | L | 13 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2Gain | 5 | 0 | 1
H1 | ISC Common | L | 15 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In1En
H1 | ISC Common | L | 15 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | In2En
H1 | ISC Common | L | 15 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | OutSw
H1 | ISC Common | L | 15 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComBoost | 0 | 0 | 1
H1 | ISC Common | L | 16 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComBoost | 1 | 0 | 1
H1 | ISC Common | L | 16 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComComp
H1 | ISC Common | L | 16 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComExcEn
H1 | ISC Common | L | 16 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComOpt
H1 | ISC Common | L | 17 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowPol
H1 | ISC Common | L | 17 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | ComFilter
H1 | ISC Common | L | 17 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastEn
H1 | ISC Common | L | 17 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastPol
H1 | ISC Common | L | 18 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowOpt
H1 | ISC Common | L | 18 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowBypass
H1 | ISC Common | L | 18 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowOfs5V
H1 | ISC Common | L | 18 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowOfsEn
H1 | ISC Common | L | 19 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowComp
H1 | ISC Common | L | 19 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowBoost
H1 | ISC Common | L | 19 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | SlowFilter
H1 | ISC Common | L | 19 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndReflServoOut | FastLimiter
H1 | ISC Common | M | 1 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrServoIn | SumMon
H1 | ISC Common | M | 1 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrServoIn | SplitMon
H1 | ISC Common | M | 2 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrServoIn | FastMon
H1 | ISC Common | M | 2 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrServoIn | SlowfbMon
H1 | ISC Common | M | 3 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrServoIn | SlowMon
H1 | ISC Common | M | 4 | El4132 | Channel | 1 | Output | PLC2 | Standard | Outputs | WORD | AlsEndFibrServoOut | ComOfs
H1 | ISC Common | M | 4 | El4132 | Channel | 2 | Output | PLC2 | Standard | Outputs | WORD | AlsEndFibrServoOut | SlowOfs
H1 | ISC Common | M | 5 | El4132 | Channel | 1 | Output | PLC2 | Standard | Outputs | WORD | AlsEndFibrServoOut | SlowoutOfs
H1 | ISC Common | M | 7 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndFibrServoIn | FastLimit
H1 | ISC Common | M | 7 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndFibrServoIn | Ok
H1 | ISC Common | M | 8 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 0 | 0 | 1
H1 | ISC Common | M | 8 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 1 | 0 | 1
H1 | ISC Common | M | 8 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 2 | 0 | 1
H1 | ISC Common | M | 8 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 3 | 0 | 1
H1 | ISC Common | M | 9 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 4 | 0 | 1
H1 | ISC Common | M | 9 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastGain | 5 | 0 | 1
H1 | ISC Common | M | 9 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastExcEn
H1 | ISC Common | M | 9 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastOpt
H1 | ISC Common | M | 10 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | LatchEn
H1 | ISC Common | M | 10 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowExcEn
H1 | ISC Common | M | 10 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Pol
H1 | ISC Common | M | 10 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Pol
H1 | ISC Common | M | 11 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 0 | 0 | 1
H1 | ISC Common | M | 11 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 1 | 0 | 1
H1 | ISC Common | M | 11 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 2 | 0 | 1
H1 | ISC Common | M | 11 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 3 | 0 | 1
H1 | ISC Common | M | 12 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 4 | 0 | 1
H1 | ISC Common | M | 12 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1Gain | 5 | 0 | 1
H1 | ISC Common | M | 12 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 0 | 0 | 1
H1 | ISC Common | M | 12 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 1 | 0 | 1
H1 | ISC Common | M | 13 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 2 | 0 | 1
H1 | ISC Common | M | 13 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 3 | 0 | 1
H1 | ISC Common | M | 13 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 4 | 0 | 1
H1 | ISC Common | M | 13 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2Gain | 5 | 0 | 1
H1 | ISC Common | M | 15 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In1En
H1 | ISC Common | M | 15 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | In2En
H1 | ISC Common | M | 15 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | OutSw
H1 | ISC Common | M | 15 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComBoost | 0 | 0 | 1
H1 | ISC Common | M | 16 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComBoost | 1 | 0 | 1
H1 | ISC Common | M | 16 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComComp
H1 | ISC Common | M | 16 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComExcEn
H1 | ISC Common | M | 16 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComOpt
H1 | ISC Common | M | 17 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowPol
H1 | ISC Common | M | 17 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | ComFilter
H1 | ISC Common | M | 17 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastEn
H1 | ISC Common | M | 17 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastPol
H1 | ISC Common | M | 18 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowOpt
H1 | ISC Common | M | 18 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowBypass
H1 | ISC Common | M | 18 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowOfs5V
H1 | ISC Common | M | 18 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowOfsEn
H1 | ISC Common | M | 19 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowComp
H1 | ISC Common | M | 19 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowBoost
H1 | ISC Common | M | 19 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | SlowFilter
H1 | ISC Common | M | 19 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrServoOut | FastLimiter
H1 | ISC Common | R | 1 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | ReferenceMon
H1 | ISC Common | R | 1 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | DividerMon
H1 | ISC Common | R | 2 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | OutputMon
H1 | ISC Common | R | 2 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | ReferenceTemp
H1 | ISC Common | R | 3 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | DividerTemp
H1 | ISC Common | R | 3 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | OutputTemp
H1 | ISC Common | R | 4 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndVcoIn | TuneMon
H1 | ISC Common | R | 5 | El4132 | Channel | 1 | Output | PLC2 | Standard | Outputs | WORD | AlsEndVcoOut | TuneOfs
H1 | ISC Common | R | 7 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndVcoIn | ExcitationSwitch
H1 | ISC Common | R | 7 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndVcoIn | PowerOk
H1 | ISC Common | R | 8 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndVcoOut | ExcitationEn
H1 | ISC Common | R | 10 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 0 | 0 | 1
H1 | ISC Common | R | 10 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 1 | 0 | 1
H1 | ISC Common | R | 10 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 2 | 0 | 1
H1 | ISC Common | R | 10 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 3 | 0 | 1
H1 | ISC Common | R | 11 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 4 | 0 | 1
H1 | ISC Common | R | 11 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 5 | 0 | 1
H1 | ISC Common | R | 11 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 6 | 0 | 1
H1 | ISC Common | R | 11 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 7 | 0 | 1
H1 | ISC Common | R | 12 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | DelayStep | 8 | 0 | 1
H1 | ISC Common | R | 12 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_APhaseOut | LatchEn
H1 | ISC Common | R | 12 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 0 | 0 | 1
H1 | ISC Common | R | 12 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 1 | 0 | 1
H1 | ISC Common | R | 13 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 2 | 0 | 1
H1 | ISC Common | R | 13 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 3 | 0 | 1
H1 | ISC Common | R | 13 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 4 | 0 | 1
H1 | ISC Common | R | 13 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 5 | 0 | 1
H1 | ISC Common | R | 14 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 6 | 0 | 1
H1 | ISC Common | R | 14 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 7 | 0 | 1
H1 | ISC Common | R | 14 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | DelayStep | 8 | 0 | 1
H1 | ISC Common | R | 14 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_APhaseOut | LatchEn

## End 2
Ifo | Chassis | Rail | # | Terminal | Channel | # | Name | PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len
H1 | End 2 | L | 1 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | LaserDiode1PowerMonitor
H1 | End 2 | L | 1 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | LaserDiode2PowerMonitor
H1 | End 2 | L | 1 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | LaserCrystalTECErrorSignal
H1 | End 2 | L | 1 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | DoublingCrystalTECErrorSignal
H1 | End 2 | L | 2 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | LaserDiode1TECErrorSignal
H1 | End 2 | L | 2 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | LaserDiode2TECErrorSignal
H1 | End 2 | L | 2 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserHeadIn | NoiseEaterMonitor
H1 | End 2 | L | 2 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD
H1 | End 2 | L | 3 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndLaserHeadIn | LaserDiode1TempGuard
H1 | End 2 | L | 3 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndLaserHeadIn | LaserDiode2TempGuard
H1 | End 2 | L | 3 | EL1124 | Channel | 3 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndLaserHeadIn | Interlock
H1 | End 2 | L | 3 | EL1124 | Channel | 4 | Input | PLC2 | Standard | Inputs | BOOL
H1 | End 2 | L | 5 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFiber_ADemodIn | RFMon
H1 | End 2 | L | 5 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFiber_ADemodIn | LOMon
H1 | End 2 | L | 5 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_ADemodIn | RFMon
H1 | End 2 | L | 5 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_ADemodIn | LOMon
H1 | End 2 | L | 6 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndRefl_ADemodIn | RFMon
H1 | End 2 | L | 6 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndRefl_ADemodIn | LOMon
H1 | End 2 | L | 6 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_BDemodIn | RFMon
H1 | End 2 | L | 6 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_BDemodIn | LOMon
H1 | End 2 | L | 7 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndFiber_ADemodIn | PowerOk
H1 | End 2 | L | 7 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndFiber_ADemodIn | Sign
H1 | End 2 | L | 7 | EL1124 | Channel | 3 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndSpare_ADemodIn | Sign
H1 | End 2 | L | 7 | EL1124 | Channel | 4 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndRefl_ADemodIn | PowerOk
H1 | End 2 | L | 9 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | IscRfEndAmp71mIn | OutputMon
H1 | End 2 | L | 9 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | IscRfEndAmp21m4In | OutputMon
H1 | End 2 | L | 9 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | IscRfEndAlsDiv39m9In | OutputMon
H1 | End 2 | L | 9 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | IscRfEndSpareIn | OutputMon
H1 | End 2 | L | 10 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | IscRfEndAmp71mIn | PowerOk
H1 | End 2 | L | 10 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | IscRfEndAmp21m4In | PowerOk
H1 | End 2 | L | 10 | EL1124 | Channel | 3 | Input | PLC2 | Standard | Inputs | BOOL | IscRfEndAlsDiv39m9In | PowerOk
H1 | End 2 | L | 10 | EL1124 | Channel | 4 | Input | PLC2 | Standard | Inputs | BOOL | IscRfEndSpareIn | PowerOk
H1 | End 2 | L | 12 | EL3104 | AI Standard Channel | 1 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | UpperVoltage
H1 | End 2 | L | 12 | EL3104 | AI Standard Channel | 2 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | UpperRtd
H1 | End 2 | L | 12 | EL3104 | AI Standard Channel | 3 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | UpperPcb
H1 | End 2 | L | 12 | EL3104 | AI Standard Channel | 4 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | UpperCurrent
H1 | End 2 | L | 13 | EL3104 | AI Standard Channel | 1 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | LowerVoltage
H1 | End 2 | L | 13 | EL3104 | AI Standard Channel | 2 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | LowerRtd
H1 | End 2 | L | 13 | EL3104 | AI Standard Channel | 3 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | LowerPcb
H1 | End 2 | L | 13 | EL3104 | AI Standard Channel | 4 | Value | PLC3 | Standard | Inputs | WORD | TcsRHEIn | LowerCurrent
H1 | End 2 | L | 14 | EL4134 | AI Standard Channel | 1 | Value | PLC3 | Standard | Outputs | WORD | TcsRHEOut | SetUpperDriveCurrent
H1 | End 2 | L | 14 | EL4134 | AI Standard Channel | 2 | Value | PLC3 | Standard | Outputs | WORD | TcsRHEOut | SetUpperDriveAwg
H1 | End 2 | L | 14 | EL4134 | AI Standard Channel | 3 | Value | PLC3 | Standard | Outputs | WORD | TcsRHEOut | SetLowerDriveCurrent
H1 | End 2 | L | 14 | EL4134 | AI Standard Channel | 4 | Value | PLC3 | Standard | Outputs | WORD | TcsRHEOut | SetLowerDriveAwg
H1 | End 2 | L | 15 | EL1124 | Channel | 1 | Input | PLC3 | Standard | Inputs | BOOL | TcsHWSEOut | DalsaCameraSwitch
H1 | End 2 | L | 15 | EL1124 | Channel | 2 | Input | PLC3 | Standard | Inputs | BOOL | TcsHWSEOut | RcxClinkSwitch
H1 | End 2 | L | 15 | EL1124 | Channel | 3 | Input | PLC3 | Standard | Inputs | BOOL
H1 | End 2 | L | 15 | EL1124 | Channel | 4 | Input | PLC3 | Standard | Inputs | BOOL
H1 | End 2 | M | 1 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt2In | HV10MonitorPZTX
H1 | End 2 | M | 1 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt2In | PositionSensorPZTX
H1 | End 2 | M | 1 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt2In | HV10MonitorPZTY
H1 | End 2 | M | 1 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt2In | PositionSensorPZTY
H1 | End 2 | M | 2 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt1In | HV10MonitorPZTX
H1 | End 2 | M | 2 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt1In | PositionSensorPZTX
H1 | End 2 | M | 2 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt1In | HV10MonitorPZTY
H1 | End 2 | M | 2 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndPzt1In | PositionSensorPZTY
H1 | End 2 | M | 3 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndRefl_BDcIn | DCPower
H1 | End 2 | M | 3 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | LscEndTr_ADcIn | DCPower
H1 | End 2 | M | 3 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserGRDcIn | DCPower
H1 | End 2 | M | 3 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFiber_ADcIn | DCPower
H1 | End 2 | M | 4 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_CDcIn | DCPower
H1 | End 2 | M | 4 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_DDcIn | DCPower
H1 | End 2 | M | 4 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_EDcIn | DCPower
H1 | End 2 | M | 4 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_FDcIn | DCPower
H1 | End 2 | M | 5 | EL4132 | AI Standard Channel | 1 | Value | PLC2 | Standard | Outputs | WORD | AlsEndLaserOut | CrystalTemperature
H1 | End 2 | M | 5 | EL4132 | AI Standard Channel | 2 | Value | PLC2 | Standard | Outputs | WORD | AlsEndLaserOut | DoublerTemperature
H1 | End 2 | M | 7 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | IscEndExtraIn | BI1
H1 | End 2 | M | 7 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL | IscEndExtraIn | BI2
H1 | End 2 | M | 7 | EL1124 | Channel | 3 | Input | PLC2 | Standard | Inputs | BOOL | IscEndExtraIn | BI3
H1 | End 2 | M | 7 | EL1124 | Channel | 4 | Input | PLC2 | Standard | Inputs | BOOL | IscEndExtraIn | BI4
H1 | End 2 | M | 8 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | IscEndExtraOut | BO1
H1 | End 2 | M | 8 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | IscEndExtraOut | BO2
H1 | End 2 | M | 8 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | IscEndExtraOut | BO3
H1 | End 2 | M | 8 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | IscEndExtraOut | BO4
H1 | End 2 | M | 9 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | IscEndExtraIn | AI1
H1 | End 2 | M | 9 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | IscEndExtraIn | AI2
H1 | End 2 | M | 9 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | IscEndExtraIn | AI3
H1 | End 2 | M | 9 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | ReceiveFromRt | I[1]
H1 | End 2 | M | 10 | EL4134 | AI Standard Channel | 1 | Value | PLC2 | Standard | Outputs | WORD | IscEndExtraOut | AO1
H1 | End 2 | M | 10 | EL4134 | AI Standard Channel | 2 | Value | PLC2 | Standard | Outputs | WORD | IscEndExtraOut | AO2
H1 | End 2 | M | 10 | EL4134 | AI Standard Channel | 3 | Value | PLC2 | Standard | Outputs | WORD | IscEndExtraOut | AO3
H1 | End 2 | M | 10 | EL4134 | AI Standard Channel | 4 | Value | PLC2 | Standard | Outputs | WORD | SendToRt | O[1]
H1 | End 2 | M | 12 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndLaserIRDcIn | Gain | 0 | 0 | 1
H1 | End 2 | M | 12 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndLaserIRDcIn | Gain | 1 | 0 | 1
H1 | End 2 | M | 12 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrRejectedDcIn | Gain | 0 | 0 | 1
H1 | End 2 | M | 12 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrRejectedDcIn | Gain | 1 | 0 | 1
H1 | End 2 | M | 13 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrTransDcIn | Gain | 0 | 0 | 1
H1 | End 2 | M | 13 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndFibrTransDcIn | Gain | 1 | 0 | 1
H1 | End 2 | M | 13 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_BDcIn | Gain | 0 | 0 | 1
H1 | End 2 | M | 13 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndSpare_BDcIn | Gain | 1 | 0 | 1
H1 | End 2 | M | 14 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndLaserIRDcIn | DCPower
H1 | End 2 | M | 14 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrRejectedDcIn | DCPower
H1 | End 2 | M | 14 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndFibrTransDcIn | DCPower
H1 | End 2 | M | 14 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndSpare_BDcIn | DCPower
H1 | End 2 | M | 15 | EL1124 | Channel | 1 | Input | PLC2 | Standard | Inputs | BOOL | AlsEndRefl_ADcIn | Status
H1 | End 2 | M | 15 | EL1124 | Channel | 2 | Input | PLC2 | Standard | Inputs | BOOL
H1 | End 2 | M | 15 | EL1124 | Channel | 3 | Input | PLC2 | Standard | Inputs | BOOL
H1 | End 2 | M | 15 | EL1124 | Channel | 4 | Input | PLC2 | Standard | Inputs | BOOL
H1 | End 2 | M | 16 | EL2124 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_ADcOut | Gain | 0 | 0 | 1
H1 | End 2 | M | 16 | EL2124 | Channel | 2 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_ADcOut | Gain | 1 | 0 | 1
H1 | End 2 | M | 16 | EL2124 | Channel | 3 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndRefl_ADcOut | Gain | 2 | 0 | 1
H1 | End 2 | M | 16 | EL2124 | Channel | 4 | Output | PLC2 | Standard | Outputs | BOOL
H1 | End 2 | M | 17 | EL3102 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndRefl_ADcIn | DCPower
H1 | End 2 | M | 17 | EL3102 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 1 | DCMReady | PLC1 | Standard | Inputs | WORD | SysEndBDiv_AIn | DCMReady
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 1 | DCMWarning | PLC1 | Standard | Inputs | WORD | SysEndBDiv_AIn | DCMWarning
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 1 | DCMError | PLC1 | Standard | Inputs | WORD | SysEndBDiv_AIn | DCMError
H1 | End 2 | R | 1 | EL7332 | DCM Control Channel | 1 | DCMEnable | PLC1 | Standard | Ouptuts | WORD | SysEndBDiv_AOut | DCMEnable
H1 | End 2 | R | 1 | EL7332 | DCM Velocity Channel | 1 | Velocity | PLC1 | Standard | Inputs | WORD | SysEndBDiv_AOut | DCMVelocity
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 2 | DCMReady | PLC1 | Standard | Inputs | WORD
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 2 | DCMWarning | PLC1 | Standard | Inputs | WORD
H1 | End 2 | R | 1 | EL7332 | DCM Status Channel | 2 | DCMError | PLC1 | Standard | Inputs | WORD
H1 | End 2 | R | 1 | EL7332 | DCM Control Channel | 2 | DCMEnable | PLC1 | Standard | Ouptuts | WORD
H1 | End 2 | R | 1 | EL7332 | DCM Velocity Channel | 2 | Velocity | PLC1 | Standard | Inputs | WORD
H1 | End 2 | R | 2 | EL1094 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndBDiv_AIn | ClosedPositionSwitch
H1 | End 2 | R | 2 | EL1094 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndBDiv_AIn | OpenPositionSwitch
H1 | End 2 | R | 2 | EL1094 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL
H1 | End 2 | R | 2 | EL1094 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL
H1 | End 2 | R | 3 | EL1094 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndShutter_AIn | State
H1 | End 2 | R | 3 | EL1094 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndShutter_BIn | State
H1 | End 2 | R | 3 | EL1094 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL
H1 | End 2 | R | 3 | EL1094 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL
H1 | End 2 | R | 4 | EL2008 | Channel | 1 | Output | PLC1 | Standard | Outputs | BOOL | SysEndShutter_AOut | Close
H1 | End 2 | R | 4 | EL2008 | Channel | 2 | Output | PLC1 | Standard | Outputs | BOOL | SysEndShutter_BOut | Close
H1 | End 2 | R | 4 | EL2008 | Channel | 3 | Output | PLC1 | Standard | Outputs | BOOL | SysEndShutter_BOut | Open
H1 | End 2 | R | 4 | EL2008 | Channel | 4 | Output | PLC1 | Standard | Outputs | BOOL | SysEndShutter_COut | Open
H1 | End 2 | R | 4 | EL2008 | Channel | 5 | Output | PLC1 | Standard | Outputs | BOOL
H1 | End 2 | R | 4 | EL2008 | Channel | 6 | Output | PLC1 | Standard | Outputs | BOOL
H1 | End 2 | R | 4 | EL2008 | Channel | 7 | Output | PLC1 | Standard | Outputs | BOOL
H1 | End 2 | R | 4 | EL2008 | Channel | 8 | Output | PLC1 | Standard | Outputs | BOOL
H1 | End 2 | R | 5 | EL3102 | AI Standard Channel | 1 | Value | PLC1 | Standard | Inputs | WORD | SysEndPico_AIn | TEMP_X
H1 | End 2 | R | 5 | EL3102 | AI Standard Channel | 2 | Value | PLC1 | Standard | Inputs | WORD | SysEndPico_AIn | TEMP_Y
H1 | End 2 | R | 6 | EL1014 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | FAULT_X
H1 | End 2 | R | 6 | EL1014 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | FAULT_Y
H1 | End 2 | R | 6 | EL1014 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | ON
H1 | End 2 | R | 6 | EL1014 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | POWER
H1 | End 2 | R | 7 | EL3102 | AI Standard Channel | 1 | Value | PLC1 | Standard | Inputs | WORD | SysEndPico_BIn | TEMP_X
H1 | End 2 | R | 7 | EL3102 | AI Standard Channel | 2 | Value | PLC1 | Standard | Inputs | WORD | SysEndPico_BIn | TEMP_Y
H1 | End 2 | R | 8 | EL1014 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | FAULT_X
H1 | End 2 | R | 8 | EL1014 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | FAULT_Y
H1 | End 2 | R | 8 | EL1014 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | ON
H1 | End 2 | R | 8 | EL1014 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | POWER
H1 | End 2 | R | 9 | EL1872 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[1]
H1 | End 2 | R | 9 | EL1872 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[2]
H1 | End 2 | R | 9 | EL1872 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[3]
H1 | End 2 | R | 9 | EL1872 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[4]
H1 | End 2 | R | 9 | EL1872 | Channel | 5 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X{5]
H1 | End 2 | R | 9 | EL1872 | Channel | 6 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[6]
H1 | End 2 | R | 9 | EL1872 | Channel | 7 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[7]
H1 | End 2 | R | 9 | EL1872 | Channel | 8 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_X[8]
H1 | End 2 | R | 9 | EL1872 | Channel | 9 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[1]
H1 | End 2 | R | 9 | EL1872 | Channel | 10 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[2]
H1 | End 2 | R | 9 | EL1872 | Channel | 11 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[3]
H1 | End 2 | R | 9 | EL1872 | Channel | 12 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[4]
H1 | End 2 | R | 9 | EL1872 | Channel | 13 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[5]
H1 | End 2 | R | 9 | EL1872 | Channel | 14 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[6]
H1 | End 2 | R | 9 | EL1872 | Channel | 15 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[7]
H1 | End 2 | R | 9 | EL1872 | Channel | 16 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_AIn | MOTOR_Y[8]
H1 | End 2 | R | 9 | EL1872 | InfoData |  | State | PLC1 | Fast | Inputs | BOOL | SysEndPico_AFastIn | INFO_DATA
H1 | End 2 | R | 10 | EL2872 | Channel | 1 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | MOTOR[3]
H1 | End 2 | R | 10 | EL2872 | Channel | 2 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | MOTOR[2]
H1 | End 2 | R | 10 | EL2872 | Channel | 3 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | MOTOR[1]
H1 | End 2 | R | 10 | EL2872 | Channel | 4 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | MOTOR[4]
H1 | End 2 | R | 10 | EL2872 | Channel | 5 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | IDLE
H1 | End 2 | R | 10 | EL2872 | Channel | 6 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | ENABLE
H1 | End 2 | R | 10 | EL2872 | Channel | 7 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | DIR_X
H1 | End 2 | R | 10 | EL2872 | Channel | 8 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_AOut | DIR_Y
H1 | End 2 | R | 10 | EL2872 | Channel | 9 | Output | PLC1 | Fast | Outputs | BOOL | SysEndPico_AFastOut | CLK_X
H1 | End 2 | R | 10 | EL2872 | Channel | 10 | Output | PLC1 | Fast | Outputs | BOOL | SysEndPico_AFastOut | CLK_Y
H1 | End 2 | R | 11 | EL1872 | Channel | 1 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[1]
H1 | End 2 | R | 11 | EL1872 | Channel | 2 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[2]
H1 | End 2 | R | 11 | EL1872 | Channel | 3 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[3]
H1 | End 2 | R | 11 | EL1872 | Channel | 4 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[4]
H1 | End 2 | R | 11 | EL1872 | Channel | 5 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X{5]
H1 | End 2 | R | 11 | EL1872 | Channel | 6 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[6]
H1 | End 2 | R | 11 | EL1872 | Channel | 7 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[7]
H1 | End 2 | R | 11 | EL1872 | Channel | 8 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_X[8]
H1 | End 2 | R | 11 | EL1872 | Channel | 9 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[1]
H1 | End 2 | R | 11 | EL1872 | Channel | 10 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[2]
H1 | End 2 | R | 11 | EL1872 | Channel | 11 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[3]
H1 | End 2 | R | 11 | EL1872 | Channel | 12 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[4]
H1 | End 2 | R | 11 | EL1872 | Channel | 13 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[5]
H1 | End 2 | R | 11 | EL1872 | Channel | 14 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[6]
H1 | End 2 | R | 11 | EL1872 | Channel | 15 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[7]
H1 | End 2 | R | 11 | EL1872 | Channel | 16 | Input | PLC1 | Standard | Inputs | BOOL | SysEndPico_BIn | MOTOR_Y[8]
H1 | End 2 | R | 11 | EL1872 | InfoData |  | State | PLC1 | Fast | Inputs | BOOL | SysEndPico_BFastIn | INFO_DATA
H1 | End 2 | R | 12 | EL2872 | Channel | 1 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | MOTOR[3]
H1 | End 2 | R | 12 | EL2872 | Channel | 2 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | MOTOR[2]
H1 | End 2 | R | 12 | EL2872 | Channel | 3 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | MOTOR[1]
H1 | End 2 | R | 12 | EL2872 | Channel | 4 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | MOTOR[4]
H1 | End 2 | R | 12 | EL2872 | Channel | 5 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | IDLE
H1 | End 2 | R | 12 | EL2872 | Channel | 6 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | ENABLE
H1 | End 2 | R | 12 | EL2872 | Channel | 7 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | DIR_X
H1 | End 2 | R | 12 | EL2872 | Channel | 8 | Output | PLC1 | Standard | Outputs | BOOL | SysEndPico_BOut | DIR_Y
H1 | End 2 | R | 12 | EL2872 | Channel | 9 | Output | PLC1 | Fast | Outputs | BOOL | SysEndPico_BFastOut | CLK_X
H1 | End 2 | R | 12 | EL2872 | Channel | 10 | Output | PLC1 | Fast | Outputs | BOOL | SysEndPico_BFastOut | CLK_Y

## End 3
Ifo | Chassis | Rail | # | Terminal | Channel | # | Name | PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len
H1 | End 3 | L | 1 | EL4134 | AI Standard Channel | 1 | Value | PLC3 | Standard | Outputs | WORD | PCalEndOut | OpticalFollowerServoGain
H1 | End 3 | L | 1 | EL4134 | AI Standard Channel | 2 | Value | PLC3 | Standard | Outputs | WORD | PCalEndOut | OpticalFollowerServoOffset
H1 | End 3 | L | 1 | EL4134 | AI Standard Channel | 3 | Value | PLC3 | Standard | Outputs | WORD | PCalEndOut | LaserPowerControl
H1 | End 3 | L | 1 | EL4134 | AI Standard Channel | 4 | Value |  | Standard | Outputs | WORD
H1 | End 3 | L | 2 | EL3104 | AI Standard Channel | 1 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | WorkingStandardPowerLevel
H1 | End 3 | L | 2 | EL3104 | AI Standard Channel | 2 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | TransmitterModuleTemperature
H1 | End 3 | L | 2 | EL3104 | AI Standard Channel | 3 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | ReceiverModuleTemperature
H1 | End 3 | L | 2 | EL3104 | AI Standard Channel | 4 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | OpticalFollowerServoOscillation
H1 | End 3 | L | 3 | EL3104 | AI Standard Channel | 1 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | LaserDiodeCurrent
H1 | End 3 | L | 3 | EL3104 | AI Standard Channel | 2 | Value | PLC3 | Standard | Inputs | WORD | PcalEndIn | LaserDiodeTemperature
H1 | End 3 | L | 3 | EL3104 | AI Standard Channel | 3 | Value |  | Standard | Inputs | WORD
H1 | End 3 | L | 3 | EL3104 | AI Standard Channel | 4 | Value |  | Standard | Inputs | WORD
H1 | End 3 | L | 4 | EL2624 | Channel | 1 | Output | PLC3 | Standard | Outputs | BOOL | PCalEndOut | OpticalFollowerServoEnable
H1 | End 3 | L | 4 | EL2624 | Channel | 2 | Output | PLC3 | Standard | Outputs | BOOL | PCalEndOut | ShutterPowerEnable
H1 | End 3 | L | 4 | EL2624 | Channel | 3 | Output | PLC3 | Standard | Outputs | BOOL | PCalEndOut | LaserEnable
H1 | End 3 | L | 4 | EL2624 | Channel | 4 | Output |  | Standard | Outputs | BOOL
H1 | End 3 | L | 5 | EL1094 | Channel | 1 | Input | PLC3 | Standard | Inputs | BOOL | PcalEndIn | ShutterClosed
H1 | End 3 | L | 5 | EL1094 | Channel | 2 | Input | PLC3 | Standard | Inputs | BOOL | PcalEndIn | ShutterOpen
H1 | End 3 | L | 5 | EL1094 | Channel | 3 | Input | PLC3 | Standard | Inputs | BOOL | PcalEndIn | OpticalFollowerServoActive
H1 | End 3 | L | 5 | EL1094 | Channel | 4 | Input |  | Standard | Inputs | BOOL
H1 | End 3 | L | 6 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[1].RFMon
H1 | End 3 | L | 6 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[1].LOMon
H1 | End 3 | L | 6 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[2].RFMon
H1 | End 3 | L | 6 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[2].LOMon
H1 | End 3 | L | 7 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[3].RFMon
H1 | End 3 | L | 7 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[3].LOMon
H1 | End 3 | L | 7 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[4].RFMon
H1 | End 3 | L | 7 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_ADemodIn | Seg[4].LOMon
H1 | End 3 | L | 8 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[1].RFMon
H1 | End 3 | L | 8 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[1].LOMon
H1 | End 3 | L | 8 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[2].RFMon
H1 | End 3 | L | 8 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[2].LOMon
H1 | End 3 | L | 9 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[3].RFMon
H1 | End 3 | L | 9 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[3].LOMon
H1 | End 3 | L | 9 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[4].RFMon
H1 | End 3 | L | 9 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfs_BDemodIn | Seg[4].LOMon
H1 | End 3 | L | 11 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt2In | HV10MonitorPZTX
H1 | End 3 | L | 11 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt2In | PositionSensorPZTX
H1 | End 3 | L | 11 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt2In | HV10MonitorPZTY
H1 | End 3 | L | 11 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt2In | PositionSensorPZTY
H1 | End 3 | L | 12 | EL3104 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt1In | HV10MonitorPZTX
H1 | End 3 | L | 12 | EL3104 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt1In | PositionSensorPZTX
H1 | End 3 | L | 12 | EL3104 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt1In | HV10MonitorPZTY
H1 | End 3 | L | 12 | EL3104 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | WORD | AlsEndWfsPzt1In | PositionSensorPZTY
H1 | ALS Laser Table Relay |  | 1 | EP2624 | Channel | 1 | Output | PLC2 | Standard | Outputs | BOOL | AlsEndLaserHeadOut | NoiseEaterRelayOn
H1 | ALS Laser Table Relay |  | 1 | EP2624 | Channel | 2 | Output | PLC1 | Standard | Outputs | BOOL | IlluminatorOut | IlluminatorRelayOn
H1 | ALS Laser Table Relay |  | 1 | EP2624 | Channel | 3 | Output |  | Standard | Outputs | BOOL
H1 | ALS Laser Table Relay |  | 1 | EP2624 | Channel | 4 | Output |  | Standard | Outputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 1 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 2 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 3 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 4 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 5 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 6 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 7 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Input | 8 | Input | PLC3 | Standard | Inputs | BOOL
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 1 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[3] | Gain | 0 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 2 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[3] | Gain | 1 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 3 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[4] | Gain | 0 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 4 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[4] | Gain | 1 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 5 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[1] | Gain | 0 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 6 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[1] | Gain | 1 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 7 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[2] | Gain | 0 | 0 | 1
H1 | AOS Baffle ITMX |  | 1 | EP2316 | Output | 8 | Input | PLC3 | Standard | Outputs | BOOL | AosBaffleDcOut[2] | Gain | 1 | 0 | 1
H1 | AOS Baffle ITMX |  | 2 | EP3174 | AI Standard Channel | 1 | Value | PLC2 | Standard | Inputs | INT | AosBaffleDcIn[3] | DcPower
H1 | AOS Baffle ITMX |  | 2 | EP3174 | AI Standard Channel | 2 | Value | PLC2 | Standard | Inputs | INT | AosBaffleDcIn[4] | DcPower
H1 | AOS Baffle ITMX |  | 2 | EP3174 | AI Standard Channel | 3 | Value | PLC2 | Standard | Inputs | INT | AosBaffleDcIn[1] | DcPower
H1 | AOS Baffle ITMX |  | 2 | EP3174 | AI Standard Channel | 4 | Value | PLC2 | Standard | Inputs | INT | AosBaffleDcIn[2] | DcPower

## Internal
PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len | PLC | Task | IO | Size | Struct | Parameter | Bit | Pos | Len
PLC2 | Standard | Inputs | BOOL | AlsEndFiber_ADemodOut | PowerOk |  |  |  | PLC2 | Standard | Inputs | BOOL | AlsEndSpare_ADemodIn | PowerOk
PLC2 | Standard | Inputs | BOOL | AlsEndRefl_ADemodOut | PowerOk |  |  |  | PLC2 | Standard | Inputs | BOOL | AlsEndSpare_BDemodIn | PowerOk
PLC1 | Fast | Inputs | UINT | Device 0 (EtherCAT) | Frm0WcState |  |  |  | PLC1 | Fast | Inputs | UINT | Device0FastIn | FrmWcState[0]
PLC1 | Fast | Inputs | UINT | Device 0 (EtherCAT) | Frm1WcState |  |  |  | PLC1 | Fast | Inputs | UINT | Device0FastIn | FrmWcState[1]
PLC1 | Fast | Inputs | UINT | Device 0 (EtherCAT) | Frm2WcState |  |  |  | PLC1 | Fast | Inputs | UINT | Device0FastIn | FrmWcState[2]
PLC1 | Fast | Inputs | UINT | Device 0 (EtherCAT) | Frm3WcState |  |  |  | PLC1 | Fast | Inputs | UINT | Device0FastIn | FrmWcState[3]
PLC1 | Standard | Inputs | UINT | Device 0 (EtherCAT) | SlaveCount |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | SlaveCount
PLC1 | Standard | Inputs | UINT | Device 0 (EtherCAT) | DevState |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | DeviceStatus
PLC1 | Standard | InfoData | UINT | Device 0 (EtherCAT) | ChangeCount |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | ChangeCount
PLC1 | Standard | InfoData | UINT | Device 0 (EtherCAT) | DevId |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | DeviceId
PLC1 | Standard | InfoData | UINT | Device 0 (EtherCAT) | AmsNetId |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | NetId
PLC1 | Standard | InfoData | UINT | Device 0 (EtherCAT) | CfgSlaveCount |  |  |  | PLC1 | Standard | Inputs | UINT | Device0In | ConfigSlaveCount