## Slide 1: The LIGO Vacuum Equipment and Beam Tubes: Retrospective and Prospective
Mike Zucker, LIGO Laboratory @ MIT
IUVSTA-LVD @ Institute for Plasma Research, Gandhinagar , India
19-22 March 2013

## Slide 2
About LIGO and LIGO India
Vacuum Requirements & Constraints
Vacuum Equipment
Beam Tubes
Paths for Improvement
Outline

## Slide 3
Newton’s puzzle:
“instantaneous action at a distance”
General Relativity
Spacetime itself is a medium
Geometry carries information
Gmn= 8pTmn
Why must there be gravitational waves?

## Slide 4: Gravitational Waves
K. Thorne (Caltech) , T. Carnahan (NASA GSFC)
Changes of matter anywhere ripple the geometry everywhere

## Slide 5: The GW Spectrum
10-9 Hz
10-4 Hz
100 Hz
103 Hz
Relic radiation
Cosmic Strings
Supermassive BH Binaries
BH and NS Binaries
Binaries coalescences
Extreme Mass Ratio
Inspirals
Supernovae
Spinning NS
10-16 Hz
Inflation Probe
Pulsar timing
Space detectors
Ground interferometers
Laser Interferometer
Gravitational Wave
Observatory

## Slide 6: Some Expected Astrophysical Sources
Compact binary inspiral: “chirps”
  NS-NS, NS-BH, BH-BH
Supernovas or GRBs: “bursts”
  GW signals observed in coincidence with EM or neutrino detectors
Pulsars in our galaxy: “periodic waves”
  Rapidly rotating neutron stars
  Modes of NS vibration
Cosmological: “stochastic background”
  Probe back to the Planck time (10-43 s)

## Slide 7: A New ‘Sense’- A New Universe
Gravitational Waves will provide complementary information, as different from what we know as sound is from sight.

## Slide 8
Great promise, but a great challenge…
A wave’s strength is characterized by its strain
We can calculate the expected strain at Earth for, say, an orbiting binary system;
If we make our interferometer 4,000 meters long,

## Slide 9: Laser Interferometer Gravitational-Wave Detector
4 kilometer long “arm” cavities
1 Megawatt circulating laser power
Free-floating 40 kg quartz test mirrors
Nd:YAG LASER
Beamsplitter
Photodetector

## Slide 10
LIGO Observatory Sites

## Slide 11
LIGO Scientific Collaboration
IUCAA

## Slide 12: No Confirmed Detections Yet…
First generation detectors reached about 100 galaxies
Current predictions in range of
10-4 CBI yr-1 galaxy-1
Need better sensitivity!

## Slide 13: Advanced LIGO
Initial detectors completed in 2000
Design sensitivity achieved 2005
Ran ~ 2.5 years
  No confirmed detection
Facilities, vacuum system designed to be compatible with “ultimate” future interferometers
Advanced LIGO detector upgrade funded ‘08, now being installed
  Design 10x more sensitive
  1,000x greater observable volume (or event rate)

## Slide 14: Advanced LIGO: 10x More Range
Advanced detectors will reach about 100,000 galaxies
Roughly 1 CBI event per month expected
Initial LIGO Range
Advanced LIGO Range

## Slide 15: Source Localization and MultiMessenger Astrophysics
GW detectors are nearly omni-directional
Array working together can determine source location
  Analogous to “aperture synthesis” in radio astronomy
Accuracy tied to diffraction limit
1
2
q
DL = c dt

## Slide 16
LIGO
GEO
Virgo
Future Global Detector Network
KAGRA
US, Europe and Japan detectors are close to co-planar—not optimal
LIGO-India
India site out of plane breaks degeneracy, improving sky coverage

## Slide 17: A Network with LIGO India, LIGO US and Virgo
HLVI contains four 3-way networks
Baselines with India are1.5x longer
~ 4x better directionality (for given SNR)
B. S. Sathyaprakash et al, ‘Scientific benefits of moving one of LIGO Hanford detectors to India’, LIGO-T1200219-v1

## Slide 18: Timing Errors without LIGO-India

## Slide 19: Timing Errors with LIGO-India

## Slide 20: LIGO India Core Team
IPR, Gandhinagar – Facilities, Beam Tubes, Vacuum Equipment, Controls
RRCAT, Indore – Interferometer Optimization, Installation, Commissioning
IUCAA, Pune – Site survey, Astrophysical Data Analysis & Computing
LIGO-US WILL CONTRIBUTE COMPLETE INTERFEROMETER PLUS TECHNOLOGY TRANSFER AND TECHNICAL SUPPORT

## Slide 21
LIGO Vacuum Equipment

## Slide 22
“Vacuum Equipment:”
Chambers, pumps, instruments
  Houses detector apparatus
  Isolation (valves), access (doors)
  Electrical, mechanical, optical penetrations
  Pumping & instrumentation
  Somewhat “conventional”
  F:A ~ 10-2 ls-1cm-2
Beam tubes
  A long hole in the air;
      Never to be vented
  Highly “unconventional”
    20 million liters (per site)
    600 million cm2 (per site)
    200 l/s char. conductance
    F:A ~ 10-5 ls-1cm-2
LIGO Really Has Two Vacuum Systems

## Slide 23
LIGO Vacuum Requirements
(partial list)
Light scattering phase noise from residual gas
    Function of molecular polarizability, transit speed and partial pressure
    Primary goals for beam tubes:
      P(H2) < 10-9 Torr
      P(H2O) < 10-10 Torr
Contamination of optics
    Mirror absorption < 0.1 ppm
    Hydrocarbons: < 1 monolayer/10 years
      Aggressive cleaning and vacuum bake of every component
    Particles: < one 10 µm particle on any mirror
      ISO Class 5 or better cleanroom protocol for worker access, internal components, surface exposure
Vibration-free environment
      No mechanical, turbo or closed-cycle cryo pumps in steady state operation
NB: Unlike accelerator, plasma, or aerospace applications, we have no radiation, thermal, or ion loading ; in LIGO outgassing is passive at ambient temperature

## Slide 24: Vacuum System Schematic
4,000 m
50 m
10 m
~
~
~

## Slide 25: Corner Station Layout
50 m

## Slide 26: End Station Arrangement
1.2m Ø gate valves
Coaxial
LN2 trap
1.4e5 l/s (H2O)
BSC chamber
LN2 storage
(outdoors)
ion pump
2500 l/s (N2)
beamtube

## Slide 27: Beamtube Gate Valves
40” & 44” ID gate valves to isolate beamtubes, sections, LN2 traps
Double O-ring gates & bonnet seals with pumped annulus*
Two actuator varieties: electric (ballscrew) and pneumatic (cylinder)
Custom design by GNB Corp.
* Principal volume is vulnerable when gate is open!

## Slide 28
In US, we installed beam tubes first to allow time for bakeout
VE contractor supplied valves & pumps to tube contractor

## Slide 29
BSC chamber
(Basic Symmetric Chamber)
2.8m Ø x 5.5m h for large cavity optics
Upper third is a removable dome
Thin (10-15mm) 304L SS shell with welded stiffeners, F&D heads
Combination of GTAW and plasma welding
Major weldments stress-relieved*
  *NOTE SURFACE FINISH!
Ports < 35cm Ø: ConFlatTM
Ports > 35cm Ø: Dual O-ring
  Treated Viton elastomer
  DRY (no grease)
  Isolated pumped annulus between inner and outer seal
  Permeation and damage tolerant

## Slide 30: In Situ Chamber Bakeout

## Slide 31: Particulate control: movable ISO Class 5 cleanrooms
Part of VE contract due to special features required for chamber access

## Slide 32: BSC Equipment Installation

## Slide 33
HAM chamber
(Horizontal Access Module)
House complex input/output optics
2.1m Ø x 2m w
More than 70% of area is removable access doors

## Slide 34
L1 ISI 1/31/08
Notes: HAM ISI insertion in HAM6 at Livingston with guidance from Jeff Kissel (LSU, left) and Brian O’Reilly (LLO, right). March 2008.

## Slide 35
Notes: L1 OMC in its double pendulum suspension, on the HAM ISI platform inside LHAM6, 3/14/2008 (photo credit: Sam Waldman)

## Slide 37: End Station Pressure Evolution after Backfill

## Slide 38: Beam Tubes
304L SS, 3.2 mm thick with external stiffeners
raw coiled stock air baked 36h @ 455C to deplete hydrogen
  JH2 < 1e-13 Tl/s/cm2
  permits ultimate P without distributed pumps
  process developed by LIGO
prepared coil spiral-welded into 1.2m tube on modified culvert mill
16m sections cleaned, leak checked, and capped
FTIR analysis to confirm HC-free
sections field butt-welded together in travelling clean room
Over 50 linear km of weld—

## Slide 40: Depleting H from raw SS before fabrication: An economical alternative to high T vacuum bakeout
SS sheet from mill is baked in air 36 hours at 455 ºC
(Hotter treatment deemed inadvisable due to carbide formation)
Total dissolved hydrogen is reduced ~ 3x
Remaining H is tightly bound, high activation T
Care is required in welding to avoid re-introduction of H
Sample Temperature (ºC)
H2 Flux (A.U.)
Raw SS
After
Airbake
data courtesy of Virgo

## Slide 41
Cleaned with pressurized hot water and detergent
QA by FTIR sampling

## Slide 42
Leak Test “Coffin”

## Slide 43
Leak Test “Coffin”

## Slide 44
position
butt weld
Beam Tube Field Assembly
leak check
field fitup
next section
transport

## Slide 45
Field butt weld made in movable shelter
Internal “dam” shields inside of weld with inert gas
Dam is later filled with He for leak test
Finally, garbed worker crawls in to remove dam & place optical baffles
Tube
Assembly

## Slide 46
Welding Shelter

## Slide 47: I2R Bakeout to Desorb Water
Glass wool insulation
IDC = 2,000 A
~ 3 weeks @ 160ºC
Final JH20 < 2e-17 Tl/s/cm2
Tubes never to be vented

## Slide 48: How Often do you get a Second Chance?? Some Updates Under Study for LIGO India
Move to “conventional” large gate valve design ?
  Reduce cost with single 0-ring gate seal, metallic bonnet seal?
  Increase reliability with “standard product”?
  Reduce complexity of two actuator styles
Remove oxide from stress-relieved vessel walls?
  Oxide patina was good for suppressing IR scatter, good for vacuum performance, but began to flake off with age– had to be removed!
  Considering conventional treatments such as passivation or electropolish
Is spiral-welding still the most economical tube construction?
Are there alternative ways to achieve ultralow dissolved hydrogen?
Should we upgrade monitoring & instrumentation along tubes?
Should we consider 316L in place of 304L for higher corrosion resistance?
Should we consider flanging or other means to reversibly isolate tube sections after installation?

## Slide 49: Cautionary Tale: The LLO Y Beamtube Leak
Discovered only late last year, but…
Reconstruct t=0 in 2008, F = 2.5e-4 torr-liter/sec (!)
  Unnoticed due to sparse instrumentation, masked by detector outgassing
  Approximately 1e5 x specification; MUST BE REPAIRED
    Legacy of water vapor deposited since ‘08 may persist, even after repairs
Localized near Y midpoint by gradient methods
Followed by He MSLD test: VERY DIFFICULT ON THIS SIZE SYSTEM !
At least 4 distinct leaks discovered to date, in at least 3 zones
  Confirms this is no fluke or isolated defect, but a progressive problem
  Most likely a spectrum of sources will be found
  Largest breach is now sealed but about 1e-5 torr-liter/sec remains unaccounted for
Those discovered to date coincide with
  Welds (both a spiral weld and a stiffener fillet), plus
  Animal residues (mice or mud wasps; emit “corrosion accelerants”), plus
  Local history of persistent water incursion
Team (incl. yours truly) has been assigned to find, repair, prevent recurrence
  Outside metallurgical & welding specialists under contract
  Additional diagnostics on known leaks, representative fab samples
  Too soon to reliably bound full $ and schedule impact!

## Slide 50
mouse latrine
mud dauber nest
evacuated temporary cover patch
pressure transient vs. distance for He injection

## Slide 52
Final Remarks
LIGO facilities are among the largest high-vacuum systems ever built, and have stringent requirements
Novel and cost-effective methods were developed to meet these challenges successfully in the US
In trouble-free operation over a decade, LIGO is now installing second-generation instruments in Louisiana and Washington
A third identical interferometer was built and is now designated for India when the site is available
With benefit of hindsight, new technology, and Indian innovation, we believe the LIGO India vacuum system will be even better!

## Slide 53: Thank you IUVSTA!
Thank you IPR!

## Slide 54
--Reference Slides--

## Slide 55: LIGO Livingston Observatory

## Slide 56: LIGO Hanford Observatory

## Slide 57
First detectors reached design sensitivity in 2005
Now installing Advanced detectors
Vacuum requirement <10-9 torr H2 <10-10 torr H2O
Limits to Sensitivity

## Slide 58: Residual Gas Index Fluctuation Noise
ρ = gas number density (~ pressure)
α = optical polarizability (~ index)
w = beam radius
v0 = most probable thermal speed
L0 = arm length
ΔL = arm optical path difference
Statistical model verified by interferometer experiment
S. Whitcomb and MZ, Proc. 7th Marcel Grossmann Meeting on GR, R. Jantzen and G. Keiser, eds. World Scientific, Singapore (1996).

## Slide 59: Residual Gas Pressure Limits in Beam Tubes
To avoid optical phase noise in laser path

## Slide 62: Pressure evolution for major species during 160ºC beam tube module bakeout

## Slide 65: Vertex Pressure Evolution after Backfill

## Slide 66
Large Flange Design
Dual “dry” o-ring
Pumped annulus between
(independent of main volume, except in open gate valve)
Seal faces single-point machined
Controlled circumferential 32 μinch “tooth” finish
Custom Viton (Flourel) cord formulation (55 gallon min. order)
Cleaned & baked after molding to remove volatile compounds and mold release wax

## Slide 67: Pumping
All “dry” pumping
Initial evacuation by blowers & maglev turbomolecular pumps
Maintained by noble diode ion pumps and coaxial LN2

## Slide 68: 80K LN2 Cryopump
Special “low vibration” design
  LN2 reservoir suspended by compliant springs with Flourel dampers
  Dual-phase liquid delivery (horizontal vacuum-jacketed lines) maintain continuous liquid and vapor flow without slugging
  Sloped “chute” introduces new liquid to reservoir
  Reservoir shape & free surface area designed to preclude boiling
  Continuous level control (PID with differential pressure level sensor)
Low-emissivity aluminum cold surface
Low-emissivity tube liners reduce thermal flux
Outdoor storage dewars refilled periodically by truck delivery

## Slide 69: Cryopump Design