## Slide 1: Cryogenic Test Masses for LIGO Upgrades
Brett Shapiro
Stanford University
LIGO

## Slide 2: Summary
Cryo test mass problem statement
Thoughts on LIGO III quad pendulum upgrades
Stanford experiments
  Preliminary rapid cooling
  Preview of upcoming experiments
Rapid cooling thoughts for LIGO III
List of problems to solve
LIGO

## Slide 3: Cyro Test Mass Problem Statement
* For LIGO III, reduce suspension and coating thermal noise by cooling the lower quad to 120 K (-153.15 C)
  Get to 120 K in a timely manner
  Then maintain 120 K
Include a warm-up scheme (don’t forget!)
Do not increase the test mass lossiness
  Emissive coatings, heat links, thick sus fibers, etc
Do not compromise passive seismic isolation
  Cables, hoses, links, etc
Keep about the same seismic isolation platforms (ISIs, HEPIs)
  Limit the amount of extra weight on the ISI
  Leave the rest of the BSC warm

## Slide 4: Possible LIGO III Mechanical Upgrades
10 kg
20 kg
Main
chain
Reaction
chain
143 kg
80 kg
Main
chain
Reaction
chain
Advanced LIGO quad pendulum
Preliminary LIGO III quad pendulum
G1300196-v1
20 kg
2.14 m
1.626 m
40 kg
20 kg
40 kg
20 kg
Fused silica, 295 K
Silicon, 120 K

## Slide 5: Why These Test Mass Changes?
Thermal elastic noise and CTE of silicon goes to zero at ≈120 K (and 18 K)
Silicon’s mechanical loss decreases at lower temperatures, silica’s increases
Silicon has high thermal conductivity – decreasing thermal lensing with high laser power
Higher mass reduces radiation pressure sensitivity and helps reduce thermal noise with less surface/volume ratio
References: [1] Strawman report – T1200031
[2] “Test mass materials for a new generation of gravitational wave detectors”; 2003; Rowan, Byer, Fejer, et al.

## Slide 6: Possible LIGO III Cryogenic Upgrade
120 K
295 K
Heat shield
77 K
ISI stage 2
Vacuum flange
Flexible liquid nitrogen hose
(77 K = -196 C)
clamp
1560 nm laser rather than 1064 nm
See also G1200246-v1
≈ 3 W at 120 K

## Slide 7: Test Mass Radiation Simulation
Ideal radiation shield:
Fully encloses test mass
77 K (liquid N2)
Shield emissivity = 1
Test mass barrel emissivity = 0.5
Laser power = 0
Power α (Temperature)^4

## Slide 8: A Lot of Heat to Remove
60 W
143 kg Test Mass:
14 MJ to get cold (warm)
2.7 days to 120 K (295 K) at a constant 60 W
Maybe there is another way to cool the test mass quickly

## Slide 9: First Si Test Mass Prototype

## Slide 10: Preliminary Stanford Tests Last Year
Goals: 1. Test ways to cool objects using liquid nitrogen.
2. Compare results with models to see if we understand.
3. Get a feel for what liquid nitrogen will and will not do.

## Slide 13: Heat conduction through Cu wires
Copper heat conducting wires
29 g silicon sample

## Slide 14: Measurements from Preliminary Setup

## Slide 15: Simulation-Measurement Comparison

## Slide 16: New Experimental Setup – Cold Links and Radiation

## Slide 17: New Experimental Setup
Silicon ‘test mass’ prototype
Flexible cold link

## Slide 18: List of Upcoming Experiments
Experiments - 1st set:
cold link (pincher) test
  Comparison with models
  Does residual air pressure enhance thermal conductivity?
radiative test
  Emissive coatings
  Measure of max heat input at 120 K.
Experiments - 2nd set:
suspension spring test to see if we can control the spring temp and Si temp simultaneously
Optical temperature measurement of Si - measure temperature dependent mode frequencies

## Slide 19: LIGO III Movable Cold Link
Test mass
Test mass front view
Test mass top view
liquid N2 pipe
Cu alloy cold link
cold link mover
flexible N2 pipe
liquid N2 pipe
Cu alloy cold link
cold link mover
flexible N2 pipe

## Slide 20: Simulation of Test Mass Cool Down
20 cold link loops:
50 mm diameter
50 mm wide
0.1 mm thick
77 K cold end
≈ 14 MJ
Power α (Temperature difference)

## Slide 21: Other Problems To Solve
Liquid N2 hoses flexible enough for ISI under vacuum
Temperature/height control of blade springs
Test mass temperature control
How to measure temperature?
  Measure acoustic modes – Young’s modulus is temp. dependent
  Infrared camera
Emissivity of optical coatings
Lossiness of emissive coatings
Good emissivity estimates/measurements of Si?
Power absorption in Si (ppm, W, etc)?
How noisy is bubbling nitrogen: seismic, Newtonian? Do boiling chips help?
Optical coating thermal noise at 120 K

## Slide 22: LIGO Cryo Resources
Strawman website – https://nodus.ligo.caltech.edu:30889/wiki/doku.php?id=strawman
Strawman report - T1200031
Google doc - https://docs.google.com/spreadsheet/ccc?key=0AqRDYyFEjUXXdHpYc3dyTlh2TXNDem5oWkhvY0NudVE#gid=0

## Slide 23: Conclusions
LIGO folks and larger GW community thinking more seriously about cryo test masses
Experiments underway at Stanford to test cooling technology
Lots of problems to solve!
LIGO

## Slide 24: Backups
LIGO

## Slide 25: aLIGO vs LIGO III Quad TF

## Slide 26: aLIGO vs LIGO III Quad TF

## Slide 27
Silicon ‘test mass’ prototype
Test mass support frame
Flexible cold links
Copper strips
Insulating Kapton strips
Cold link manipulating cables
Cable roller bearing
Pivoting copper rod
Extension spring

## Slide 28: Questions to Answer – G1200803
How do we get the cold to the test mass?
  Long, flexible heat links running down the ISI and the suspension will have low thermal conductivity and may compromise isolation performance
  Might bring N2 into BSC with vacuum proof hoses, but these will be even stiffer than heat links.
  Heat pipes have excellent thermal conductivity, but are even stiffer still.
  Radiative cooling bypasses at least some of the heat link issues, but can introduce scattering, parasitic modes, electrostatic coupling.
How much temp regulation and tolerance?
How to measure temp?
  Measure acoustic modes – modulus of elasticity is temperature dependent
  Use IR camera
How do heat links compromise performance?
  Lossiness
  Seismic isolation degradation
Thermal parameters of optical coatings
Lossiness of emissive coatings
Good emissivity estimates/measurements of Si?
Power absorption in Si (ppm, W, etc)?
Strength of Silicon fibers/ribbons?
What information are we missing?
Notes: Alan says studying strength of Si. Minimum strength of 573 MPa

## Slide 29: Reasons not to have open LN2 in BSC
The mirror coatings may have a thermal shock issue
The nitrogen must be very clean before contacting the test mass.
Exposure to vacuum will cause liquid nitrogen to explode.