## SR2-04 (second check 3Sep13)
DCC: T-1300653-v1
Measurement Uncertainties in Determining the Radius of Curvature for SR2-04
R. Martin and G. Billingsley
Date: 9/3/2013
We determined the Radius of Curvature from phase maps taken at CIT by R. Martin and G. Billingsley with the Wyko inteferometer. For different mirror orientations we measured the third order Zernike terms and the distance between the optic and the transmission sphere. Radii of Curvature for the Transmission Spheres (TS) were provided by the manufacturer. Based on measurement uncertainties we placed error bars on the calculated values.
For SR2-04, the measured ROC is very close to the lower end of the designed range (absolute value). The cumulative measurement error is less than 0.1%, and is conservative. Measured Zernike terms have the least effect. The uncertainty in measuring the gap between TS and optic contributes to 1 mm change in ROC per 1 mm distance.
[6 empty] | Effect of TS-ROC (0 deg) |  | Effect of Gap (0 deg) |  | Effect of Measured Zernike-3 (0 deg)
SR2-04 (used average of all good files at each rotation) |  |  | 50 mm aperture |  |  | (from ZYGO Report) |  | (conservative, measured +/-1 mm) |  | from Z86 database, for all rotations
|  |  | 0 deg (arrow up) | 90 deg (arrow right) | 225 deg | ROC-x (max) | ROC-y (min) | +2mm | -2mm | Z[3]min=0.003954wv | Z[3]max=0.007719wv
Radius of curvature of TS | Assumes converging TS and cvx part | m | -6.4989 | -6.4989 | -6.4989 | -6.5021 | -6.4956 | -6.4989 | -6.4989 | -6.4989 | -6.4989
Gap between TS and part to be measured | Vertex to vertex, no tilt | m | 0.09462 | 0.09462 | 0.09532 | 0.09462 | 0.09462 | 0.09662 | 0.09262 | 0.09462 | 0.09462
Zernike power coefficient | Assumes "Fringe" Zernikes | nm | 5.5637 | 7.0107 | 6.8947 | 5.5637 | 5.5637 | 5.5637 | 5.5637 | 4.2071 | 8.213
Part measured aperture |  | mm | 50.274 | 50.274 | 50.274 | 50.274 | 50.274 | 50.274 | 50.274 | 50.274 | 50.274
Wavefront "sag" at part plus power |  | mm | -0.049321 | -0.049318 | -0.049324 | -0.049296 | -0.049346 | -0.049336 | -0.049306 | -0.049324 | -0.049316
Part radius |  | m | -6.4057 | -6.4061 | -6.4053 | -6.4089 | -6.4025 | -6.4037 | -6.4077 | -6.4053 | -6.4064
|  |  | -6.4057 |  |  | -0.0032114 | 0.0032215 | 0.0020009 | -0.0020009 | 0.00035236 | -0.00068826
|  |  | 7.0921e-06 | -0.00036881 | 0.00036171 | Tolerances due to RS-ROC |  | Tolerances due to gap TS-optic |  | Tolerances due to Z-3 fluctuations
|  |  | +/- 0.37 mm |  |  | +/- 3.2 mm |  | +/- 2.0 mm |  | +/- 0.7 mm
Cumulative: + 5.6 mm /- 5.9 mm (<0.1%)
Specified Radius |  | m | -6.43
|  | (+/-0.03) | -6.4 | min
|  |  | -6.46 | max
[6 empty] | Effect of TS-ROC (0 deg) |  | Effect of Gap (0 deg) |  | Effect of mask centering on measured Zernike-3 (0 deg)
|  |  | 25 mm aperture |  |  | (from ZYGO Report) |  | (conservative, measured +/-1 mm)
|  |  | 0 deg (arrow up) | 90 deg (arrow right) | 225 deg | ROC-x (max) | ROC-y (min) | +2mm | -2mm | x+5.32mm, y0 | x-5.32mm, y0 | x0, y+5.32mm | x0, y-5.32mm
Radius of curvature of TS | Assumes converging TS and cvx part | m | -6.4989 | -6.4989 | -6.4989 | -6.5021 | -6.4956 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989
Gap between TS and part to be measured | Vertex to vertex, no tilt | m | 0.09462 | 0.09462 | 0.09532 | 0.09462 | 0.09462 | 0.09662 | 0.09262 | 0.09462 | 0.09462 | 0.09462 | 0.09462
Zernike power coefficient | Assumes "Fringe" Zernikes | nm | 0.66926 | 0.56179 | 0.43624 | 0.66926 | 0.66926 | 0.66926 | 0.66926 | 1.2151 | 1.3364 | 1.5258 | 0.81396
Part measured aperture |  | mm | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22 | 25.22
Wavefront "sag" at part plus power |  | mm | -0.012413 | -0.012413 | -0.012415 | -0.012407 | -0.01242 | -0.012417 | -0.012409 | -0.012412 | -0.012412 | -0.012412 | -0.012413
Part radius |  | m | -6.4049 | -6.4048 | -6.404 | -6.4081 | -6.4017 | -6.4029 | -6.4069 | -6.4055 | -6.4056 | -6.4058 | -6.4051
|  |  | -6.4046 |  |  | -0.0032107 | 0.0032207 | 0.0020004 | -0.0020004 | -0.00056332 | -0.00068852 | -0.00088401 | -0.00014933
|  |  | -0.00035048 | -0.00023959 | 0.00059007 | Tolerances due to RS-ROC |  | Tolerances due to gap TS-optic |  | Tolerances due to Z-3 fluctuations over the selected aperture location
|  |  | +/- 0.59 mm |  |  | +/- 3.2 mm |  | +/- 2.0 mm |  | +0 mm/ -0.9 mm
Cumulative: +5.2 mm / - 6.1 mm (<0.1%)