## SRM-05 check 2Sep13
T1300745-v1
SRM-05 ROC Tolerances
R. Martin and G. Billingsley
2013-09-03
Checked indiv files, gaps, recalculated average for each rotation, recalculated ROC
|  | Individual Files | [7 empty] | Average Files |  | Effect of Gap |  | Effect of Zernike uncertainty |  | Effect of ROC-TC tolerance
|  | 06m_SRM-05-0_010 | 06m_SRM-05-0_042 | 06m_SRM-05-0_063 | 06m_SRM-05-0_070 | 06m_SRM-05-225_013 | 06m_SRM-05-225_043 | 06m_SRM-05-225_075 | 06m_SRM-05-225_091 | 0 deg (arrow up) | 225 deg | +2 mm | -2 mm | Zmin=+0.001893 | Zmax=+0.00499 | ROC-x (max) | ROC-y (min)
Radius of curvature of TS | m | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.4989 | -6.5021 | -6.4956
Gap between TS and part to be measured | m | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86262 | 0.86462 | 0.86062 | 0.86262 | 0.86262 | 0.86262 | 0.86262
Zernike power coefficient | nm | 5.1221 | 4.2592 | 2.8558 | 4.2741 | 2.6164 | 3.8027 | 4.1762 | 3.9549 | 3.441 | 3.5123 | 3.441 | 3.441 | 2.0142 | 5.3094 | 3.441 | 3.441
Part measured aperture | mm | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324
Wavefront "sag" at part plus power | mm | -0.020383 | -0.020385 | -0.020388 | -0.020385 | -0.020388 | -0.020386 | -0.020385 | -0.020386 | -0.020387 | -0.020387 | -0.020394 | -0.020379 | -0.02039 | -0.020383 | -0.020375 | -0.020398
Part radius | m | -5.6391 | -5.6386 | -5.6378 | -5.6386 | -5.6377 | -5.6383 | -5.6385 | -5.6384 | -5.6381 | -5.6382 | -5.6361 | -5.6401 | -5.6374 | -5.6392 | -5.6414 | -5.6349
[10 empty] | -5.6382 |  | 2.0014 | -2.0014 | 0.78909 | -1.0336 | -3.2122 | 3.2222
Specified Radius | m | -5.69
| (+/-0.03) | -5.63 | [10 empty] | Cumulative | -6.2471 | mm
|  | -5.75 | [10 empty] | Tolerance: | 6.0126
SRM-05
- Measured on 6-Sep-2011 at CIT with Wyko interferometer by R. Martin and G. Billingsley
- Measured Gap = 85.8 cm (+ 4.62 mm)
- The optic was measured at 0 deg (arrow up), 90 deg (arrow right) and arrow at 225 deg cw, all when looking at the HR side of the optic. Data at 90 deg were lost when the old computer died
- Transmission Sphere Data File - C1203092
- Related Surface Figure report - E1200570