## PUNCH LIST
# | Priority | Description | Holding this review? | Punch-list | Assigned or Action | Installation Subset | E1400152-v12 aLIGO LLO INSTALLATION "PUNCH LIST"
1 | H | Visual inspection of view-ports in-situ was not complete refer to bug list. Bug 761 https://services.ligo-wa.caltech.edu/integrationissues/show_bug.cgi?id=761. View-ports are also not in ICS. This should be completed as part of inventory. | No | Yes | Project | All Chambers
2 | M | Phase 3 testing on Op Lev was not complete. | No | Yes | Project | All Relevant Chambers
3 | H | The ESD system is still undergoing installation and should not be evaluated at this time. That said, there are short circuits in the cabling-connectors. | No | Yes | Project | BSC4, BSC5 (BSC1 & BSC3)
4 | M | There are many references to G1001032. It is not clear whether this document will be brought to completion and maintained at both sites. So, it might be worthwhile before the systems review to either finish it at both sites, or omit it from LLO documentation. In either case, a blanket reference to it is less helpful than would be a pointer to the particular racks referred to in the particular review document. | No | Yes | Project | General Note
5 | L | Bug 20. No ESD on ITM's. Refer to LIGO-L1200291-v5 for Dennis Coyne's comments from 4/23/2014: The aLIGO project may decide to implement the change/correction for the L1 ITMy CP in order to make the ITMy ESD operational. However it does not currently appear to be necessary for completion of the project. This deviation from intended design/implementation shall be noted in the acceptance documentation so that we don't forget the discrepancy. However for the time being no correction for the L1 CPy will be pursued | No | Yes | Ops | LBSC1
6 | L | We also note that the cable to connect the ESD for ITMY is not in place. I tried to find the relevant log entry. But I believe that we changed out the feedthrough for a blank because the feedthrough was leaking. I need to add a note to the LBSC1 issue tracker, but want to find some log entry first. | No | Yes | Ops | LBSC1
7 | M | Bug 615 "Unresponsive GS13 (V2) on ITMY (BSC1)" | No | Yes | Ops (unless opportunity arrises). Joe G notes this is a concerning item!!! | LBSC1
8 | M | Note that this ICS entry is screwed up. It has LHAM1, LHAM3, LHAM5, LHAM6, LBSC1 and LBSC3 chamber assemblies all as indented under BSC-ISI Unit 1, which in turn is under LBSC2 chamber. The LBSC1 chamber assembly is missing the TCS optics. Use the serial number tracking memo M1000051 as a guide for deciding what is essential. | No | Yes | Project | LBSC1
9 | H | E1300847-v7 is missing reference to a BS Optical Lever testing report | No | Yes | Project. Op Lev Report - This is on Op Lev Punch-list (Eric G running) | LBSC2
10 | M | There are some issues with the ICS entry for D0900428 for LBSC2. Some TCS and SLC records have been added but do not appear. Still need to add viewports, Oplev periscope and perhaps misc. other items to the ICS records. | No | Yes | Project. Note that this could be done now if person could be identified. | LBSC2
11 | M | Baffling of the BS Baffle. Bug 505. Currently re-designing on Ops. Plan to install at next opportunity. | No | Yes | Ops (unless opportunity arrises) | LBSC2
12 | M | Note that this ICS entry is screwed up. It has LHAM1, LHAM3, LHAM5, LHAM6, LBSC1 and LBSC3 chamber assemblies all as indented under BSC-ISI Unit 1, which in turn is under LBSC2 chamber. The LBSC1 chamber assembly is missing the TCS optics. Use the serial number tracking memo M1000051 as a guide for deciding what is essential. | No | Yes | Project | LBSC3
13 | H | Bug #738. SEI Cable. Plan to fix at next opportunity. | No | Yes | Performed at next opportunity | LBSC4
14 | H | Bug 856, ESD short in vacuum (or in the feedthrough) |  |  | copy from above | LBSC4
15 | H | IR QPD failure on the TMS | No | Yes | Project | LBSC4
16 | H | a stuck pico-motor | No | Yes | Project | LBSC4
17 | H | Bug #80. Possibility of damage to ESD pattern on ERMs and CPs due to arcing. An interlock system based on vacuum pressure is installed and working at LHO. The LLO vacuum pressure sensors leak. In addition the LLO vacuum engineer has concerns about the chosen pressure gauges, so an alternate sensor is being sought. Modifications to the electronics may be required. | No | Yes | Project | LBSC4 and LBSC5
18 | H | Bug 878, ESD short in vacuum (or in the feedthrough) (similar to bug 856 for LBSC4, but a different pin) | No | Yes | Project | LBSC5
19 | M | ICS Record is lacking. | No | Yes | Project | LHAM1
20 | H | LHAM1 HEPI has not yet been commissioned. | No | Yes | Project | LHAM1
21 | L | HAM1 installation & alignment documents missing. Write HAM1 INS doc and cite E1300857, aLIGO ISC Custom Optics Requirements" within. | No | Yes | Project | LHAM1
22 | M | LHAM2: The two IO Tables, IOT2L and IOT2R, are supposed to be defined in: D0902284-v11, IO Table Layouts for IOT2L and IOT2R ("current default layouts", i.e. initial layouts), but only the layout for IOT2L is provided in the DCC; The initial layout for IOT2R is missing. In addition, the final, as-built layouts for the two IO Tables are to be provided in D1300356, As Built Layouts for ALIGO L1 IOT2L and IOT2R D1300357, As Built Layouts for ALIGO H1 IOT2L and IOT2R which implies that the layouts will be different for the two IFOs. They should be the same, with small dimensional differences included as notes on the same common drawing. In addition, both of these drawings are missing in the DCC. | No | Yes | Project | LHAM2
23 | M | Two S# documents have been filed with the same rack designation (L1-SUS-R1): S1200749 and S1200523. One must be incorrect. Resolve the discrepancy. | No | Yes | Project | LHAM2
24 | M | The DCC entry for S1200744 for rack L1-SEI-C2 is empty. | No | Yes | Project | LHAM2
25 | H | ISS PD Array. Currently re-designing on Ops. Plan to install at next opportunity. An ECR (and bug) will be issued very soon. | No | Yes | Performed at next opportunity | LHAM2
26 | H | Testing-PSL ISS outer loop PD array no installation test document | No | Yes | Project | LHAM2
27 | M | Note that this ICS entry is screwed up. It has LHAM1, LHAM3, LHAM5, LHAM6, LBSC1 and LBSC3 chamber assemblies all as indented under BSC-ISI Unit 1, which in turn is under LBSC2 chamber. The LHAM5 chamber assembly only has the SR3, SRM and cable information. Use the serial number tracking memo M1000051 as a guide for deciding what is essential. | No | Yes | Project | LHAM5
28 | L | Bug #175. SR3 magnet off optic. Won't fix now. | No | Yes | Ops | LHAM5
29 | L | Bug #82. The OFI unit in LHAM5 is not in accordance with the final design (it has a light-weighted bench and lower structural resonances). We do not plan to swap out to the final design unless an serendipitous opportunity arises, or we find that the resonances are a problem. [N.B.: Bug 82 is missing from LHAM5 report.] | No | Yes | Performed at next opportunity, if conclude needed | LHAM5
30 |  | Bug 214 may also apply here. (see comment below in LHAM6) | See below |  |  | LHAM5
31 | L | ICS Record is lacking. | No | Yes | Project | LHAM6
32 | H | LHAM6 HEPI has not yet been commissioned. | No | Yes | Project | LHAM6
33 | L | -only ISI installation document/HA given; please note explicitly there is no installation procedure written for HAM6 optical components like the OMC. For such a critical part, we must add to the punchlist for India the writing of a procedure (low priority). Refer to E1300857, aLIGO ISC Custom Optics Requirements" within. | No | Yes | Project | LHAM6
34 | L | -As-built alignment procedure: was there an exhaustive search of the aLOG for alignment aLOGs? This might be a particularly useful aLOG to record as there is no baseline alignment procedure in existence. | No | Yes | Project | LHAM6
35 | L | Bug 214. The beam was missed OM1 in LHAM6 using the nominal drawing locations. The position of OM1 was changed to make the system wok. The drawing discrepancy is not yet understood. | No | Yes | Ops | LHAM6
36 | M | Missing rack S-number documents and drawings. Specifically: S1202977 and D1201118 for L1-PSL-C1 S1202978 and D1201120 for L1-PSL-R1 S1202979 and D1201121 for L1-PSL-R2 | No | Yes | Project | PSL
37 | M | Environmental testing/characterization (temperature, humidity, particulate levels, magnetic fields, etc.) for the LAE was not documented, but should have been. | No | Yes | Project | PSL
38 | M | 1) We need a hierarchical list of aLIGO-provided VE modifications, for the components that were actually installed at LLO, listing the makes/models, installation location, drawings, serial number (if applicable), and any test results. Links to the aLOG are OK, but they need to be described and to refer to specific entries. | No | Yes | Project (gather while chance exists) | VE
39 |  | 2) All gas-containing ISC components should be leak-tested in-house at the next opportunity. The overall plan should be run by the VRB, as was done for the seismic pods. | No | Yes | Project (gather review documentation and post) | VE
40 | H | Brian, Calum, Dennis to update installation documentation with the notes found by the reviewers (these notes which are not part of the punch-list can be found in the additional tabs of this excel sheet.) | No | Yes | Project | All Chambers

## HAM123 Nots
| DONE | PUNCH | ONES MARKED PUNCH NOW ADDED TO PUNCH LIST
| TO DO | NOTE | THE ONES MARKED NOTE ARE FOR AUTHORS TO CONSIDER. THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
| Chamber | Section | Paragraph(s) | Comment | Calum's Remarks 5/7/14 | BML reply
NOTE | LHAM1 | 1 | Vacuum Chamber(s)/Electronics Rack Designations(s)/other | Acceptable as is. But recommendation to have a link to a documents with all the components that go into the chamber ( such as the initial installation document or equivalent).
Already in punch list | LHAM1 | 1 | Electronics Rack Designations(s) | If there is not official designation for Capacitive Position Sensor readout boxes, please state how they are identified.
NOTE | LHAM1 | 2 | Initial alignment procedures | LHAM1 points to E1100782 which is empty.
NOTE | LHAM1 | 2 | As-Built alignment procedures | Add aLOG link or equivalent to specify how alignment was performed
NOTE | LHAM1 | 4 | Serial Number Records | Incorrect pointer to DCC # D0900421; Assembly name is called out as LHAM6 and not LHAM1. ICS entry dicrpancy is on the punch list
NOTE | LHAM1 | 4 | Serial Number Records | Incorrect pointer to DCC # D1000513 for HEPI assembly. Should be D1000514?
NOTE | LHAM1 | 5 | Testing-SEI | No Phase 3 SEI teting documents. On punch list?
NOTE | LHAM1 | 5 | Testing-AOS/SLC | Add reference to bug #761
Already in punch list | LHAM1 | 6 | Installation Completeness | As stated in report, ICS records for HAM1 may be incomplete. No records for passive stack isolation. What records are provided for I1 for instance?
NOTE | LHAM2 | 1 | Vacuum Chamber(s)/Electronics Rack Designations(s)/other | Acceptable as is. But recommendation to have a link to a documents with all the components that go into the chamber ( such as the initial installation document or equivalent).
Already in punch list | LHAM2 | 1 | Electronics Rack Designations(s) | If there is not official designation for Capacitive Position Sensor readout boxes, please state how they are identified.
NOTE | LHAM2 | 2 | As-Builtinstallation procedures | Are there any other sources that would have as-built notes recorded?
NOTE | LHAM2 | 2 | As-Built installationprocedures | typo? elog #1298 should read elog #9298
NOTE | LHAM2 | 2 | As-Built alignment procedures | typo? The most up do sate as-built alignment procedure is E1200640-v8 |  | Typo- should be E1100783. E1200640 refers to LBSC1 alignment
NOTE | LHAM2 | 5 | Testing-HAM-ISI | Change title of document E1200105; Title says Phase 2 testing, but the report has in-chamber results included. Are these not Phase 3 tests?
NOTE | LHAM2 | 5 | Testing-IO HAMAUXiliary | add link to elog #9743; Q measurements are missing but not required
NOTE | LHAM2 | 5 | Testing-AOS/SLC-VP | add reference to bug list # for visual inspection of View port
NOTE | LHAM2 | 5 | Testing-AOS/OptLev | what is the bug report reference # to the missing test report
NOTE | LHAM2 | 5 | IO-Faraday Isolator | add link to elog # 8608
PUNCH | LHAM2 | 5 | Testing-PSL ISS outer loop PD array | no installation test document | ADDED TO PUNCH LIST
NOTE | LHAM3 | 1 | Vacuum Chamber(s)/Electronics Rack Designations(s)/other | Acceptable as is. But recommendation to have a link to a documents with all the components that go into the chamber ( such as the initial installation document or equivalent). | Could you clarify which documents? | e.g. E120025 which lists all the chamber components plus a thumbnail drawing of each. This just a suggestion. The document is referenced a few boxes down in the report. My comment was just that it would be convenient to have up front at first. It would be nice to have an overview of the structures and components that go inside the chamber.
Already in punch list | LHAM3 | 1 | Electronics Rack Designations(s) | If there is not official designation for Capacitive Position Sensor readout boxes, please state how they are identified. | That is point they are not. Perhaps this becomes punch-list for ML. | Ha! My point is that there should be some identifier that connects each box used to the chamber where its installed. The same comment was made for HAM1 and HAM2.
NOTE | LHAM3 | 2 | Initial alignment procedures | Not signed and dated as indicated in the document. Either sign and date or remove lines from document. | Signed, where? Can you confirm where you found this? | E1100783 contains date, signature lines, and comment entry lines in numberous sections wthroughout the document.
Already in punch list | LHAM3 | 4 | Serial Number Records | Recommend that the ICS entry be edited to be correct. | Add to PUNCH-LIST. No action on me as an author. | OK
NOTE | LHAM3 | 4 | Serial Number Records | Incorrect pointer to DCC # D1000513 for HEPI assembly. Should be D1000514? | Agreed. I will fix. | OK
NOTE | LHAM3 | 5 | Testing-HAM-ISI | Change title of document E1200104; Title says Phase 2 testing, but the report has in-chamber results included. Are these not Phase 3 tests? | Suggest don't fix as a lot of docs for such a small item. | I agree it's a pain to fix over a lot of documents. There should be some place (one place?) to note this discrepancy for the benefit of future operators.
NOTE | LHAM3 | 5 | Testing-AOS/SLC-VP | Status of 2 VP's are not listed. Add reference to bug list # for visual inspection of View port | Bug is listed just not in this location. However, I will fix. | OK
NOTE | LHAM3 | 5 | ISC QPD Sled | Acceptce test are not 'approved'. Add links to T1200065 (green) and T1200063 (IR) | Good links. I will add. | OK
NOTE | LHAM3 | 5 | B&K Hammer Testing | add entry to elog #7238? | It is already linked through 9383 so suggest no need to add. | all right. (I know you're overloaded with work)
Already in punch list | LHAM3 | 6 | ICS Assembly Record | misc. items should be specified. Link to M1000051 is only policy , it is not a lnk to the ICS. Put on punch list? | I don't agree. The person with the BOM and policy will define the misc items. So recommend don't fix in document but agree should be on punch list. | OK to punch list.

## HAM456 Notes
TO DO | THE ONES MARKED NOTE ARE FOR AUTHORS.
DONE | THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
Calum/Dennis/Brian,
Here are comments regarding simple corrections to the LHAM4-6/PSL acceptance documents. In general they look very complete and carefully reviewed previously.
-Mike
* * *
LHAM4
------
NOTE | -if you search the document for HAM5, you find three typos which should be corrected to HAM4. There are additional references to HAM5 which are real however. Good to correct the former, for clarity, e.g. in Section 2, alignments of HAM4 and HAM5 are included in the same document, however the alog entry 6871 is specifically regarding SR2 in LHAM4 (and yet that alog even makes reference to LHAM5 SRM and SR3).
Sec 3:
NOTE | "D1001227" link is top-level drawing (floor occupancy) for TCS. For HAM6, this equivalent section has an actual table layout. What is the goal in referencing these drawings? To vet floor occupancy, or the actual table layouts, or both? What level of consistency is expected between documents?
Sec 6:
NOTE | AOS/SLC/Baffles: the document shows a hyperlink to "LLO aLOG 7144". However, this active link points to LLO aLOG 7140, regarding HAM5. Looking at aLOG 7144, it too refers to LHAM5, so it is not simply that the linked aLOG was not properly updated in this HAM4 document. We have to find the appropriate HAM4 testing aLOG.
Sec 7:
NOTE | repeated entry for bug 399
LHAM5
------
Sec 4:
PUNCH | ICS entry issue lists as low; can we upgrade to medium to ensure it gets traction at some point. | ADDED TO PUNCH LIST
Sec 5:
NOTE | AOS/OFI: correctly refers to LLO aLOG 7140 regarding OFI B&K hammer testing. However, should also reference aLOG 7144 regarding HAM5 baffle B&K tests.
Sec 7:
NOTE | repeated entry for bug 17
LHAM6
------
Sec 1:
NOTE | -typo "LHAM5" should read LHAM6.
NOTE | -optical table layout for ISCHT6R is given drawing; for other HAMs, floor occupancy drawings are given. Which is relevant? Or should both be present?
Sec 2:
PUNCH | -only ISI installation document/HA given; please note explicitly there is no installation procedure written for HAM6 optical components like the OMC. For such a critical part, we must add to the punchlist for India the writing of a procedure (low priority). |  |  |  | ADDED TO PUNCH LIST
NOTE | -The OMC cable link is an original link. There is a followup cable-replacement aLOG which could be referenced too, https://alog.ligo-la.caltech.edu/aLOG/index.php?callRep=7631
PUNCH | -As-built alignment procedure: was there an exhaustive search of the aLOG for alignment aLOGs? This might be a particularly useful aLOG to record as there is no baseline alignment procedure in existence. |  |  | ADDED TO PUNCH LIST
Sec 6:
PUNCH | -No top level ICS record, and in bold red, "...ISC information is sorely lacking...", yet the priority of this item in the checklist is low. Can this too be moved to medium. |  | ADDED TO PUNCH LIST
PSL
---
Sec 3:
NOTE | -Electronic rack drawings link to DCC records, but not drawings
Sec 4:
NOTE | -are these serial no. records complete?

## VE Notes
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
§2, Procedures and Modifications:
NOTE | - Several times the reader is referred to §1 of E130045O for procedures and modifications def'n. This is more-or-less a list of around 40 DCC docs without any road map. For example, it is not possible for me to work out exactly which drawings were built and installed where at LLO.
NOTE | - Several times the reader is asked to carry out an alog search for author Scott McCormick, and is also pointed to several specific entries from Mike Meyer. I am not satisfied that this is the way to archive records of installations. The search under McCormick returns 166 hits and probably misses some things entered by Harry or others.
NOTE | On the other hand, LLO's personnel turn-over in recent years makes this archive difficult to compile.
§3, Drawings:
NOTE | - The drawings look good, but given their dates it is not clear to me where one would find the as-built information.
§4, Serial No. Records:
NOTE | - None are listed.
§5, Testing:
NOTE | - The GNP-provided inspection reports (Q1300009) refer to various S/Ns. So, I would expect §4 to list which S/Ns got installed where at LLO. The drawings in §3 don't seem to have this information.
NOTE | - the reader is asked to carry out an alog search for author Scott McCormick, and is also pointed to several specific entries from Mike Meyer. I am not satisfied that this is the way to archive records of installations. The search under McCormick returns 166 hits and probably misses some things entered by Harry or others.
§6, Installation completeness:
§7, Issues and ECRs:
NOTE | - There are ISC component(s) in LHAM1 that are relevant to VE operation and contamination control. An WFS seal is thought to be leaking Helium (see LLO alog entry 12386). We don't know whether any hydrocarbons have been introduced into the vacuum system that might necessitate cleanup.
NOTE | - The traveller document for what is thought to be the relevant WFS, S1301246, does not mention leak checking, and lists the unit as a "Spare Unit." The test procedure, T1200347, lists only electronic tests.
NOTE | - The serial numbers and traveller info for the contents of HAM1 should be in JIRA, so that we can trace this sort of thing.
My inclination is to request the following punch list items:
PUNCH | 1) We need a hierarchical list of aLIGO-provided VE modifications, for the components that were actually installed at LLO, listing the makes/models, installation location, drawings, serial number (if applicable), and any test results. Links to the aLIGO are OK, but they need to be described and to refer to specific entries. |  | LINE CLOSED
PUNCH | 2) All gas-containing ISC components should be leak-tested in-house at the next opportunity. The overall plan should be run by the VRB, as was done for the seismic pods. | LINE CLOSED

## Synopsis
DONE | ALL ADDED TO PUNCH LIST
SYNOPSIS FROM DENNIS, BRIAN AND CALUM
CLOSED
LBSC1 / LBSC3
Moved to punch list
LBSC2
Moved to punch list
LBSC4
Moved to punch list
LBSC5
Moved to punch list
LHAM1
no significant concerns, issues or actions
LHAM2
Moved
LHAM3/4
no significant concerns, issues or actions
LHAM5
Moved to punch list
LHAM6
Moved to punch list
VE / DAQ / PSL
No comment.
CLOSED

## Gen Notes
Gen notes from Joe G
All now moved to Punch-list.
This is what I took for Joe's LBSC4,5 input.
CLOSED

## LBSC2 n DAQ
Calum's comments on LBSC2 and DAQ
All now moved to Punch-list (first tab)
CLOSED

## BSC1,3 notes
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
Hi Michael,
PUNCH | Here are my notes and comments on my review of LBSC1-ITMY and LBSC3-ITMX. Generally in good shape. The data is there, although it may require some digging. My contribution for consideration to the punch list consists in adding a task for the ESD system.
NOTE | A reference to the CDS cabling and rack drawings to supplement G1001032 would be a good thing.
Sorry it is so late.
Regards,
Vern
aLIGO Installation Acceptance Review Comments for LBSC1 (Y), E1400179-v3
E1400179-v3 --> 1 Installation Instance/Subset Definition
NOTE | They need a reference in addition to G1001032 to trace signal paths and find engineering level information. Something
like a collection of the CDS engineers’ cable drawings, e.g.,
D0901301 “BSC_ISI” documents the wiring and rack layout for the BSC1,2,3 racks and cabling
D1001725 “ITM_SUS_CONTROLS_WIRING” documents the wiring and rack layouts for the SUS part of the BSCs.
E1400179-v3 --> 2. Procedures --> As-Built/Installed Procedures
NOTE | For accountability and reference, we need to know the names of the people who produced the red-lined documents recording the installation “as-built”, under E1200344-v4 à Other Files: à Redlines on docs used during install (BSC2BSC1INSTALL.pdf, 3.8 MB)
(see https://dcc.ligo.org/DocDB/0093/E1200634/009/E1200634-v9%20LBSC1%20installation%20procedure-v9.pdf )
E1400179-v3 --> 2. Procedures --> As-Built/Aligned Procedure --> E1200640-v7
NOTE | This document is not filled in consistently. There are voids in the data, perhaps because of the history of the difficulty in setting up and performing the measurements. I have to assume the fields that are filled in are the most important and the empty fields are either irrelevant or redundant. An expert may want to review the missing information. Examples: No offset recorded on p. 9. P. 10, Table 6.2.3 is empty. . . . The first entry of a number is on p. 13 for the ITM yaw error. Think about the people reading this a few years from now. The CP-to-ITMY gap distance (p. 14) is filled in. Important measurements are there and are summarized in the alog. I have learned that this document is not the single place to go to for the “final” measurements, but instead the measurements are spread over the alog entries, in E1400179-v3, and in E1200640-v7. The ability to find the entries in the alog, which are typically the final numbers, requires knowing about E1400179-v3 or later. It makes E1400179-v_latest a very important resource. This may be one of the most important things learned during this review, where to find summary information. [I have discussed this point at length to illustrate a common issue with much of the documentation. In practice, it is “good enough”. I think the Livingston documents and the Hanford documents together will provide an adequate record to enable the construction of the third IFO and to document the LHO and LLO instruments as far as we understand them at the time of acceptance. This later may have consequences at a time of discovery in the future.]
E1400179-v3 --> 2. Procedures --> 4 Serial Number Records
CLOSED | ICS, JIRA was unresponsive and I could not verify any ICS records. This also applies to the LBSC3 review.
E1400179-v3 --> 2. Procedures --> 5 Testing
NOTE | Noticed a refinement of procedure as installation and testing progressed. This makes one a little more lenient at the beginning and more critical at the end of the review period.
PUNCH | On the electrostatic drive (ESD): The ESD system is still undergoing installation and should not be evaluated at this time. That said, there are short circuits in the cabling-connectors. Consider elevating this to the “punch list”.
aLIGO Installation Acceptance Review Comments for LBSC3 (X), E1400181-v2
CLOSED | Many of the comments made for LBSC1 also apply here. LBSC3 is in better shape, reflecting the experience of the installation folks.
E1400181-v2 --> 2. Procedures --> As-Built/Installed Procedures
NOTE | Same concern as with LBSC1, the need for names of people making the notes. I don’t mind the informality of a scanned document, it is better to have the information.
NOTE | OpLever final testing remains incomplete (at least as far as information posted to the DCC).