## Slide 1: Progress on Cryogenic Test Masses for aLIGO Upgrades
Brett Shapiro
Stanford University
LIGO
Notes: Don’t hold onto the podium so much
Look around the audience more, don’t just stare at one person
Take a few more moments to explain plots. E.g. this is a plot of the temperature vs time …. The goal is for the test mass to reach the dotted line at 124 Kas fast as possible
Explain the steady state vs rapid cooling operation more clearly

## Slide 2: Summary
Moving from Advanced LIGO to LIGO III
LIGO III cryogenic test mass suspensions
Stanford experiment
LIGO III simulation
Lessons learned / changes made
Future work

## Slide 3: Advanced LIGO Timeline
1st science run
Livingston, LA
Hanford, WA
Sept-Dec 2015
Last core optics
March 2014
Full lock testing
Aug-Sept 2014
Build LIGO III
GW
astronomy
Soon after
Far after

## Slide 4: Predicted Advanced LIGO Sensitivity

## Slide 5: Predicted Advanced LIGO Sensitivity
LIGO III aims to do better by a factor of 3

## Slide 6: Predicted Advanced LIGO Sensitivity
LIGO III aims to do better by a factor of 3
Use cryogenics to push thermal noise down

## Slide 7: LIGO III cryo work distribution
Courtesy of Nicolas Smith-Lefebvre
Caltech - cryogenic reference cavities; direct thermal noise measurements
Jena/Glasgow/Moscow - mechanical loss
MIT – high emissivity coatings
KAGRA – 20 K sapphire suspensions
INPE Brazil – Cryogenic multi-nested pendulum
Stanford – optical coatings (Riccardo Bassiri’s talk); cryogenic technology

## Slide 8: Cyro Test Mass Problem Statement
* For LIGO III, reduce suspension and coating thermal noise by cooling the lower quad to 124 K (-149.15 C)
  Si test masses (blue team in LIGO-T1200031)
  Get to 124 K in a timely manner <- Stanford experiment
  Then maintain 124 K
Include a warm-up scheme (don’t forget!)
Do not increase the test mass lossiness
  Emissive coatings, heat links, thick sus fibers, optical coatings, substrate, suspension fiber bonding, etc
Do not compromise passive seismic isolation
  Cables, hoses, links, etc
The same seismic isolation platforms (ISIs, HEPIs)
  Limit the amount of extra weight on these plaforms
  Leave the rest of the vacuum chamber warm

## Slide 9: LIGO III Steady State Cooling
See also G1200246-v1
124 K
295 K
Heat shield at
80 K maintains 124 K through radiative coupling
Active Seismic Isolation Table
ISI stage 2
Vacuum flange
Flexible liquid nitrogen hoses:
Supply and return
clamp
1560 nm MW laser
≈ 3 W absorption at 124 K
≤ 10 m
143 kg Si test mass
…
…

## Slide 10: Initial Cool Down Cold Link – 2 Designs
return
Test mass
Conductive cooling, low pressure N2 gas
Pros and Cons
liquid N2 pipe flexures
supply
thermally conductive plate with variable gap
Pros
Operates in partial vacuum.
Low heat transfer between cold and warm parts of vacuum system.
Fine temperature control – just back away when at desired temperature
Versatility, design permits both conductive and convective cooling.
Cons
Requires moving parts:
  flexible pipes
  actuators
Physically contacts barrel of test mass

## Slide 11: Initial Cool Down Cold Link – 2 Designs
Pros and Cons
Convective cooling, up to 1 atm N2 gas
Test mass
liquid N2
pipes
thermally conductive plate with large fixed gap
return
supply
convective N2 gas
Pros
No moving parts or actuators
No contact with test mass
Faster cooling than conduction
Cons
Convection between cold and warm parts of vacuum system
No fine temperature control – must return to vacuum to ‘turn off’ cold link.
Does not operate under vacuum

## Slide 12: Experimental Setup
Test dewar
Heat shield

## Slide 14: Close up of cold link
Test
mass
Cold link
Flexure ‘point’ contacts
Variable gap

## Slide 15: Measurement – cold link engaged
140 min time constant

## Slide 16: Test mass temperature modeling

## Slide 17: Test mass temperature modeling

## Slide 18: Test mass temperature modeling
Heat transfer goes through residual N2 gas between test mass and cold link!

## Slide 19: Test mass temperature modeling
Heat transfer goes through residual N2 gas between test mass and cold link!
This is a perfect exponential decay…

## Slide 20: Exponential Temperature Decay
These are both functions of temperature.
In general, the solution is not truly exponential since the time ‘constant’ changes.

## Slide 21: Exponential Temperature Decay

## Slide 22: Exponential Temperature Decay

## Slide 23: Exponential Temperature Decay

## Slide 24: Measurement – cold link disengaged
160 min time constant

## Slide 25: Measurement – cold link disengaged
160 min time constant
Simulation assuming only N2 gas conduction

## Slide 26: Conductive vs Convective Regimes

## Slide 27: Finite Element Modeling
Images courtesy of Tim MacDonald
≈ 43 min into cool down
Cold link
Test mass
N2 gas
Due to complexity, LIGO III designs must be verified with FEM
Below: FEM of conduction through N2 gas to cold link for Stanford experiment
Convective FEM is proving to require large amounts of computing power

## Slide 28: Cold link on a LIGO III test mass
Test mass radius = 280 mm
Test mass thickness = 250 mm
Cold link covers 2/3 of top barrel
Conduction limit, no convection

## Slide 29: Experimentation lessons learned
Air dominates most heat flow across contacts
Cold links should have distributed contacts
Solder is not leak tight against high pressure cryogenic fluids – welding is probably best
Cryogenic fluid should have 1 flow path
Send fluid from bottom up
Use fatter pipes to minimize fluid pressure
Minimize the number of materials in the plumbing – joining and contraction issues
Leave room for differential contraction
Silicon diode temperature sensors are great

## Slide 30: Flexible cold link evolution

## Slide 31: Future work
Get bigger pump to reduce pressure
Better computing for FEM of convection
Next generation experiment on the prototype aLIGO in-vacuum seismic isolation table.
  More realistic LIGO setup
  Measure temperature drifts on LIGO hardware
  Measure seismic noise of nitrogen delivery
  Use welded joints to prevent leaks.
LIGO

## Slide 32: LVC STANFORD August 25-29 2014

## Slide 33: Backups
LIGO

## Slide 34: Advanced LIGO Layout
FI
Signal Recycling Cavity
ITM
ETM
Input
Mode
Cleaner
Output
Mode
Cleaner
Power Recycling
BS
4 km
Laser
Φ
m
PD
GW readout
FI
ITM
ETM
125 W
5.7 kW
815 kW
CP
Installation complete in 2014
Observing GWs in 2015

## Slide 35: Possible LIGO III Mechanical Upgrades
10 kg
20 kg
Main
chain
Reaction
chain
143 kg
Main
chain
Reaction
chain
Advanced LIGO quad pendulum
Preliminary LIGO III quad pendulum
2.14 m
1.626 m
40 kg
20 kg
40 kg
20 kg
Fused silica, 295 K
Silicon, 120 K
? kg
? kg
? kg

## Slide 36: 3 Quad Conceptual Designs
T1300786
Higher payload
Lighter
Test mass
Ideal masses with PUM springs

## Slide 37: How to get a LN2 Hose to ST2
Stage 0
Stage 1
Stage 2
Stage A
LN2 hoses: Supply and return
Actuators
Sensors
Extra stage, A, in parallel with stage 1 carries hose. Stage A is actuated to follow stage 2 so the hose has does not short seismic isolation. Stage A sensor noise is set by the stage 2 isolation requirement (so it follows stage 2 and not the sensor noise).
Vacuum flange
clamps

## Slide 38: Test Mass Radiation Simulation
Ideal radiation shield:
Fully encloses test mass
77 K (liquid N2)
Shield emissivity = 1
Test mass barrel emissivity = 0.5
Laser power = 0
Power α (Temperature)^4

## Slide 39: A Lot of Heat to Remove
60 W
143 kg Test Mass:
14 MJ to get cold (warm)
2.7 days to 120 K (295 K) at a constant 60 W
Maybe there is another way to cool the test mass quickly

## Slide 40: Dewar pressure during measurements

## Slide 41: Effect of pressure on test mass temp

## Slide 43: N2 gas therm. cond. vs pressure

## Slide 44: Temperature Sensor Locations

## Slide 45: Test Mass Temperature Equations

## Slide 46
ref: Harry, Bodiya, Desalvo. Optical Coatings and Thermal Noise in Precision Measurement. 2012. pg 113.

## Slide 47: Si CTE vs Temperature
Thermoelastic component of thermal noise goes to zero with CTE.

## Slide 48: Si Thermal Conductivity vs Temp.

## Slide 49: Si Specific Heat vs Temperature

## Slide 50: Thermal Conductivity of Materials

## Slide 51: Beam Tube Heat Shield Length
ε
Beam tube
Blackbody disc model
Heat Shield
rj = 0.28m
ri = 0.6225m
Test Mass
L (m)
What is the optimal heat shield length in the beam tube?
Tj = 120 K
Ti = 295 K
T = 77 K
Ε = 1
Ε = 0.1
See T1300556

## Slide 52: Heat shield length in beam tube
Beam tube shield length L (m)
Black body power from beam tube arriving at test mass
≈1/3 drop for every factor of 2 L
See T1300556
Test mass emissivity = 0.1

## Slide 53: Other Problems To Solve
Liquid N2 hoses flexible enough for ISI under vacuum
Temperature/height control of blade springs
Test mass temperature control
How to measure temperature?
  Measure acoustic modes – Young’s modulus is temp. dependent
  Infrared camera
Emissivity of optical coatings
Lossiness of emissive coatings
Good emissivity estimates/measurements of Si?
Power absorption in Si (ppm, W, etc)?
How noisy is bubbling nitrogen: seismic, Newtonian? Do boiling chips help?
Optical coating thermal noise at 120 K