## Sheet1
E1400204-v4
IAS LLO Primary Optics As Built Alignment Summary
Optic | X Axis (mm) |  |  |  | Y Axis (mm) |  |  |  | Z Axis (mm) |  |  |  | Pitch (µrad)1 |  |  |  | Yaw (mrad)2
| Actual | Target | Error | Tolerance | Actual | Target | Error | Tolerance | Actual | Target3 | Error | Tolerance | Actual | Target | Error | Tolerance | Actual | Target | Error | Tolerance
BS | -203 | -202.5 | -0.5 | ±3.0 | -183.3 | -184.2 | 0.9 | ±1.0 | -83.3 | -82.7 | -0.6 | ±1.0 | 237 | 212 | 25 | ±55 | 784.821 | 784.983 | -0.162 | ±0.190
ITMx | 5015 | 5013 | 2 | ±3.0 | -199.2 | -200 | 0.8 | ±1.0 | -82 | -81.4 | -0.6 | ±1.0 | 241 | 312 | -71 | ±1324 | 0.015 | 0 | 0.015 | ±0.1834
CPx5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1505 | 1534 | -29 | ±1400 | 0.044 | 0 | 0.044 | ±1.400
ITMy | -199.4 | -200 | 0.6 | ±1.0 | 4983.3 | 4983.1 | 0.2 | ±3.0 | -82.5 | -83 | 0.5 | ±1.0 | 615 | 611 | 4 | ±1324 | 0.072 | 0 | 0.072 | ±0.1834
CPy5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 959 | 1833 | -874 | ±1400 | 0.929 | 0 | 0.929 | ±1.400
PRM | -20187.5 | -20189.6 | 2.1 | ±3.0 | -627.2 | -628.1 | 0.9 | ±2.0 | -95.9 | -96.7 | 0.8 | ±1.0 | -244 | -269 | 25 | ±100 | 5.831 | 5.876 | -0.045 | ±0.100
PR2 | -3578.7 | -3579.2 | 0.5 | ±3.0 | -530.1 | -530.5 | 0.4 | ±1.0 | -91.6 | -92.3 | 0.7 | ±1.0 | 50 | 0 | 50 | ±100 | -7.996 | -8.011 | 0.015 | ±0.100
PR3 | -19740.9 | -19740 | -0.9 | ±3.0 | -175.7 | -176.5 | 0.8 | ±3.0 | -90.1 | -87.9 | -2.2 | ±3.0 | 50 | 0 | 50 | ±50 | -11.156 | -11.146 | -0.01 | ±0.050
SRM-s6 | 306.9 | 305 | 1.9 | ±4.6 | -19909.8 | -19908.7 | -1.1 | ±3.0 | -81.7 | -81.1 | -0.6 | ±3.0 | 27 | 30 | -3 | ±520 | 57.098 | 57.094 | 0.004 | ±0.820
SR2 | -596.4 | -594.1 | -2.3 | ±5.1 | -4177 | -4178.2 | 1.2 | ±3.0 | -83.4 | -81.6 | -1.8 | ±3.4 | -155 | 0 | -155 | ±520 | 42.358 | 42.108 | 0.25 | ±0.820
SR3 | -175.3 | -175.3 | 0 | ±3.0 | -19614.9 | -19616 | 1.1 | ±3.0 | -80.9 | -82.1 | 1.2 | ±3.0 | 8 | 0 | 8 | ±560 | 13.762 | 13.783 | -0.021 | ±0.200
ETMx | 3999499.2 | 3999498 | 1.2 | ±3.0 | -200.6 | -200 | -0.6 | ±1.0 | -80.5 | -80 | -0.5 | ±1.0 | 323 | 315 | 8 | ±50 | -0.005 | 0 | -0.005 | ±0.050
ERMx5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1716 | 1711 | 5 | ±1470 | -0.03 | 0 | -0.03 | ±1.470
ETMy | -200.5 | -200 | -0.5 | ±1.0 | 3999468.3 | 3999468.1 | 0.2 | ±3.0 | -80.25 | -80 | -0.25 | ±1.0 | -1 | 19 | -20 | ±50 | 0.02 | 0 | 0.02 | ±0.050
ERMy5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 660 | 1241 | -581 | ±1470 | 0.893 | 0 | 0.893 | ±1.470
1: Pitch sign convention follows the standard SUS convention for pitch, i.e. positive = down, negative = up (see E1000617, E1100108, & E1100109)
2: Yaw sign convention follows the standard SUS convention for yaw, i.e. positive = counterclockwise, negative = clockwise (see E1000617, E1100108, & E1100109). Yaw is reported as the angle from the axis closest to the HR surface normal of the optic
3: All Z axis target positions are reported in coordinates local to the building that houses the optic. I.E. the BS Z axis is local to the LLO corner station, while the ETMx Z axis is local to the LLO End X station.
4: This is the pitch/yaw tolerance for L1 ITMx and ITMy at the time of install. This was changed to ±50 µrad after these 2 optics were installed and aligned, hence the difference in spec between ITM and ETM
5: SUS sets the position of the CPx, CPy, ERMx, and ERMy. IAS is only repsonsible for aligning the pitch and yaw of these optics so that is all that is recorded here. The pitch/yaw of these optics is set so the optic is parallel with the AR surface of their associated TM.
6: This is the SRM surrogate optic that is being used in the early low power stages of aLIGO commissioning. The real SRM optic will be installed and aligned at a later date.