LIGO Laboratory / LIGO Scientific Collaboration

LIGO- T1400393 V5	                 LIGO                              September 18th, 2014

L1 Optical Lever Commissioning and Testing

T1400393 – V5

Mike Vargas, Thomas Vo

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Scope and Introduction

This document covers the methods and results of each L1 aLIGO optical levers testing and commissioning phase.  All measurements to date are done with HEPI floating and ISI damping for whenever applicable, if HEPI is  not  yet  commissioned,  then  it  will  be  locked  prior  to  the  beginning  of  the  measurement. Additionally, all available suspension damping loops will be engaged wherever applicable. These results will serve the purpose of verifying that each optical lever is passing specs outlined in  E1200719-v1 as well as letting the end user know the characteristics of each lever such as the linear range and structural resonances. If an optical lever does not meet spec, changes may be made to the physical system (mechanical, electronics, software, etc) in order to comply with requirements.  Those changes to the system will require changes to this document (new graphs or data).

Required attributes that will deem an optical lever commissioned and passed testing phase:

Calibration of the optical levers into micro radians.

Beam profiles in vertical and horizontal directions at the QPD. The beam scan used for taking this data will be a THORLABS BP209-VIS.

Calibrated spectra comparing the Optical Lever signals to OSEMS or ISI position sensors.

Longitudinal to angular coupling coefficients (calculated via T1300130):

o dɵ'/dx (yaw)   = ds' (yaw) / (2 L) = (1 / L) sin(ɵ) cos(φ)

o dφ'/dx (pitch) = ds' (pitch) / (2 L) = (1 / L) sin(ɵ) sin(φ)

Long-term drift testing results.

Dark current noise floor measurement.

Since the HAM optical levers do not have a translation stage with a micrometer, we need to find a separate method to measure the linear response curve.

The process for HAM and BS OpLev calibration will have two varieties:

The first way we can take the calibration data from another optical lever and scale the geometry, beam power, and beam width so that it matches HAM2. The geometry scaling can be derived from the Zeemax layouts and the lever arm is directly proportional to the sensitivity of the optical lever. The beam power for the optical levers can be found by looking at a time average of the sum of counts. The beam widths can be found by looking at the beam profile at the QPD. Both the power and the widths are linearly related to the sensitivity so they can be scaled simply as well. A potential source of error would associated with the performance of the QPDs on an individual basis, these will be tested to look at the variance of power and sensitivity in order to create and error via this method.

A second way to map the linear response curve is to drive the ISI and correlate the optical lever pitch and yaw signals to the position sensors (RX, RY, and RZ). This requires a set actuation at a few different amplitudes to prove linearity in the actuators as well as a coherence function to show that we see good correlation. For HAM2, the RY sensors will correspond to the pitch OptLev signals and the RZ sensors will correspond to the yaw OptLev signals.

Summary:

|  | Target | Laser SN | Profiled | Calibrated | Issues |
| --- | --- | --- | --- | --- | --- |
| L1 | HAM2 | - | No | No | Requires non-glitching laser. |
| L1 | HAM3 | 136 | No | No | Requires non-glitching laser. |
| L1 | HAM4 | - | No | No | Requires non-glitching laser. |
| L1 | HAM5 | - | No | No | Requires non-glitching laser. |
| L1 | BS | 186 | No | No | Requires non-glitching laser. |
| L1 | ITMX | 192 | Yes | Yes | Beam size is larger than required and will be adjusted. |
| L1 | ITMY | 190 | Yes | Yes | Beam size is slightly larger than required. Adjustment possibly needed. |
| L1 | ETMX | 125 | No | Yes | Beam scan reveals prominent rings. Telescope adjustment needed. Requires non-glitching laser. |
| L1 | ETMY | 139 | Yes | No | Calibration values are ~100, a problem is suspected. Requires non-glitching laser. |
| L1 | PR3 | 143 | Yes | Yes | Requires non-glitching laser. |
| L1 | SR3 | 187 | No | Yes | Requires non-glitching laser. |

1   ITMY

|  | Gaussian_Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch | 4.22 | 23.88 |  | 0.56 |  |
| Yaw | 4.18 | 23.98 |  | 0.30 |  |

1.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:  23.88 [uradians/counts]

Yaw:  23.98 [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-ITMY_L3_OPLEV_PIT_GAIN

L1-SUS-ITMY_L3_OPLEV_YAW_GAIN

1.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) | 4.22 | 95.3 |
| Y-Axis (right side) | 4.18 | 94.3 |

1.3  Longitudinal to Angular  Coupling

- L = 66.1207 m

- ɵ =  1.22 degrees

- φ  = 59.62 degrees

dɵ'/dx = 0.56 urad/m

dφ'/dx = 0.3  urad/m

1.4  Calibrated  Spectra

1.5  Calibrated  Long Term Drift

1.6  QPD Sensor Dark Noise

2   ETMY

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch | 2.44 |  |  | 20.91 |  |
| Yaw | 2.36 |  |  | 11.9 |  |

2.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-ETMY_L3_OPLEV_PIT_GAIN

L1-SUS-ETMY_L3_OPLEV_YAW_GAIN

2.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) | 2.44 | 90.9 |
| Y-Axis (right side) | 2.36 | 91.6 |

2.3  Longitudinal to angular coupling

- L = 11.16 m

- ɵ =  7.65 degrees

- φ  = 60.25 degrees

dɵ'/dx = 20.91 urad/m

dφ'/dx = 11.90 urad/m

2.4  Calibrated  Spectra

2.5  Calibrated  Long Term Drift

2.6  QPD Sensor Dark Noise

3   ITMX

|  | Gaussian_Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch | 5.33 | 33.27 |  | 0.56 |  |
| Yaw | 5.37 | 32.88 |  | 0.30 |  |

3.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:  33.27 [uradians/counts]

Yaw:   32.88 [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-ITMX_L3_OPLEV_PIT_GAIN

L1-SUS-ITMX_L3_OPLEV_YAW_GAIN

3.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) | 5.33 | 93.5 |
| Y-Axis (right side) | 5.37 | 93.1 |

3.3  Longitudinal to Angular  Coupling

- L = 66.1147 m

- ɵ =  1.22 degrees

- φ  = 59.62 degrees

dɵ'/dx = 0.56 urad/m

dφ'/dx = 0.3  urad/m

3.4  Calibrated  Spectra

3.5  Calibrated  Long Term Drift

3.6  QPD Sensor Dark Noise

4   ETMX

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  | 46.51 |  | 20.89 |  |
| Yaw |  | 46.73 |  | 11.9 |  |

4.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:  46.51 [uradians/counts]

Yaw:   46.73 [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-ETMX_L3_OPLEV_PIT_GAIN

L1-SUS-ETMX_L3_OPLEV_YAW_GAIN

4.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

4.3  Longitudinal to angular coupling

- L = 11.162 m

- ɵ =  7.65 degrees

- φ  = 60.25 degrees

dɵ'/dx = 20.89 urad/m

dφ'/dx = 11.90 urad/m

4.4  Calibrated  Spectra

4.5  Calibrated  Long Term Drift

4.6  QPD Sensor Dark Noise

5   Beam Splitter

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  |  |  | 0.00 |  |
| Yaw |  |  |  | 26.3 |  |

5.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.

These results are inputted into the SUS gain filters:

L1-SUS-BS_L3_OPLEV_PIT_GAIN

L1-SUS-BS_L3_OPLEV_YAW_GAIN

5.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

5.3  Longitudinal to angular coupling

- L = 4.021 m

- ɵ =  3.03 degrees

- φ  = 0 degrees

dɵ'/dx = 0.00 urad/m

dφ'/dx = 26.3 urad/m

5.4  Calibrated  Spectra

5.5  Calibrated  Long Term Drift

5.6  QPD Sensor Dark Noise

6   PR3

|  | Gaussian_Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch | 3.09 | 35.86 |  | 4.59 |  |
| Yaw | 3.18 | 33.54 |  | 0.00 |  |

6.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:  35.86 [uradians/counts]

Yaw:  33.54 [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-PR3_L3_OPLEV_PIT_GAIN

L1-SUS-PR3_L3_OPLEV_YAW_GAIN

6.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) | 3.09 | 90.27 |
| Y-Axis (right side) | 3.18 | 88.16 |

6.3  Longitudinal to Angular  Coupling

- L = 27.726 m

- ɵ =  3.64 degrees

- φ  = 89.99 degrees

dɵ'/dx = 4.59 urad/m

dφ'/dx = 0.00 urad/m

6.4  Calibrated  Spectra

6.5  Calibrated  Long Term Drift

6.6  QPD Sensor Dark Noise

7   SR3

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  | 51.20 |  | 2.48 |  |
| Yaw |  | 53.11 |  | 0.00 |  |

7.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:  51.20 [uradians/counts]

Yaw:  53.11 [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.  A DC method was used to obtain the y-axis data; we first translate the QPD a fixed amount as measured with a micrometer and use the geometry of the setup to correlate an angular deviation to the pitch or yaw signals.

These results are inputted into the SUS gain filters:

L1-SUS-SR3_L3_OPLEV_PIT_GAIN

L1-SUS-SR3_L3_OPLEV_YAW_GAIN

7.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

7.3  Longitudinal to angular coupling

- L = 27.50 m

- ɵ =  1.95 degrees

- φ  = 89.94 degrees

dɵ'/dx = 2.48 urad/m

dφ'/dx = 0.00 urad/m

7.4  Calibrated  Spectra

7.5  Calibrated  Long Term Drift

7.6  QPD Sensor Dark Noise

8   HAM2

|  | Gaussian_Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  |  |  | 0.02 |  |
| Yaw |  |  |  | 0.20 |  |

8.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.

These results are inputted into the SUS gain filters:

L1-ISI-HAM2_L3_OPLEV_PIT_GAIN

L1-ISI-HAM2_L3_OPLEV_YAW_GAIN

8.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

8.3  Longitudinal to Angular  Coupling

- L = 26.643 m

- ɵ =  0.19 degrees

- φ  = 3.58 degrees

dɵ'/dx = 0.02 urad/m

dφ'/dx = 0.20 urad/m

8.4  Calibrated  Spectra

8.5  Calibrated  Long Term Drift

8.6  QPD Sensor Dark Noise

9   HAM3

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  |  |  | 0.02 |  |
| Yaw |  |  |  | 0.30 |  |

9.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.

These results are inputted into the SUS gain filters:

L1-ISI-HAM3_L3_OPLEV_PIT_GAIN

L1-ISI-HAM3_L3_OPLEV_YAW_GAIN

9.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

9.3  Longitudinal to angular coupling

- L = 26.580 m

- ɵ =  0.22 degrees

- φ  = 3.01 degrees

dɵ'/dx = 0.02 urad/m

dφ'/dx = 0.30 urad/m

9.4  Calibrated  Spectra

9.5  Calibrated  Long Term Drift

9.6  QPD Sensor Dark Noise

10   HAM4

|  | Gaussian_Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  |  |  | 0.01 |  |
| Yaw |  |  |  | 0.30 |  |

10.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.

These results are inputted into the SUS gain filters:

L1-ISI-HAM4_L3_OPLEV_PIT_GAIN

L1-ISI-HAM4_L3_OPLEV_YAW_GAIN

10.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

10.3  Longitudinal to Angular  Coupling

- L = 26.559 m

- ɵ =  0.2 degrees

- φ  = 3 degrees

dɵ'/dx = 0.01 urad/m

dφ'/dx = 0.30 urad/m

10.4  Calibrated  Spectra

10.5  Calibrated  Long Term Drift

10.6  QPD Sensor Dark Noise

11   HAM5

|  | Gaussian Width (mm) | Calibration (Counts/urad) | Low- Frequency Resonances (Hz) | Long-to-ang coupling coefficient (urad/m) | Peak-to-peak drift (uradians/10 minutes) |
| --- | --- | --- | --- | --- | --- |
| Pitch |  |  |  | 0.01 |  |
| Yaw |  |  |  | 0.20 |  |

11.1  Calibration Parameters

The slope of the linear region of the graph for pitch and yaw:

Pitch:   [uradians/counts]

Yaw:   [uradians/counts]

The geometric calculation for these parameters uses a Zeemax calculation for the path lengths of the lever arms from  E1200836--v1.

These results are inputted into the SUS gain filters:

L1-ISI-HAM5_L3_OPLEV_PIT_GAIN

L1-ISI-HAM5_L3_OPLEV_YAW_GAIN

11.2  Beam Profile

|  | Gaussian Width (mm) | Gaussianity (%) |
| --- | --- | --- |
| X-Axis (left side) |  |  |
| Y-Axis (right side) |  |  |

11.3  Longitudinal to angular coupling

- L = 26.760 m

- ɵ  =  0.18 degrees

- φ  = 2.97 degrees

dɵ'/dx = 0.01 urad/m

dφ'/dx = 0.20 urad/m

11.4  Calibrated  Spectra

11.5  Calibrated  Long Term Drift

11.6  QPD Sensor Dark Noise