## Slide 1: Optical lever lasers - To Fix or Not to Fix!!
Suresh Doravari
LIGO-G1500071

## Slide 2
Optical Lever Lasers -
To fix or not to fix!
What do we mean by better performance?
Do we need better performance?
How do we obtain better performance?
A comparison of diode lasers and HeNe lasers
Options ahead of us
A possible plan of action

## Slide 3
Our current oplev lasers can perform ten times better
Oplev laser power fluctuations : factory spec RIN is 1% , after the fix < 0.1%

## Slide 4
Lower noise across the entire spectrum by an order of magnitude

## Slide 5
Spectrograms show an absence of glitches after the fix.

## Slide 6
School One
The oplevs are used to obtain lock and then the control is
transferred to wavefront sensors.
So glitchy lasers are okay.
School Two
When the weather gets bad
    High wind conditions in LHO
    Storms in LLO
We will need to engage oplev damping even during lock.
Even obtaining lock may be difficult if the oplev noise is high.
So we do need lasers which do not glitch.
Do we need better performance?

## Slide 7
An example of how laser power glitches kick the optic:

## Slide 8
How can we obtain better performance?
  Fix the ones we have
    Tweak alignment to reduce optical feed back
    Tweak operating temperature set point
    Choose power level for glitch free operation
    Reduce fan speed
    Passive thermal shielding
  Replace them with
    Better diode lasers
      Get the present supplier to deliver glitch free lasers
      Source from Thorlabs or another such standard company
        (will require some local tuning for glitch free operation)
    He-Ne lasers
      Fiber fed He-Ne laser with stabilisation

## Slide 9
Comparison between He-Ne and Diode lasers

## Slide 10
Comparison of He-Ne and diode lasers : Spectra

## Slide 11
Comparison between He-Ne and Diode lasers : Spectrograms

## Slide 12
Options ahead of us
  Do nothing and live with the current glitchy lasers
  Train the current supplier to build glitch free lasers
  Obtain better engineered diode lasers another well established supplier (Thorlabs for eg.) and tune them for glitch free performance our selves
  Switch over to He-Ne lasers (either fiber coupled or free space) and do the work needed to stabilize them

## Slide 13
A plan of action
Fix the diode lasers currently with us and use them till their end-of-life (at least two years)
In the mean time explore and test other vendors and lasers
  Diode lasers from Thorlabs
  Stabilised HeNe Lasers with fibers couplers from Newport
  (Both these will require some in-house work for stable operation)
In a couple of years time replace the diode lasers as they die with the new solution