LIGO HANFORD OBSERVATORY

PO BOX 159

RICHLAND WA 99352

TEL: 509.372.8106

FAX: 509.372.8137

# MEMORANDUM

DATE:		August 26, 2015

| TO: | ISC team |
| --- | --- |
| FROM: | Daniel Sigg, Koji Arai |
| SUBJECT: | EOM/AOM Diver Modifications |
| Refer to: | LIGO-E1500353-v1 |

This modifications apply to D0900761-B, D0900847-B, D0900848-C and D1000216-B to bring them to revisions C, C, D and C, respectively.

The affected serial numbers are S1500117 to S1500128.

Power Board D0900848-C

Change 1:

Capacitors need to be added to prevent the AD829 driving the pass transistor from oscillating.

C20, C35, C36, C37, C50, C51, C52, and C63 → 1nF

C91 → add a 100pF between U6 op27 (PIN6) and VREFP (PIN7) 
	 should fit between the pads of C22 and R32

W1 → remove solder jumper

Change 2:

Remove the protection diodes which are no longer needed with the use of the power sequencing relays

D1, D2, D3, D4 → remove

Change 3:

Fix the gain in the differential output stages.

R66, R68 → 5.49K, shift and use only the pad above R70 and R74

J1, J2 → add a jumper wire from the free end to ground (see pic)

Servo Board D0900847-B

Change 1:

Reduce the high frequency gain of the servo to allow for better phase margin.

R69 → 200

C49 → 1nF

R71 → 3.3K

R67 → 3.3K

Change 2:

Add capacitor at the output drive to ground to reduce the backfeed of RF from the attenuator board.

C58 → 1nF, add between signal and ground pin of SMA connector J5

Attenuator Board D1000216-B

Change 1:

Add capacitor at the control input to ground to reduce the backfeed of RF to the servo board.

C2 → 1nF, add between signal and ground pin of SMA connector J3

Controller Board D0900761-B

No change.

BOM (for 12 units):

| Qty | Item | Distributor | Description |
| --- | --- | --- | --- |
| 108 | 311-1122-1-ND | Digi-Key | 1nF, 0805, NPO Power board: C20, C35, C36, C37, C50, C51, C52, and C63 Servo board: C49 |
| 12 | 311-1111-1-ND | Digi-Key | 100pF, 0805, NPO Power board: C91 |
| 24 | P5.49KDACT-ND | Digi-Key | 5.49K, 0805, 0.1% Power board: R66, R68 |
| 12 | P200DACT-ND | Digi-Key | 200, 0805, 0.1% Servo board: R69 |
| 24 | P3.3KDACT-ND | Digi-Key | 3.3K, 0805, 0.1% Servo board: R67, R71 |
| 24 | 478-1492-1-ND | Digi-Key | 1nF, 1206, NPO Servo board:C58 Attenuator board: C2 |
| 1 | 6710 BK005-ND | Digi-Key | Hook-up wire |