## Slide 1: Improving seismic isolation in Advanced LIGO using a ground rotation sensor
04/16/2016
Krishna Venkateswara
for
UW- Michael Ross, Charlie Hagedorn, and Jens Gundlach
aLIGO – SEI team
LIGO-G1600083

## Slide 2: Contents
Introduction: Why tilt affects LIGO
BRS at LIGO Hanford and Tilt Data
Next Gen/Compact BRS
Summary

## Slide 3: Introduction

## Slide 4: Tilt-problem limits duty cycle of detectors
Earthquakes, High Wind and High Microseism cause significant down-time of both detectors.
All are primarily related to the problem of tilt-horizontal coupling producing too much ISI motion.
Source: B. Lantz T1500610
Source: S. Dwyer G1501365
Source: DetChar Summary Pages

## Slide 5: Tilt versus Horizontal displacement
Conventional seismometers and tiltmeters cannot differentiate between horizontal displacement and ground tilt.
x  ax
x  g
Tilt response to horizontal displacement response for all seismometers = -
Tilt is confused with horizontal motion at low frequencies (below ~ 0.1 Hz).
Solution: Inertial rotation sensors, Tilt-free seismometers or Ring-laser gyroscopes…

## Slide 6: Beam Rotation Sensors

## Slide 7: BRS Concept
Principle:
Ground tilt is measured by measuring angle between ground and low frequency beam balance.
Horizontal acceleration can be rejected by locating center of mass at the pivot.
Autocollimator
Venkateswara, Krishna, et al. "A high-precision mechanical absolute-rotation sensor." Review of Scientific Instruments 85.1 (2014): 015005.

## Slide 8: BRS Concept
Principle:
Ground tilt is measured by measuring angle between ground and low frequency beam balance.
Horizontal acceleration can be rejected by locating center of mass at the pivot.
Autocollimator
Venkateswara, Krishna, et al. "A high-precision mechanical absolute-rotation sensor." Review of Scientific Instruments 85.1 (2014): 015005.

## Slide 9: BRS Parameters
10-25 m-thick Cu-Be Flexures
Mass | 4-5 kg
Length | 0.9 m
Moment of Inertia I | 0.6 kg m2
Flexure stiffness | 10-3 N m
Quality factor | 2000-3000
Distance between pivot and center of mass | Can be tuned to less than 1 µm
Autocollimator
Displacement rejection ratio

## Slide 10: Installation at EX VEA
Ref: LHO 13250

## Slide 11
Ground tilt during low wind speeds
Requirement
Autocollimator sensitivity is ~0.1 nrad/Hz

## Slide 12
Ground tilt during 20-30 mph winds
Requirement
High coherence between ground seismometer and BRS

## Slide 13
Second example
Requirement
Primary microseism is visible after tilt-subtraction!

## Slide 14
Nominal Config
Windy Config
Order of magnitude less RMS motion
Sacrifice only 2-3 here
Same at QUAD Resonances and above
Improving Isolation
Source: J. Kissel, G1500475

## Slide 15
Ground tilt from M6.7 Earthquake near Australia
ASD of ground seismometer and BRS Data.
Band-passed time-series signals

## Slide 16: Next Gen/Compact BRS

## Slide 17
Compact-BRS
New features
1. Cross Shape (~0 quadrupole moment) ensures first order insensitivity to gravity gradient noise.
2. New compact interferometric readout with ~10X better sensitivity.
30 cm

## Slide 18
Schematic

## Slide 19
Readout sensitivity
Readout with “hard control” Piezo – stacks locked at ~ 5 Hz
Tilt Signal is in the drive
Readout with “soft control” Piezo – stacks locked at ~ 5 mHz
Tilt signal is in the PhotoDiodes
~ 3 pm/rt(Hz)
Individual interferometer limited by frequency noise. Differential signal (angle channel) reduced frequency noise by matched cavities.

## Slide 20
1. BRS allows an unambiguous separation of tilt and horizontal acceleration between 0.01 to ~30 Hz.
2. BRS-1 and BRS-2 can improve total RMS differential motion of the Optical Platforms in LIGO and improve isolation, especially under windy conditions.
3. A new compact prototype (cBRS) is under development. When placed on the Optical Platforms, it can significantly boost angular isolation. It may also prove useful in Newtonian Noise measurement at 10 Hz.
Summary

## Slide 21: Thank you!

## Slide 22: Extra slides

## Slide 23: Target sensitivity
from “Requirements for a ground rotation sensor to improve Advanced LIGO” (B. Lantz et al., 2009)

## Slide 24: Damper Installation
A turn-table applies a gravitational torque on the balance to actively damp the beam-balance when amplitude exceeds a threshold.
gravitational feedback system!
Ref: LHO 14388

## Slide 25
BRS-2 Assembly pictures
Assembled beam and weights
Flexures
Integrated with vacuum chamber

## Slide 26
BRS-2 Assembly pictures
Capacitor plates
Beckhoff modules
Basler CCD camera
Foam box

## Slide 27
Sensitivity comparison
Current BRS read-out noise
Potential c-BRS read-out noise
Source: J. Warner, LHO 17197
Frequency (Hz)
Angle ASD (rad/rt(Hz))

## Slide 28
cBRS Tilt data
Current floor is too noisy to assess performance well
Frequency noise is consistent with before
Expectation of ‘sensitivity’ with 10% matching