## Slide 1: Mirror scattering loss analysis by integrating sphere measurement
Scattering and loss by test mass
  Discrepancy between the measured arm loss, 50ppm/mirror, and the loss based on optics measurement, 25ppm/mirror
  Loss by mirror is estimated as
    Loss by surface figure = 10ppm,
    Larger angle scattering = 10ppm = 5ppm by micro roughness + 5ppm by point scattering,
    Misc and uncertainty = sevetal ppm
Large angle scattering measurement covers down to 1°. What if TIS(θ<1°) is comparable to TIS(θ≥1°)
Integrating sphere measurement with extension to measure TIS in large and small angle (L. Zhang)
  Preliminary and proof of concept
  Small angle (θ≤1°) and large angle (θ≥1°) scattering
    Missing energy in the small angle scattering
    Information defect size/distribution
  Micro roughness on coated surface is correlated to uncoated surface

## Slide 2: Three scattering sources Figure error, micro roughness and point scattering
 | Figure error | Micro roughness | Point scattering
λspatial | > 5mm | < 5mm | < 0.1mm
data | phasemap | PSD | data
cause | Reflect back into cavity | Reflect out of cavity to a fixed direction | Reflect uniformly, mostly out of cavity
effect | Change of resonating mode | Hit specific part of IFO, e.g., beam tube baffle. | Additional loss
16 cm
0.2 mm
ETM07
coated
polished

## Slide 3: aLIGO ETM TIS(θ≥1°) using integrating sphere
W=150μm
θ≥1°
95%
Micro roughness
Point scattering

## Slide 4: Micro roughness : coated ~ polished low end of TIS(θ≥1°) ~ 3-4 x polished surface
ETM16 case
peak
TIS(λs<75μm) / (4π σ(λs<80μm)/λlaser)2
80μm > λs > 0.9μm
Zygo 50x PMM
1/1000
1/10
~2
PSD
1mm
1mm
Around the peak

## Slide 5: Angular distribution of reflected field depends on the defect size
θ
Mie
Geometry
r
r
√π r
3% of solid angle
Larger defects induce sharper reflection

## Slide 6: Far field and small size defects small defects cannot be characterized by PSD
reflected power = power density at defect ∙ defect size ∙ |A|2
circle with radius a
rectangle size of du x dv
axis : k a θ = 0.1 a(μm) θ(degree)
log( power )
small
k a θ =1 when
a = 1μm, θ=10°
or
a=10μm, θ=1°
Defect size looks different seen at different angle

## Slide 7: Clustered defects behave like a single defect
Power distribution generated by small defects in a square
Same distribution generated by

## Slide 8: Setup for measuring TIS(θ≤1°) and TIS(θ≥1°)
150μm

## Slide 9: Very preliminary results mirror : initial LIGO ETM04
TIS (0.7°≤ θ≤1°)
TIS (θ≥1°)
scratch
dΩ=1.5e-4

## Slide 10: TIS(θ≤1°) vs TIS(θ≥1°) very preliminary
Quadrant with scratch
Quadrants no scratch
(2) Scratch
Large defect on surface
(4) Large defect
(3) Clustered small defects
Large TIS(θ≥1°), small TIS(θ≤1°)
(1) Isolated small defect or micro roughness
brighter
total size>1.5 μm

## Slide 11: Revised setup
Hole in the mirror
θ > 0.7°
θ > 0.3°
Lens moved out of the path to reduce noise
Beam diverging toward the second mirror,
which induces larger tail noise of the undisturbed beam.

## Slide 12: Comparison of results : PD2(0.7°≤ θ≤ 1°) vs PD2(0.3°≤ θ≤ 1°)
PD2(0.7°≤ θ≤ 1°)
PD2(0.3°≤ θ≤ 1°)
3) Tail of gaussian
Beam (1mm beam
going thr 2.5mmx2
hole)
Mirror with larger
hole being prepared
2) Efficiency difference?
1) Cleaning issue?
Why different?
iLIGO
ETM04
PD1(1°≤ θ≤ 75°)
lens in path
lens out path

## Slide 13: Summary
Extended Integrating Sphere measurement
  Preliminary result – proof of concept
  In the small solid angle (10-4 , θ<1°), comparable energy observed as in the large solid angle(θ≥1°)
  Improvement of measurement setup and understanding of systematic uncertainties are necessary to quantify the conclusion
Large angle (θ≥1°) TIS may underestimate the loss
Measurement of aLIGO test masses coated by LMA
  Necessary to quantify the loss contribution from defects
Quantitative comparison of larger and small angle TIS may provide further information of defects
  Size, clustering, …

## Slide 14: End of slides

## Slide 15: Defect size and total loss assuming defect shape is circle
Defect size vs TIS(θ≥1°)
Total / TIS(θ≥1°) vs TIS(θ≥1°)

## Slide 16: Loss function