## Slide 1
Laser Power meter calibrations at LNE
LCM-LNE-CNAM
J. Dubard, O. Enouf, F. Issard, P. Pinot, Z. Silvestri
GW (Gravitational Wave) Metrology Workshop, March 14-15, 2019

## Slide 2: OUTLINE
Introduction
Calibration capabilities
Traceability
New development in laser power measurement based on radiation pressure
Conclusion
2

## Slide 3: INTRODUCTION
French metrology is managed by LNE and is based on:
4 National institutes: LNE, CNAM, CEA-LNHB, SYRTE-OP
6 DI’s (ENSAM, IRSN, CETIAT, LTFB, LADG, TRAPIL)
Radiometry-Photometry:
CNAM: set-up the references according to CCPR KC
LNE: set-up complementary references (laser power, fiber optics, radiance,…), and transfert to industry
3

## Slide 4: Calibration capabilities
Power measurement
Wavelengths (nm): 532, 633, 1064, 1070, 10600
Power range: 0.1 mW to 1 kW
Laser sources |  |  |  | 
Spectra-Physics | 127 | 633 | 40 mW | 2 mm
Spectron Laser Systems | SL 301 | 1064 | 4 W single 50 W multi | 4 mm
MPB tech | GN-802-MES | 10600 | 15 W | div
Spectra-Physics | Millennia eV25 | 532 | 25 W | 3 mm
IPG PHOTONICS | P30-001298 | 1070 | 1100 W | 14 mm
4

## Slide 5: Calibration capabilities
Energy measurement
Wavelengths (nm): 1064, 532, 355
Energy: up to 700 mJ
Laser: QUANTEL, Brilliant B
Pulse width: 7 ns
Repetition rate: 20 Hz
5

## Slide 6: Calibration capabilities
CMC
6

## Slide 7: TRACEABILITY
Traceability through home made laser power meters
7

## Slide 8: TRACEABILITY
Previous traceability was through the electrical quantities. This has shown a systematic error of 2 % (EUROMET comparison 156)
Nowdays traceability is to the cryogenic radiometer through a Si trap detector:
3 Hamamatsu S1227 SI detector
Calibrated at a power levels of 100 µW and 500 µW
Linearity has been demonstrated at 5 mW (cryogenic radiometer) and at 20 mW (comparison to PMOD scale with 0.1% agreement)
« Nonlinearity of the quantum efficiency of Si reflection trap detector at 633 nm », K-D. Stock et al, Metrologia, 35, 1998
8

## Slide 9: TRACEABILITY
Cryogenic radiometer
Si Trap detector
λ: 488, 514, 633
P: 100 -500 µW
Radiometer WL001
Conduction
λ: 532
P: 500 µW -5 mW
Radiometer WLR2
Conduction + electrical substitution
λ: 532
P: 100 mW + Attenuator
WL001
WLR2
9

## Slide 10: TRACEABILITY
Radiometer WLR2
Conduction + electrical substitution
Radiometer CLR01
Conduction + electrical substitution
λ: 532
P: 10 W+ Attenuator
Radiometer CLR5
Conduction + electrical substitution
λ: 1070
P: 300 W+ Attenuator
Radiometer CLR30
Conduction + electrical substitution
λ: 1070
P: 1000 W+ Attenuator
10

## Slide 11
Cryogenic radiometer is not necessary
Alternative to thermopiles and complementarity to photodiodes
Simple, compact and reliable design
Possible use on an operational laser beam for laser intensity and stability control (surgery, additive manufacturing ...)
> Development of an absolute standard of an optical power meter
Directly traceable to the new SI
Based on pressure radiation using magnets + pyrolytic graphite (PyC)
Range 100 mW – 10 W
Wavelenght range 0,5 µm <  < 10,6 µm
Expected relative uncertainty 1 %
NEW DEVELOPMENT: radiation pressure
11

## Slide 12: Pyrolytic graphite
Main characteristics of Pyrolitic Craphite
  CVD fabrication
  Consisting of a stack of layers of graphene
  Ultra pure and non-porous
  Anisotropic properties :
   Thermal and electrical conductivities
   Diamagnetic (c z=-4,5  10-4 c x,y=-8,5  10-5)
x
y
z
PyC in levitation on NdFeB magnets
at room temperature
supermagnete.fr
supermagnete.fr
PyC
NdFeB magnets
0,5 mm
Magnet
Diamagnetic material
Repulsive
force
Repulsive
force
NEW DEVELOPMENT: radiation pressure
12

## Slide 13
NEW DEVELOPMENT: radiation pressure
ki : spring constant
Arrangement of 9 NdFeB magnets
2D distribution of B
Simulation Ansys
P. Pinot et Z. Silvestri, Optical power meter using radiation pressure measurement, Measurement, Vol. 131, (2019)
Mass of the optical part: 1.18 g
13

## Slide 14: First results
kg m-2
= 1,1 s
= 1 cm
S = 1,75 V/mm
r = 0,999
: period
Relative uncertainty : Now  7 %
Objective  1 %
NEW DEVELOPMENT: radiation pressure
14

## Slide 15: CONCLUSION
Calibration capabilities up to 1 kW CW with uncertainty from 1.3% (500 µW) to 2.7% (1 kW)
New development of a power meter based on radiation pressure using pyrolitic graphite levitation property
Intermediate range of power (100 mW - 10 W)
Compact, simple and transportable
Directly traceable to the new SI
Temperature and vibration sensitive
Uncertainty down to 1%
15

## Slide 16
Thank you for your attention
16