| AUTHORS | DATE | Document Change Notice, Release or Approval |
| --- | --- | --- |
| AUTHOR: G. Billingsley | 05/31/19 | E1900155 x0 |

| Applicable Documents |  |
| --- | --- |
| LIGO-D1900146-v1 | Mirror Blank Drawing A+ Filter Cavity End Mirror |
| MIL-G-174-B | Glass, Optical |
| Requirements |  |
| Physical Dimensions | Right Circular Cylinder: 78 mm x 153 mm ø Per LIGO-D1900146-v1 |
| Clear Aperture | Central 120 mm |
| Serial Number | Blanks shall be serialized as FEXX, where XX increments starting at 01 |
| Material | High Purity Fused Silica |
| Final shaping | Shaping shall be performed using a progression of grit size ending with a 320 or smaller grit tool. |
| Defect depth | Maximum on any surface or corner is less than 0.5 mm |
| Homogeneity | ≤ 4 x 10-6 P-V in clear aperture |
| Birefringence | ≤ 5 nm/cm |
| Bubble and Inclusion Cross section within the clear aperture | No bubbles or inclusions within 5mm of the flat surfaces within the clear aperture, < 1 mm total inclusion cross-section |
| Striae within the clear aperture | Class A according to MIL-G-174 |
| OH Content | ≤ 1000 ppm |

| Specification | Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Physical Dimensions | Inspection | 100% | Certification |
| Serial number | Visual Inspection | 100% | Certification |
| Material | Process Control Material Certification | 100% | Certification |
| Defect depth | Visual Inspection | 100% | Certification |
| Birefringence | MIL-G-174 Section 4.4.5 | 100% | Certification |
| Inclusions | Visual Inspection | 100% | Hand sketch indicating location and dimensions |

Table 1: MEASUREMENT MATRIX: FREQUENCY AND METHOD