## Slide 1: T2000526 Ink and First Contact Damage test Sample SN 0919
Zhang, Billingsley

## Slide 2
Witness sample 0919 ITM-type HR witness sample, Initial Absorption Scan, no absorbers found

## Slide 3
Experimental Plan
Witness sample 0919 coating run C11068 HR coating
First line applied near the bottom (label “ink cleaned three times”)
Wait several hours until completely dry
Apply a thin layer of first contact with brush
Wait several hours before removing
Second line applied (label “ink cleaned twice”)
Clean both lines as above
Third line applied (label “ink cleaned once”)
Clean all three lines as above
Fourth line applied (label “ink”)
Very thin patch of first contact applied at the bottom
Wait several hours
Pull leaving traces
Irradiate in RTS 2W/mm2 - Inspect
Irradiate in RTS 20W/mm2 - Inspect
Irradiate in RTS 200 W/mm2 - Inspect

## Slide 4
A note on First Contact
A thin layer does not allow for much solvent to work on the contaminant
  Once the contaminant is freed from the surface, it is entangled in the polymer
First contact is optimally adherent soon after drying, ideally being removed within an hour of application for a poured layer - less for a painted layer.
This experiment seeks to employ the worst of both conditions -
  a very thin layer
  an 8 hour drying time

## Slide 5: Scan at 2 W/mm2

## Slide 6: Inspection after 2 W/mm2
Only one damage site; in line 4 “ink”
Tracks of the scan can be seen in line 4 “ink”
Confirmation of damage is best seen looking through the optic.
Ink from other lines can be seen under the microscope

## Slide 7: Damage found at 2W/mm2 in one location. Easiest to see when viewed through the optic Possibly at another power level, done during stage movement.

## Slide 8: The ghost of line one
A ghost of the first ink line can be seen near some First Contact shards

## Slide 9
Scan at 20 W/mm2
The four lines can be seen
The first contact can be seen

## Slide 10: Inspection after 20 W/mm2
Excessive Damage on line 4- Ink
Damage on lines 2 and 3
No Damage on line 1
No Damage in First Contact
Perhaps one damage site that was not on a line? We usually clean before this type of test, but the sample is getting quite dirty. (below)

## Slide 11: No damage in line 1,
The remaining, visible part of line 1 - no damage
Viewed on the front of the optic in dark field. Cannot be seen through the optic (in DF)

## Slide 12: Viewed through the optic in dark field

## Slide 13: Viewed through the optic in dark field
Small damage site in line 2 First contact shard shows no damage

## Slide 14
Scan at 200 W/mm2
The optic is quite dirty at this point.
The four lines can be seen
The first contact can be seen

## Slide 15: Inspection after 200 W/mm2 - Optic now cleaned

## Slide 16
How it actually turned out - after all damage runs
Ink (line 4)
1 FC cleaning (line 3)
2 FC cleaning (line 2)
3 FC cleaning (line 1)
First Contact string
First Contact film
All lines are at least partially visible, even with the high background from the mount and the damage

## Slide 17
After cleaning
Ink (line 4)
1 FC cleaning (line 3)
2 FC cleaning (line 2)
3 FC cleaning (line 1)
First Contact string
First Contact film

## Slide 18: Absorption measurement after cleaning (and unintentional damage)

## Slide 19: Comparing before/after cleaning Accidental damage occurred after cleaning at (0,-y) when irradiated with 2000 w/mm2 Not clear if there is still absorption, or if the signal comes from distortion

## Slide 20: Damage filled with first contact after cleaning
Attempted Cleaning:
Acetone Soak in ultrasonic bath ~ 1 hour
  No noticeable change
First Contact thinner in ultrasonic bath ~ 1 hour
  No noticeable change
Images are at different magnifications

## Slide 21: Does HWS see absorption from First Contact? OR from embedded ink?

## Slide 23: Probe (HeNe) beam background, no pump beam
An RTS scan with the IR (pump) beam off. A slight increase of the background at the damaged points is due to lower He-Ne transmission, making the noise equivalent absorption appear larger.

## Slide 24: No scatter signal beyond background from Pump beam

## Slide 25: First Contact absorption on an uncoated optic 60W/mm2 No damage identified after FC removal

## Slide 26: The ink that we use for optic fiducials