LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T2500040-v1	Advanced LIGO	February 21, 2025

Delay Measurements of REFL RF Detectors

Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Executive Summary

As part of the ISI upgrade in HAM1 the cable length for the coaxial cables will change for currently installed REFL and POP RF detectors. Therefore, the delay of the associated demodulator will have to be change accordingly.

Here we propose two sets of measurements to determine the phase delay both before and after the changeover. The first measurement is using the excitation inputs to measure the total phase including twice the cable length plus the internal phase shift. Since the internal phase shift will stay the same, we can back out the delay of a single cable run. This measurement can be done ahead of time in a maintenance window, and immediately after the detectors are reinstalled as part of the electrical checkout.

The second measurement requires a Time-domain reflectometer (TDR). To make this measurement one has to disconnect the coax cables from the detectors first. This will guarantee a good reflection at the cable end. This needs to be done as one of the first in-vacuum tasks with the cables still connected to the feedthroughs. It has to be repeated after reinstalling the detectors but before the cables are connected to the sensors.

# Sensors

The effected sensors are listed in the table below:

| Detector | Change | Reason for change |
| --- | --- | --- |
| LSC-REFL_A | Longer in-air cables | Feedthrough port changed from D5 to D6 |
| LSC-REFL_B | Longer in-air cables |  |
| LSC-POP_A | No change | - |
| ASC-REFL_A | Longer in-air cables | Feedthrough port changed from D6 to D1 |
| ASC-REFL_B | Longer in-air cables |  |
| ASC-POP_X | Longer cables | Move into vacuum |

# Setup

Currently all coax cables are terminated at patch panels in rack ISC-R2 except LSC REFL_B which is only present LHO. The most recent drawing for ISC-R2 is shown in Figure 3, which includes separate electronics for LSC REFL_B and removes the patch panels for the in-air WFS.

In addition, LSC REFL_B currently implements a single coax cable only. Therefore, only a TDR measurement can be performed.

The phase measurement is performed with an RF network analyzer. It is important to keep the same measurement cables for the before and after measurement, since these cables will add to the delay. Alternatively, one can characterize the delay added by the measurement cables.

## LSC Patch Panel

Figure 1: LSC patch panel. Each LSC sensor has a coax cable for a test input (Test In), and three RF readback signals (RF1, RF2, and Test Out). RF1 is used by the 9 MHz signal, and RF2 is used by the 45 MHz signal.

Figure 3: Annotated Rack drawing of ISC-R2.

ASC Patch Panel

Figure 2: Wavefront Sensor Demodulator Patch Panel. Each wavefront sensor includes an excitation input (Test In) and eight signal outputs (9 and 45 MHz signals for segments 1 thru 4, or 36 and 45MHz for segments 1 thru 4).

Add REFL_B

Coax cables

move

ASC patch panel

REFAL_A

LSC patch panel

Figure 4: Photo of ISC-R2 rack (center part).

Figure 5: Photo of ISC-R2 rack (bottom part).

ASC patch panel

POP-X

Coax cables

move

ASC patch panel

REFL_B

Coax cables

move

# Measurements

## Phase measurements

| Sensor | Signal | Input cable | Output cable | Before (°) | After (°) |
| --- | --- | --- | --- | --- | --- |
| None | 9 MHz | Measurement setup only |  |  |  |
|  | 36 MHz |  |  |  |  |
|  | 45 MHz |  |  |  |  |
| LSC-REFL_A | 9 MHz | ISC_RF21-B/3 | ISC_RF31-B/2 |  |  |
|  | 45 MHz |  | ISC_RF31-B/1 |  |  |
| LSC-POP_A | 9 MHz | ISC_RF30-B/3 | ISC_RF30-B/2 |  |  |
|  | 45 MHz |  | ISC_RF30-B/1 |  |  |
| ASC-REFL_A | 9 MHz | ISC_RF23-B/5 | ISC_RF23-B/1 |  |  |
|  |  |  | ISC_RF22-B/3 |  |  |
|  |  |  | ISC_RF22-B/1 |  |  |
|  |  |  | ISC_RF23-B/3 |  |  |
|  | 45 MHz |  | ISC_RF23-B/2 |  |  |
|  |  |  | ISC_RF22-B/4 |  |  |
|  |  |  | ISC_RF22-B/2 |  |  |
|  |  |  | ISC_RF23-B/4 |  |  |
| ASC-REFL_B | 9 MHz | ISC_RF25-B/5 | ISC_RF25-B/1 |  |  |
|  |  |  | ISC_RF24-B/3 |  |  |
|  |  |  | ISC_RF24-B/1 |  |  |
|  |  |  | ISC_RF25-B/3 |  |  |
|  | 45 MHz |  | ISC_RF25-B/2 |  |  |
|  |  |  | ISC_RF24-B/4 |  |  |
|  |  |  | ISC_RF24-B/2 |  |  |
|  |  |  | ISC_RF25-B/4 |  |  |
| ASC-POP_X | 36 MHz | ISC_RF37-B/5 | ISC_RF37-B/1 |  |  |
|  |  |  | ISC_RF36-B/3 |  |  |
|  |  |  | ISC_RF36-B/1 |  |  |
|  |  |  | ISC_RF37-B/3 |  |  |
|  | 45 MHz |  | ISC_RF37-B/2 |  |  |
|  |  |  | ISC_RF36-B/4 |  |  |
|  |  |  | ISC_RF36-B/2 |  |  |
|  |  |  | ISC_RF37-B/4 |  |  |

## TDR Measurements

| Sensor | Signal | Output cable | Before (ns) | After (ns) |
| --- | --- | --- | --- | --- |
| None | 9 MHz | Measurement setup only |  |  |
| LSC-REFL_A | 9 MHz | ISC_RF21-B/2 |  |  |
|  | 45 MHz | ISC_RF21-B/1 |  |  |
| LSC-REFL_A | 9 MHz | ISC_RF31-B/2 |  |  |
| LSC-POP_A | 9 MHz | ISC_RF30-B/2 |  |  |
|  | 45 MHz | ISC_RF30-B/1 |  |  |
| ASC-REFL_A | 9 MHz | ISC_RF23-B/1 |  |  |
|  |  | ISC_RF22-B/3 |  |  |
|  |  | ISC_RF22-B/1 |  |  |
|  |  | ISC_RF23-B/3 |  |  |
|  | 45 MHz | ISC_RF23-B/2 |  |  |
|  |  | ISC_RF22-B/4 |  |  |
|  |  | ISC_RF22-B/2 |  |  |
|  |  | ISC_RF23-B/4 |  |  |
| ASC-REFL_B | 9 MHz | ISC_RF25-B/1 |  |  |
|  |  | ISC_RF24-B/3 |  |  |
|  |  | ISC_RF24-B/1 |  |  |
|  |  | ISC_RF25-B/3 |  |  |
|  | 45 MHz | ISC_RF25-B/2 |  |  |
|  |  | ISC_RF24-B/4 |  |  |
|  |  | ISC_RF24-B/2 |  |  |
|  |  | ISC_RF25-B/4 |  |  |
| ASC-POP_X | 36 MHz | ISC_RF37-B/1 |  |  |
|  |  | ISC_RF36-B/3 |  |  |
|  |  | ISC_RF36-B/1 |  |  |
|  |  | ISC_RF37-B/3 |  |  |
|  | 45 MHz | ISC_RF37-B/2 |  |  |
|  |  | ISC_RF36-B/4 |  |  |
|  |  | ISC_RF36-B/2 |  |  |
|  |  | ISC_RF37-B/4 |  |  |